## Summary
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:25:44 PM by version Alpha5 - Build 2023 Jul. 26
1. Configuration for 3P4W Meter
Equipment Setup
Option Name |  |  |  |  | Value
Communication Port |  |  |  |  | COM15
Communication Port Settings |  |  |  |  | 9600, 8, None, One
Timeout, Delay and Retries |  |  |  |  | 0, 50, 5000, 0, 100, 5000, 2
Equipment Version |  |  |  |  | GF302D
Control phase angle between YR (U1U2) and YB (U1U3) |  |  |  |  | Yes
Indicates RYB channels by colors |  |  |  |  | Yes
Meter Platform Name |  |  |  |  | Global Reference Meter
Serial Number |  |  |  |  | SED-IV0004
Meter Type |  |  |  |  | TYPE_3P4W
Meter Class |  |  |  |  | Class 0.5S
Meter Rating |  |  |  |  | Non-specified
Meter Maximum Current (Imax) |  |  |  |  | 40
Meter Basic Voltage (Un or Ub) |  |  |  |  | 220
Meter Basic Current (Ib or In) |  |  |  |  | 5
Meter Frequency (Fb) |  |  |  |  | 50
Meter Constant Active Power (P) |  |  |  |  | 3200
Meter Constant Reactive Power (Q) |  |  |  |  | 3200
Meter % Error Average Length |  |  |  |  | 3
Meter % Error Rounding Decimals |  |  |  |  | 3
Discard first few %error responses |  |  |  |  | Yes
Discard first few %error, number of responses will be discarded |  |  |  |  | 1
Chain voltage or current among test cases |  |  |  |  | No
Meter Configuration
Option Name |  |  |  |  | Value
Optical Port
Optical Port Settings |  |  |  |  | 9600, 8, None, One
External COM Port
External COM Port Settings |  |  |  |  | 19200, 8, None, One
Private Protocol Settings |  |  |  |  | abcd.efgh.xx@renesas.com, 0, 50, 100, 20000, 1
2. Accuracy Test
Below table summarizes all test categories and test result
No | Path In Workspace | Test Category | Passed | Failed | Inconclusive | Total | Overall Result | Accuracy Chart | Details
1 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / PF = 1.0 | PF = 1.0 | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart1. PF = 1.0.png | [>>]
2 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / PF = 0.5L | PF = 0.5L | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart2. PF = 0.5L.png | [>>]
3 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / PF = 0.8C | PF = 0.8C | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart3. PF = 0.8C.png | [>>]
4 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / 5th Harmonics | 5th Harmonics | 1 | 0 | 0 | 1 | PASSED | Temp\AccuracyChart4. 5th Harmonics.png | [>>]
5 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 1.0 | Voltage Variation, PF = 1.0 | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart5. Voltage Variatio....png | [>>]
6 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 0.5L | Voltage Variation, PF = 0.5L | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart6. Voltage Variatio....png | [>>]
7 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 1.0 | Frequency Variation, PF = 1.0 | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart7. Frequency Varia....png | [>>]
8 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 0.5L | Frequency Variation, PF = 0.5L | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart8. Frequency Varia....png | [>>]
9 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / PF = 1.0 | PF = 1.0 | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart9. PF = 1.0.png | [>>]
10 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / PF = 0.5L | PF = 0.5L | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart10. PF = 0.5L.png | [>>]
11 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / PF = 0.8C | PF = 0.8C | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart11. PF = 0.8C.png | [>>]
12 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / 5th Harmonics | 5th Harmonics | 1 | 0 | 0 | 1 | PASSED | Temp\AccuracyChart12. 5th Harmonics.png | [>>]
13 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 1.0 | Voltage Variation, PF = 1.0 | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart13. Voltage Variatio....png | [>>]
14 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 0.5L | Voltage Variation, PF = 0.5L | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart14. Voltage Variatio....png | [>>]
15 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 1.0 | Frequency Variation, PF = 1.0 | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart15. Frequency Varia....png | [>>]
16 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 0.5L | Frequency Variation, PF = 0.5L | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart16. Frequency Varia....png | [>>]
17 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / PF = 1.0 | PF = 1.0 | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart17. PF = 1.0.png | [>>]
18 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / PF = 0.5L | PF = 0.5L | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart18. PF = 0.5L.png | [>>]
19 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / PF = 0.8C | PF = 0.8C | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart19. PF = 0.8C.png | [>>]
20 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / 5th Harmonics | 5th Harmonics | 1 | 0 | 0 | 1 | PASSED | Temp\AccuracyChart20. 5th Harmonics.png | [>>]
21 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 1.0 | Voltage Variation, PF = 1.0 | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart21. Voltage Variatio....png | [>>]
22 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 0.5L | Voltage Variation, PF = 0.5L | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart22. Voltage Variatio....png | [>>]
23 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 1.0 | Frequency Variation, PF = 1.0 | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart23. Frequency Varia....png | [>>]
24 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 0.5L | Frequency Variation, PF = 0.5L | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart24. Frequency Varia....png | [>>]
25 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / PF = 1.0 | PF = 1.0 | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart25. PF = 1.0.png | [>>]
26 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / PF = 0.5L | PF = 0.5L | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart26. PF = 0.5L.png | [>>]
27 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / PF = 0.8C | PF = 0.8C | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart27. PF = 0.8C.png | [>>]
28 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / 5th Harmonics | 5th Harmonics | 1 | 0 | 0 | 1 | PASSED | Temp\AccuracyChart28. 5th Harmonics.png | [>>]
29 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 1.0 | Voltage Variation, PF = 1.0 | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart29. Voltage Variatio....png | [>>]
30 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 0.5L | Voltage Variation, PF = 0.5L | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart30. Voltage Variatio....png | [>>]
31 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 1.0 | Frequency Variation, PF = 1.0 | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart31. Frequency Varia....png | [>>]
32 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 0.5L | Frequency Variation, PF = 0.5L | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart32. Frequency Varia....png | [>>]

## 1. PF = 1.0
[<<]
Accuracy Test Report of 'PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:20:11 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 20 | 20 |  | ±0.1% | ±0.5% | 0.03% | PASSED | 2023-07-25 15:19
Yes | 1 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 12 | 20 |  | ±0.1% | ±0.5% | 0.013% | PASSED | 2023-07-25 15:20
Yes | 3 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.5% | 0.03% | PASSED | 2023-07-25 15:20
Yes | 4 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0.038% | PASSED | 2023-07-25 15:20
Yes | 7 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0.047% | PASSED | 2023-07-25 15:21
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 0.009% | 0% | 0% | 0.1%
0 | 1 | 0.062% | 0.053% | 0% | 0.1%
0 | 2 | 0.018% | -0.044% | 0.03% | 0.1%
0 | 3 | 0.009% | -0.009% | 0.03% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -0.012% | 0% | 0% | 0.1%
1 | 1 | 0% | 0.012% | 0% | 0.1%
1 | 2 | 0.012% | 0.012% | 0% | 0.1%
1 | 3 | 0.027% | 0.015% | 0.013% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0% | 0% | 0% | 0.1%
3 | 1 | 0.009% | 0.009% | 0% | 0.1%
3 | 2 | 0.022% | 0.013% | 0.01% | 0.1%
3 | 3 | 0.058% | 0.036% | 0.03% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.018% | 0% | 0% | 0.1%
4 | 1 | 0.009% | -0.009% | 0% | 0.1%
4 | 2 | 0.062% | 0.053% | 0.03% | 0.1%
4 | 3 | 0.044% | -0.018% | 0.038% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.048% | 0% | 0% | 0.1%
7 | 1 | 0.036% | -0.012% | 0% | 0.1%
7 | 2 | 0.052% | 0.016% | 0.045% | 0.1%
7 | 3 | 0.052% | 0% | 0.047% | 0.1%

## 2. PF = 0.5L
[<<]
Accuracy Test Report of 'PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:20:13 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | -0.083% | PASSED | 2023-07-25 15:22
Yes | 1 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | -0.109% | PASSED | 2023-07-25 15:22
Yes | 3 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.021% | PASSED | 2023-07-25 15:23
Yes | 4 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.05% | PASSED | 2023-07-25 15:24
Yes | 7 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.074% | PASSED | 2023-07-25 15:25
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 19 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -35.48% | 0% | 0% | 0.1%
0 | 1 | -34.68% | 0.801% | 0% | 0.1%
0 | 2 | -26.4% | 8.281% | -32.19% | 0.1%
0 | 3 | -16.75% | 9.651% | -25.95% | 0.1%
0 | 4 | -2.95% | 13.8% | -15.37% | 0.1%
0 | 5 | -0.653% | 2.297% | -6.785% | 0.1%
0 | 6 | -0.071% | 0.582% | -1.225% | 0.1%
0 | 7 | 0.375% | 0.446% | -0.116% | 0.1%
0 | 8 | 0.25% | -0.125% | 0.185% | 0.1%
0 | 9 | 0.393% | 0.143% | 0.339% | 0.1%
0 | 10 | 0.698% | 0.305% | 0.447% | 0.1%
0 | 11 | 1.498% | 0.8% | 0.863% | 0.1%
0 | 12 | 1.024% | -0.474% | 1.073% | 0.1%
0 | 13 | 0.68% | -0.344% | 1.067% | 0.1%
0 | 14 | 0.303% | -0.377% | 0.669% | 0.1%
0 | 15 | -0.089% | -0.392% | 0.298% | 0.1%
0 | 16 | -0.071% | 0.018% | 0.048% | 0.1%
0 | 17 | -0.089% | -0.018% | -0.083% | 0.1%
0 | 18 | -0.089% | 0% | -0.083% | 0.1%
Deviations of %error on test id #1. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -30.78% | 0% | 0% | 0.1%
1 | 1 | -19.82% | 10.96% | 0% | 0.1%
1 | 2 | -4.865% | 14.95% | -18.49% | 0.1%
1 | 3 | -0.548% | 4.317% | -8.41% | 0.1%
1 | 4 | 0.065% | 0.613% | -1.783% | 0.1%
1 | 5 | 0.214% | 0.149% | -0.09% | 0.1%
1 | 6 | 0.142% | -0.072% | 0.14% | 0.1%
1 | 7 | 0.393% | 0.251% | 0.25% | 0.1%
1 | 8 | 0.927% | 0.534% | 0.487% | 0.1%
1 | 9 | 0.626% | -0.301% | 0.649% | 0.1%
1 | 10 | 0.202% | -0.424% | 0.585% | 0.1%
1 | 11 | -0.041% | -0.243% | 0.262% | 0.1%
1 | 12 | -0.095% | -0.054% | 0.022% | 0.1%
1 | 13 | -0.089% | 0.006% | -0.075% | 0.1%
1 | 14 | -0.142% | -0.053% | -0.109% | 0.1%
Deviations of %error on test id #3. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -8.231% | 0% | 0% | 0.1%
3 | 1 | -5.771% | 2.46% | 0% | 0.1%
3 | 2 | 0.214% | 5.985% | -4.596% | 0.1%
3 | 3 | 0.107% | -0.107% | -1.817% | 0.1%
3 | 4 | 0.009% | -0.098% | 0.11% | 0.1%
3 | 5 | 0.16% | 0.151% | 0.092% | 0.1%
3 | 6 | 0.303% | 0.143% | 0.157% | 0.1%
3 | 7 | 0.402% | 0.099% | 0.288% | 0.1%
3 | 8 | 0.509% | 0.107% | 0.405% | 0.1%
3 | 9 | 0.366% | -0.143% | 0.426% | 0.1%
3 | 10 | 0.214% | -0.152% | 0.363% | 0.1%
3 | 11 | 0.134% | -0.08% | 0.238% | 0.1%
3 | 12 | 0.018% | -0.116% | 0.122% | 0.1%
3 | 13 | 0.036% | 0.018% | 0.063% | 0.1%
3 | 14 | -0.009% | -0.045% | 0.015% | 0.1%
3 | 15 | 0.036% | 0.045% | 0.021% | 0.1%
Deviations of %error on test id #4. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -25.61% | 0% | 0% | 0.1%
4 | 1 | -8.655% | 16.96% | 0% | 0.1%
4 | 2 | -0.018% | 8.637% | -11.43% | 0.1%
4 | 3 | 0.608% | 0.626% | -2.688% | 0.1%
4 | 4 | 0.366% | -0.242% | 0.319% | 0.1%
4 | 5 | 0.366% | 0% | 0.447% | 0.1%
4 | 6 | 0.214% | -0.152% | 0.315% | 0.1%
4 | 7 | 0.089% | -0.125% | 0.223% | 0.1%
4 | 8 | 0.036% | -0.053% | 0.113% | 0.1%
4 | 9 | 0.053% | 0.017% | 0.059% | 0.1%
4 | 10 | 0.062% | 0.009% | 0.05% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.089% | 0% | 0% | 0.1%
7 | 1 | 0.066% | -0.023% | 0% | 0.1%
7 | 2 | 0.085% | 0.019% | 0.08% | 0.1%
7 | 3 | 0.071% | -0.014% | 0.074% | 0.1%

## 3. PF = 0.8C
[<<]
Accuracy Test Report of 'PF = 0.8C'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / PF = 0.8C'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:20:16 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | -0.442% | PASSED | 2023-07-25 15:26
Yes | 1 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | -0.053% | PASSED | 2023-07-25 15:26
Yes | 3 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | -0.083% | PASSED | 2023-07-25 15:27
Yes | 4 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | -0.162% | PASSED | 2023-07-25 15:27
Yes | 7 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.03% | PASSED | 2023-07-25 15:28
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -10% | 0% | 0% | 0.1%
0 | 1 | -9.58% | 0.42% | 0% | 0.1%
0 | 2 | -4.999% | 4.581% | -8.193% | 0.1%
0 | 3 | -4.999% | 0% | -6.526% | 0.1%
0 | 4 | -1.194% | 3.805% | -3.731% | 0.1%
0 | 5 | -0.583% | 0.611% | -2.259% | 0.1%
0 | 6 | -0.124% | 0.459% | -0.634% | 0.1%
0 | 7 | 0.053% | 0.177% | -0.218% | 0.1%
0 | 8 | 0.071% | 0.018% | 0% | 0.1%
0 | 9 | 0.071% | 0% | 0.065% | 0.1%
0 | 10 | -0.177% | -0.248% | -0.012% | 0.1%
0 | 11 | -0.053% | 0.124% | -0.053% | 0.1%
0 | 12 | -0.106% | -0.053% | -0.112% | 0.1%
0 | 13 | -0.036% | 0.07% | -0.065% | 0.1%
0 | 14 | -0.177% | -0.141% | -0.106% | 0.1%
0 | 15 | -0.354% | -0.177% | -0.189% | 0.1%
0 | 16 | -0.636% | -0.282% | -0.389% | 0.1%
0 | 17 | -0.548% | 0.088% | -0.513% | 0.1%
0 | 18 | -0.46% | 0.088% | -0.548% | 0.1%
0 | 19 | -0.319% | 0.141% | -0.442% | 0.1%
Deviations of %error on test id #1. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -10.31% | 0% | 0% | 0.1%
1 | 1 | -6.057% | 4.253% | 0% | 0.1%
1 | 2 | -5.404% | 0.653% | -7.257% | 0.1%
1 | 3 | -1.055% | 4.349% | -4.172% | 0.1%
1 | 4 | -0.875% | 0.18% | -2.445% | 0.1%
1 | 5 | 0.059% | 0.934% | -0.624% | 0.1%
1 | 6 | 0.036% | -0.023% | -0.26% | 0.1%
1 | 7 | -0.089% | -0.125% | 0.002% | 0.1%
1 | 8 | -0.106% | -0.017% | -0.053% | 0.1%
1 | 9 | -0.041% | 0.065% | -0.079% | 0.1%
1 | 10 | -0.012% | 0.029% | -0.053% | 0.1%
Deviations of %error on test id #3. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -8.566% | 0% | 0% | 0.1%
3 | 1 | -8.499% | 0.067% | 0% | 0.1%
3 | 2 | -3.78% | 4.719% | -6.948% | 0.1%
3 | 3 | -3.797% | -0.017% | -5.359% | 0.1%
3 | 4 | -3.159% | 0.638% | -3.579% | 0.1%
3 | 5 | -0.53% | 2.629% | -2.495% | 0.1%
3 | 6 | -0.53% | 0% | -1.406% | 0.1%
3 | 7 | -0.257% | 0.273% | -0.439% | 0.1%
3 | 8 | -0.133% | 0.124% | -0.307% | 0.1%
3 | 9 | -0.124% | 0.009% | -0.171% | 0.1%
3 | 10 | -0.062% | 0.062% | -0.106% | 0.1%
3 | 11 | -0.062% | 0% | -0.083% | 0.1%
Deviations of %error on test id #4. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -4.1% | 0% | 0% | 0.1%
4 | 1 | -1.652% | 2.448% | 0% | 0.1%
4 | 2 | -0.381% | 1.271% | -2.044% | 0.1%
4 | 3 | -0.106% | 0.275% | -0.713% | 0.1%
4 | 4 | -0.115% | -0.009% | -0.201% | 0.1%
4 | 5 | -0.195% | -0.08% | -0.139% | 0.1%
4 | 6 | -0.177% | 0.018% | -0.162% | 0.1%
Deviations of %error on test id #7. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | -0.101% | 0% | 0% | 0.1%
7 | 1 | 0.041% | 0.142% | 0% | 0.1%
7 | 2 | 0.043% | 0.002% | -0.006% | 0.1%
7 | 3 | 0.021% | -0.022% | 0.035% | 0.1%
7 | 4 | 0.027% | 0.006% | 0.03% | 0.1%

## 4. 5th Harmonics
[<<]
Accuracy Test Report of '5th Harmonics'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / 5th Harmonics'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:20:21 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Harmonics' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
1 | 0 | 0 | 1
3. Test Cases
Total 1 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 1 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° | {UcUbUa, 5th-10%, IcIbIa, 5th-40%} | 12 | 20 |  | ±0.1% | ±1% | 0.117% | PASSED | 2023-07-25 15:29
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #1. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0.006% | 0% | 0% | 0.1%
1 | 1 | 0.223% | 0.217% | 0% | 0.1%
1 | 2 | 0.211% | -0.012% | 0.147% | 0.1%
1 | 3 | 0.113% | -0.098% | 0.182% | 0.1%
1 | 4 | 0.027% | -0.086% | 0.117% | 0.1%

## 5. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:20:23 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DP | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | 0.033% | PASSED | 2023-07-25 15:30
Yes | 1 | DP | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.028% | PASSED | 2023-07-25 15:30
Yes | 2 | DP | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.022% | PASSED | 2023-07-25 15:31
Yes | 3 | DP | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.031% | PASSED | 2023-07-25 15:31
Yes | 4 | DP | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.042% | PASSED | 2023-07-25 15:31
Yes | 5 | DP | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | 0.052% | PASSED | 2023-07-25 15:32
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 2.464% | 0% | 0% | 0.1%
0 | 1 | 0.734% | -1.73% | 0% | 0.1%
0 | 2 | 0.036% | -0.698% | 1.078% | 0.1%
0 | 3 | 0.058% | 0.022% | 0.276% | 0.1%
0 | 4 | 0.036% | -0.022% | 0.043% | 0.1%
0 | 5 | 0.004% | -0.032% | 0.033% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0.049% | 0% | 0% | 0.1%
1 | 1 | 0.018% | -0.031% | 0% | 0.1%
1 | 2 | 0.013% | -0.005% | 0.027% | 0.1%
1 | 3 | 0.053% | 0.04% | 0.028% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | 0.04% | 0% | 0% | 0.1%
2 | 1 | 0.013% | -0.027% | 0% | 0.1%
2 | 2 | 0.036% | 0.023% | 0.03% | 0.1%
2 | 3 | 0.018% | -0.018% | 0.022% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.04% | 0% | 0% | 0.1%
3 | 1 | 0.027% | -0.013% | 0% | 0.1%
3 | 2 | -0.013% | -0.04% | 0.018% | 0.1%
3 | 3 | 0.08% | 0.093% | 0.031% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.04% | 0% | 0% | 0.1%
4 | 1 | 0.071% | 0.031% | 0% | 0.1%
4 | 2 | 0.036% | -0.035% | 0.049% | 0.1%
4 | 3 | 0.018% | -0.018% | 0.042% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | 0.009% | 0% | 0% | 0.1%
5 | 1 | 0.067% | 0.058% | 0% | 0.1%
5 | 2 | 0.018% | -0.049% | 0.031% | 0.1%
5 | 3 | 0.071% | 0.053% | 0.052% | 0.1%

## 6. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:20:26 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DP | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | 0.223% | PASSED | 2023-07-25 15:33
Yes | 1 | DP | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.133% | PASSED | 2023-07-25 15:33
Yes | 2 | DP | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.238% | PASSED | 2023-07-25 15:34
Yes | 3 | DP | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.065% | PASSED | 2023-07-25 15:34
Yes | 4 | DP | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.166% | PASSED | 2023-07-25 15:35
Yes | 5 | DP | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | 0.092% | PASSED | 2023-07-25 15:35
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -21.3% | 0% | 0% | 0.1%
0 | 1 | -10.45% | 10.85% | 0% | 0.1%
0 | 2 | -3.367% | 7.084% | -11.71% | 0.1%
0 | 3 | 0.402% | 3.769% | -4.472% | 0.1%
0 | 4 | 0.554% | 0.152% | -0.804% | 0.1%
0 | 5 | 0.563% | 0.009% | 0.506% | 0.1%
0 | 6 | 0.357% | -0.206% | 0.491% | 0.1%
0 | 7 | 0.25% | -0.107% | 0.39% | 0.1%
0 | 8 | 0.294% | 0.044% | 0.3% | 0.1%
0 | 9 | 0.232% | -0.062% | 0.259% | 0.1%
0 | 10 | 0.142% | -0.09% | 0.223% | 0.1%
Deviations of %error on test id #1. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -29.04% | 0% | 0% | 0.1%
1 | 1 | -19.33% | 9.708% | 0% | 0.1%
1 | 2 | -9.5% | 9.831% | -19.29% | 0.1%
1 | 3 | -4.133% | 5.367% | -10.99% | 0.1%
1 | 4 | 0.116% | 4.249% | -4.506% | 0.1%
1 | 5 | 0.178% | 0.062% | -1.28% | 0.1%
1 | 6 | 0.357% | 0.179% | 0.217% | 0.1%
1 | 7 | 0.267% | -0.09% | 0.267% | 0.1%
1 | 8 | 0.285% | 0.018% | 0.303% | 0.1%
1 | 9 | 0.527% | 0.242% | 0.36% | 0.1%
1 | 10 | 0.294% | -0.233% | 0.369% | 0.1%
1 | 11 | 0.196% | -0.098% | 0.339% | 0.1%
1 | 12 | 0.151% | -0.045% | 0.214% | 0.1%
1 | 13 | 0.053% | -0.098% | 0.133% | 0.1%
Deviations of %error on test id #2. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -24.67% | 0% | 0% | 0.1%
2 | 1 | -14.32% | 10.35% | 0% | 0.1%
2 | 2 | -6.863% | 7.455% | -15.28% | 0.1%
2 | 3 | -2.504% | 4.359% | -7.895% | 0.1%
2 | 4 | 0.16% | 2.664% | -3.069% | 0.1%
2 | 5 | 0.366% | 0.206% | -0.659% | 0.1%
2 | 6 | 0.258% | -0.108% | 0.261% | 0.1%
2 | 7 | 0.142% | -0.116% | 0.255% | 0.1%
2 | 8 | 0.116% | -0.026% | 0.172% | 0.1%
2 | 9 | 0.241% | 0.125% | 0.166% | 0.1%
2 | 10 | 0.312% | 0.071% | 0.223% | 0.1%
2 | 11 | 0.241% | -0.071% | 0.265% | 0.1%
2 | 12 | 0.16% | -0.081% | 0.238% | 0.1%
Deviations of %error on test id #3. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -22.13% | 0% | 0% | 0.1%
3 | 1 | -8.313% | 13.82% | 0% | 0.1%
3 | 2 | -6.662% | 1.651% | -12.37% | 0.1%
3 | 3 | 0.196% | 6.858% | -4.926% | 0.1%
3 | 4 | 0.142% | -0.054% | -2.108% | 0.1%
3 | 5 | 0% | -0.142% | 0.113% | 0.1%
3 | 6 | 0.089% | 0.089% | 0.077% | 0.1%
3 | 7 | 0.267% | 0.178% | 0.119% | 0.1%
3 | 8 | 0.473% | 0.206% | 0.276% | 0.1%
3 | 9 | 0.59% | 0.117% | 0.443% | 0.1%
3 | 10 | 0.375% | -0.215% | 0.479% | 0.1%
3 | 11 | 0.33% | -0.045% | 0.432% | 0.1%
3 | 12 | 0.178% | -0.152% | 0.294% | 0.1%
3 | 13 | 0.107% | -0.071% | 0.205% | 0.1%
3 | 14 | 0.018% | -0.089% | 0.101% | 0.1%
3 | 15 | 0.071% | 0.053% | 0.065% | 0.1%
Deviations of %error on test id #4. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -10.18% | 0% | 0% | 0.1%
4 | 1 | -8.833% | 1.346% | 0% | 0.1%
4 | 2 | 0.107% | 8.94% | -6.302% | 0.1%
4 | 3 | 0.339% | 0.232% | -2.796% | 0.1%
4 | 4 | 0.276% | -0.063% | 0.241% | 0.1%
4 | 5 | 0.241% | -0.035% | 0.285% | 0.1%
4 | 6 | 0.151% | -0.09% | 0.223% | 0.1%
4 | 7 | 0.107% | -0.044% | 0.166% | 0.1%
Deviations of %error on test id #5. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -14.97% | 0% | 0% | 0.1%
5 | 1 | -10.22% | 4.75% | 0% | 0.1%
5 | 2 | -2.606% | 7.616% | -9.267% | 0.1%
5 | 3 | 0.464% | 3.07% | -4.121% | 0.1%
5 | 4 | 0.339% | -0.125% | -0.601% | 0.1%
5 | 5 | 0.33% | -0.009% | 0.378% | 0.1%
5 | 6 | 0.312% | -0.018% | 0.327% | 0.1%
5 | 7 | 0.491% | 0.179% | 0.378% | 0.1%
5 | 8 | 0.384% | -0.107% | 0.396% | 0.1%
5 | 9 | 0.572% | 0.188% | 0.482% | 0.1%
5 | 10 | 0.473% | -0.099% | 0.476% | 0.1%
5 | 11 | 0.393% | -0.08% | 0.479% | 0.1%
5 | 12 | 0.196% | -0.197% | 0.354% | 0.1%
5 | 13 | 0.142% | -0.054% | 0.244% | 0.1%
5 | 14 | 0.062% | -0.08% | 0.133% | 0.1%
5 | 15 | 0.071% | 0.009% | 0.092% | 0.1%

## 7. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:20:34 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.006% | PASSED | 2023-07-25 15:39
Yes | 1 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0% | PASSED | 2023-07-25 15:40
Yes | 2 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.024% | PASSED | 2023-07-25 15:40
Yes | 3 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.015% | PASSED | 2023-07-25 15:41
Yes | 4 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.016% | PASSED | 2023-07-25 15:41
Yes | 5 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.032% | PASSED | 2023-07-25 15:41
Yes | 6 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.022% | PASSED | 2023-07-25 15:42
Yes | 7 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.022% | PASSED | 2023-07-25 15:42
Yes | 8 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.043% | PASSED | 2023-07-25 15:43
Yes | 9 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.047% | PASSED | 2023-07-25 15:43
Yes | 10 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.047% | PASSED | 2023-07-25 15:44
Yes | 11 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.056% | PASSED | 2023-07-25 15:44
Yes | 12 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.062% | PASSED | 2023-07-25 15:44
Yes | 13 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.065% | PASSED | 2023-07-25 15:45
Yes | 14 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.058% | PASSED | 2023-07-25 15:45
Yes | 15 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.052% | PASSED | 2023-07-25 15:46
Yes | 16 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.07% | PASSED | 2023-07-25 15:47
Yes | 17 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.084% | PASSED | 2023-07-25 15:47
Yes | 18 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.1% | PASSED | 2023-07-25 15:48
Yes | 19 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.08% | PASSED | 2023-07-25 15:48
Yes | 20 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.08% | PASSED | 2023-07-25 15:49
Yes | 21 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.083% | PASSED | 2023-07-25 15:49
Yes | 22 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.085% | PASSED | 2023-07-25 15:50
Yes | 23 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.089% | PASSED | 2023-07-25 15:50
Yes | 24 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.104% | PASSED | 2023-07-25 15:51
Yes | 25 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.111% | PASSED | 2023-07-25 15:52
Yes | 26 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.12% | PASSED | 2023-07-25 15:52
Yes | 27 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.111% | PASSED | 2023-07-25 15:53
Yes | 28 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.098% | PASSED | 2023-07-25 15:53
Yes | 29 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.132% | PASSED | 2023-07-25 16:02
Yes | 30 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.131% | PASSED | 2023-07-25 16:02
Yes | 31 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.126% | PASSED | 2023-07-25 16:02
Yes | 32 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.124% | PASSED | 2023-07-25 16:03
Yes | 33 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.146% | PASSED | 2023-07-25 16:03
Yes | 34 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.132% | PASSED | 2023-07-25 16:04
Yes | 35 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.141% | PASSED | 2023-07-25 16:04
Yes | 36 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.132% | PASSED | 2023-07-25 16:05
Yes | 37 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.142% | PASSED | 2023-07-25 16:05
Yes | 38 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.145% | PASSED | 2023-07-25 16:06
Yes | 39 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.141% | PASSED | 2023-07-25 16:06
Yes | 40 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.133% | PASSED | 2023-07-25 16:06
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -0.004% | 0% | 0% | 0.1%
0 | 1 | -0.04% | -0.036% | 0% | 0.1%
0 | 2 | 0.013% | 0.053% | -0.01% | 0.1%
0 | 3 | 0.009% | -0.004% | -0.006% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -0.004% | 0% | 0% | 0.1%
1 | 1 | -0.004% | 0% | 0% | 0.1%
1 | 2 | 0.004% | 0.008% | -0.001% | 0.1%
1 | 3 | 0% | -0.004% | 0% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | 0.018% | 0% | 0% | 0.1%
2 | 1 | 0.049% | 0.031% | 0% | 0.1%
2 | 2 | 0.009% | -0.04% | 0.025% | 0.1%
2 | 3 | 0.013% | 0.004% | 0.024% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.018% | 0% | 0% | 0.1%
3 | 1 | 0.022% | 0.004% | 0% | 0.1%
3 | 2 | 0% | -0.022% | 0.013% | 0.1%
3 | 3 | 0.022% | 0.022% | 0.015% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -0.004% | 0% | 0% | 0.1%
4 | 1 | 0.044% | 0.048% | 0% | 0.1%
4 | 2 | 0% | -0.044% | 0.013% | 0.1%
4 | 3 | 0.004% | 0.004% | 0.016% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | 0.013% | 0% | 0% | 0.1%
5 | 1 | 0.044% | 0.031% | 0% | 0.1%
5 | 2 | 0.004% | -0.04% | 0.02% | 0.1%
5 | 3 | 0.049% | 0.045% | 0.032% | 0.1%
Deviations of %error on test id #6. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | 0.009% | 0% | 0% | 0.1%
6 | 1 | 0.031% | 0.022% | 0% | 0.1%
6 | 2 | 0.009% | -0.022% | 0.016% | 0.1%
6 | 3 | 0.027% | 0.018% | 0.022% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.049% | 0% | 0% | 0.1%
7 | 1 | 0.004% | -0.045% | 0% | 0.1%
7 | 2 | 0.058% | 0.054% | 0.037% | 0.1%
7 | 3 | 0.004% | -0.054% | 0.022% | 0.1%
Deviations of %error on test id #8. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | 0.013% | 0% | 0% | 0.1%
8 | 1 | 0.071% | 0.058% | 0% | 0.1%
8 | 2 | 0.013% | -0.058% | 0.032% | 0.1%
8 | 3 | 0.044% | 0.031% | 0.043% | 0.1%
Deviations of %error on test id #9. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | 0.013% | 0% | 0% | 0.1%
9 | 1 | 0.053% | 0.04% | 0% | 0.1%
9 | 2 | 0.04% | -0.013% | 0.035% | 0.1%
9 | 3 | 0.049% | 0.009% | 0.047% | 0.1%
Deviations of %error on test id #10. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | 0.049% | 0% | 0% | 0.1%
10 | 1 | 0.049% | 0% | 0% | 0.1%
10 | 2 | 0.04% | -0.009% | 0.046% | 0.1%
10 | 3 | 0.053% | 0.013% | 0.047% | 0.1%
Deviations of %error on test id #11. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | 0.093% | 0% | 0% | 0.1%
11 | 1 | 0.04% | -0.053% | 0% | 0.1%
11 | 2 | 0.089% | 0.049% | 0.074% | 0.1%
11 | 3 | 0.04% | -0.049% | 0.056% | 0.1%
Deviations of %error on test id #12. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | 0.093% | 0% | 0% | 0.1%
12 | 1 | 0.04% | -0.053% | 0% | 0.1%
12 | 2 | 0.058% | 0.018% | 0.064% | 0.1%
12 | 3 | 0.089% | 0.031% | 0.062% | 0.1%
Deviations of %error on test id #13. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | 0.093% | 0% | 0% | 0.1%
13 | 1 | 0.053% | -0.04% | 0% | 0.1%
13 | 2 | 0.058% | 0.005% | 0.068% | 0.1%
13 | 3 | 0.085% | 0.027% | 0.065% | 0.1%
Deviations of %error on test id #14. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | 0.089% | 0% | 0% | 0.1%
14 | 1 | 0.036% | -0.053% | 0% | 0.1%
14 | 2 | 0.085% | 0.049% | 0.07% | 0.1%
14 | 3 | 0.053% | -0.032% | 0.058% | 0.1%
Deviations of %error on test id #15. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | 0.076% | 0% | 0% | 0.1%
15 | 1 | 0.044% | -0.032% | 0% | 0.1%
15 | 2 | 0.071% | 0.027% | 0.064% | 0.1%
15 | 3 | 0.04% | -0.031% | 0.052% | 0.1%
Deviations of %error on test id #16. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | 0.071% | 0% | 0% | 0.1%
16 | 1 | 0.071% | 0% | 0% | 0.1%
16 | 2 | 0.076% | 0.005% | 0.073% | 0.1%
16 | 3 | 0.062% | -0.014% | 0.07% | 0.1%
16 | 4 | 0.071% | 0.009% | 0.07% | 0.1%
Deviations of %error on test id #17. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | 0.093% | 0% | 0% | 0.1%
17 | 1 | 0.107% | 0.014% | 0% | 0.1%
17 | 2 | 0.053% | -0.054% | 0.084% | 0.1%
17 | 3 | 0.093% | 0.04% | 0.084% | 0.1%
Deviations of %error on test id #18. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | 0.04% | 0% | 0% | 0.1%
18 | 1 | 0.134% | 0.094% | 0% | 0.1%
18 | 2 | 0.04% | -0.094% | 0.071% | 0.1%
18 | 3 | 0.125% | 0.085% | 0.1% | 0.1%
Deviations of %error on test id #19. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | 0.08% | 0% | 0% | 0.1%
19 | 1 | 0.053% | -0.027% | 0% | 0.1%
19 | 2 | 0.12% | 0.067% | 0.084% | 0.1%
19 | 3 | 0.067% | -0.053% | 0.08% | 0.1%
Deviations of %error on test id #20. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | 0.08% | 0% | 0% | 0.1%
20 | 1 | 0.076% | -0.004% | 0% | 0.1%
20 | 2 | 0.089% | 0.013% | 0.082% | 0.1%
20 | 3 | 0.076% | -0.013% | 0.08% | 0.1%
Deviations of %error on test id #21. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | 0.102% | 0% | 0% | 0.1%
21 | 1 | 0.08% | -0.022% | 0% | 0.1%
21 | 2 | 0.093% | 0.013% | 0.092% | 0.1%
21 | 3 | 0.076% | -0.017% | 0.083% | 0.1%
Deviations of %error on test id #22. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | 0.12% | 0% | 0% | 0.1%
22 | 1 | 0.067% | -0.053% | 0% | 0.1%
22 | 2 | 0.08% | 0.013% | 0.089% | 0.1%
22 | 3 | 0.107% | 0.027% | 0.085% | 0.1%
Deviations of %error on test id #23. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | 0.085% | 0% | 0% | 0.1%
23 | 1 | 0.107% | 0.022% | 0% | 0.1%
23 | 2 | 0.085% | -0.022% | 0.092% | 0.1%
23 | 3 | 0.076% | -0.009% | 0.089% | 0.1%
Deviations of %error on test id #24. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | 0.093% | 0% | 0% | 0.1%
24 | 1 | 0.093% | 0% | 0% | 0.1%
24 | 2 | 0.085% | -0.008% | 0.09% | 0.1%
24 | 3 | 0.134% | 0.049% | 0.104% | 0.1%
Deviations of %error on test id #25. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | 0.107% | 0% | 0% | 0.1%
25 | 1 | 0.085% | -0.022% | 0% | 0.1%
25 | 2 | 0.142% | 0.057% | 0.111% | 0.1%
25 | 3 | 0.107% | -0.035% | 0.111% | 0.1%
Deviations of %error on test id #26. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | 0.138% | 0% | 0% | 0.1%
26 | 1 | 0.138% | 0% | 0% | 0.1%
26 | 2 | 0.093% | -0.045% | 0.123% | 0.1%
26 | 3 | 0.129% | 0.036% | 0.12% | 0.1%
Deviations of %error on test id #27. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | 0.116% | 0% | 0% | 0.1%
27 | 1 | 0.089% | -0.027% | 0% | 0.1%
27 | 2 | 0.134% | 0.045% | 0.113% | 0.1%
27 | 3 | 0.111% | -0.023% | 0.111% | 0.1%
Deviations of %error on test id #28. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | 0.129% | 0% | 0% | 0.1%
28 | 1 | 0.12% | -0.009% | 0% | 0.1%
28 | 2 | 0.085% | -0.035% | 0.111% | 0.1%
28 | 3 | 0.089% | 0.004% | 0.098% | 0.1%
Deviations of %error on test id #29. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | 0.134% | 0% | 0% | 0.1%
29 | 1 | 0.102% | -0.032% | 0% | 0.1%
29 | 2 | 0.151% | 0.049% | 0.129% | 0.1%
29 | 3 | 0.142% | -0.009% | 0.132% | 0.1%
Deviations of %error on test id #30. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | 0.147% | 0% | 0% | 0.1%
30 | 1 | 0.098% | -0.049% | 0% | 0.1%
30 | 2 | 0.147% | 0.049% | 0.131% | 0.1%
30 | 3 | 0.147% | 0% | 0.131% | 0.1%
Deviations of %error on test id #31. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | 0.12% | 0% | 0% | 0.1%
31 | 1 | 0.147% | 0.027% | 0% | 0.1%
31 | 2 | 0.116% | -0.031% | 0.128% | 0.1%
31 | 3 | 0.116% | 0% | 0.126% | 0.1%
Deviations of %error on test id #32. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | 0.12% | 0% | 0% | 0.1%
32 | 1 | 0.142% | 0.022% | 0% | 0.1%
32 | 2 | 0.111% | -0.031% | 0.124% | 0.1%
32 | 3 | 0.12% | 0.009% | 0.124% | 0.1%
Deviations of %error on test id #33. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | 0.12% | 0% | 0% | 0.1%
33 | 1 | 0.156% | 0.036% | 0% | 0.1%
33 | 2 | 0.156% | 0% | 0.144% | 0.1%
33 | 3 | 0.125% | -0.031% | 0.146% | 0.1%
Deviations of %error on test id #34. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | 0.156% | 0% | 0% | 0.1%
34 | 1 | 0.12% | -0.036% | 0% | 0.1%
34 | 2 | 0.156% | 0.036% | 0.144% | 0.1%
34 | 3 | 0.12% | -0.036% | 0.132% | 0.1%
Deviations of %error on test id #35. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | 0.129% | 0% | 0% | 0.1%
35 | 1 | 0.147% | 0.018% | 0% | 0.1%
35 | 2 | 0.156% | 0.009% | 0.144% | 0.1%
35 | 3 | 0.12% | -0.036% | 0.141% | 0.1%
Deviations of %error on test id #36. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | 0.147% | 0% | 0% | 0.1%
36 | 1 | 0.147% | 0% | 0% | 0.1%
36 | 2 | 0.098% | -0.049% | 0.131% | 0.1%
36 | 3 | 0.151% | 0.053% | 0.132% | 0.1%
Deviations of %error on test id #37. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | 0.142% | 0% | 0% | 0.1%
37 | 1 | 0.147% | 0.005% | 0% | 0.1%
37 | 2 | 0.147% | 0% | 0.145% | 0.1%
37 | 3 | 0.138% | -0.009% | 0.144% | 0.1%
37 | 4 | 0.142% | 0.004% | 0.142% | 0.1%
Deviations of %error on test id #38. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | 0.116% | 0% | 0% | 0.1%
38 | 1 | 0.138% | 0.022% | 0% | 0.1%
38 | 2 | 0.147% | 0.009% | 0.134% | 0.1%
38 | 3 | 0.151% | 0.004% | 0.145% | 0.1%
Deviations of %error on test id #39. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | 0.147% | 0% | 0% | 0.1%
39 | 1 | 0.129% | -0.018% | 0% | 0.1%
39 | 2 | 0.147% | 0.018% | 0.141% | 0.1%
39 | 3 | 0.147% | 0% | 0.141% | 0.1%
Deviations of %error on test id #40. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | 0.138% | 0% | 0% | 0.1%
40 | 1 | 0.147% | 0.009% | 0% | 0.1%
40 | 2 | 0.102% | -0.045% | 0.129% | 0.1%
40 | 3 | 0.151% | 0.049% | 0.133% | 0.1%

## 8. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:20:53 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.25% | PASSED | 2023-07-25 16:07
Yes | 1 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.214% | PASSED | 2023-07-25 16:08
Yes | 2 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.396% | PASSED | 2023-07-25 16:08
Yes | 3 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.306% | PASSED | 2023-07-25 16:08
Yes | 4 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.062% | PASSED | 2023-07-25 16:09
Yes | 5 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.232% | PASSED | 2023-07-25 16:10
Yes | 6 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.196% | PASSED | 2023-07-25 16:10
Yes | 7 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.035% | PASSED | 2023-07-25 16:11
Yes | 8 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.056% | PASSED | 2023-07-25 16:11
Yes | 9 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.05% | PASSED | 2023-07-25 16:12
Yes | 10 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.053% | PASSED | 2023-07-25 16:12
Yes | 11 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.024% | PASSED | 2023-07-25 16:13
Yes | 12 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.253% | PASSED | 2023-07-25 16:14
Yes | 13 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.024% | PASSED | 2023-07-25 16:14
Yes | 14 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.107% | PASSED | 2023-07-25 16:15
Yes | 15 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.009% | PASSED | 2023-07-25 16:16
Yes | 16 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.006% | PASSED | 2023-07-25 16:16
Yes | 17 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.027% | PASSED | 2023-07-25 16:17
Yes | 18 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.33% | PASSED | 2023-07-25 17:24
Yes | 19 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.042% | PASSED | 2023-07-25 16:18
Yes | 20 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.246% | PASSED | 2023-07-25 17:24
Yes | 21 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.035% | PASSED | 2023-07-25 16:19
Yes | 22 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.003% | PASSED | 2023-07-25 16:20
Yes | 23 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0% | PASSED | 2023-07-25 16:20
Yes | 24 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.018% | PASSED | 2023-07-25 16:21
Yes | 25 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.021% | PASSED | 2023-07-25 16:22
Yes | 26 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.024% | PASSED | 2023-07-25 17:25
Yes | 27 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.018% | PASSED | 2023-07-25 17:25
Yes | 28 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.039% | PASSED | 2023-07-25 16:25
Yes | 29 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.05% | PASSED | 2023-07-25 16:26
Yes | 30 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.024% | PASSED | 2023-07-25 17:26
Yes | 31 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.036% | PASSED | 2023-07-25 17:27
Yes | 32 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.047% | PASSED | 2023-07-25 17:27
Yes | 33 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.128% | PASSED | 2023-07-25 17:28
Yes | 34 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.229% | PASSED | 2023-07-25 17:28
Yes | 35 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.009% | PASSED | 2023-07-25 17:29
Yes | 36 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.125% | PASSED | 2023-07-25 17:29
Yes | 37 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.092% | PASSED | 2023-07-25 17:30
Yes | 38 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.077% | PASSED | 2023-07-25 17:34
Yes | 39 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.039% | PASSED | 2023-07-25 17:31
Yes | 40 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.089% | PASSED | 2023-07-25 17:31
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -35.75% | 0% | 0% | 0.1%
0 | 1 | -26.45% | 9.296% | 0% | 0.1%
0 | 2 | -20.22% | 6.233% | -27.47% | 0.1%
0 | 3 | -11.48% | 8.738% | -19.38% | 0.1%
0 | 4 | -4.83% | 6.65% | -12.18% | 0.1%
0 | 5 | 0.321% | 5.151% | -5.33% | 0.1%
0 | 6 | 0.276% | -0.045% | -1.411% | 0.1%
0 | 7 | 0.232% | -0.044% | 0.276% | 0.1%
0 | 8 | 0.241% | 0.009% | 0.25% | 0.1%
Deviations of %error on test id #1. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -29.09% | 0% | 0% | 0.1%
1 | 1 | -27.49% | 1.6% | 0% | 0.1%
1 | 2 | -15.35% | 12.14% | -23.98% | 0.1%
1 | 3 | -11.81% | 3.543% | -18.21% | 0.1%
1 | 4 | -2.114% | 9.692% | -9.756% | 0.1%
1 | 5 | -1.203% | 0.911% | -5.041% | 0.1%
1 | 6 | 0.241% | 1.444% | -1.025% | 0.1%
1 | 7 | 0.134% | -0.107% | -0.276% | 0.1%
1 | 8 | 0.089% | -0.045% | 0.155% | 0.1%
1 | 9 | 0.089% | 0% | 0.104% | 0.1%
1 | 10 | 0.339% | 0.25% | 0.172% | 0.1%
1 | 11 | 0.366% | 0.027% | 0.265% | 0.1%
1 | 12 | 0.285% | -0.081% | 0.33% | 0.1%
1 | 13 | 0.563% | 0.278% | 0.405% | 0.1%
1 | 14 | 0.834% | 0.271% | 0.561% | 0.1%
1 | 15 | 0.509% | -0.325% | 0.635% | 0.1%
1 | 16 | 0.473% | -0.036% | 0.605% | 0.1%
1 | 17 | 0.321% | -0.152% | 0.434% | 0.1%
1 | 18 | 0.223% | -0.098% | 0.339% | 0.1%
1 | 19 | 0.098% | -0.125% | 0.214% | 0.1%
Deviations of %error on test id #2. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -30.4% | 0% | 0% | 0.1%
2 | 1 | -19.11% | 11.29% | 0% | 0.1%
2 | 2 | -16.91% | 2.205% | -22.14% | 0.1%
2 | 3 | -3.45% | 13.46% | -13.16% | 0.1%
2 | 4 | -2.623% | 0.827% | -7.66% | 0.1%
2 | 5 | 0.357% | 2.98% | -1.905% | 0.1%
2 | 6 | 0.321% | -0.036% | -0.648% | 0.1%
2 | 7 | 0.411% | 0.09% | 0.363% | 0.1%
2 | 8 | 0.455% | 0.044% | 0.396% | 0.1%
Deviations of %error on test id #3. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -29.34% | 0% | 0% | 0.1%
3 | 1 | -17.99% | 11.35% | 0% | 0.1%
3 | 2 | -15.18% | 2.807% | -20.84% | 0.1%
3 | 3 | -2.361% | 12.82% | -11.85% | 0.1%
3 | 4 | -1.815% | 0.546% | -6.453% | 0.1%
3 | 5 | 0.205% | 2.02% | -1.324% | 0.1%
3 | 6 | 0.241% | 0.036% | -0.456% | 0.1%
3 | 7 | 0.33% | 0.089% | 0.259% | 0.1%
3 | 8 | 0.348% | 0.018% | 0.306% | 0.1%
Deviations of %error on test id #4. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -19.97% | 0% | 0% | 0.1%
4 | 1 | -11.19% | 8.775% | 0% | 0.1%
4 | 2 | -6.383% | 4.81% | -12.52% | 0.1%
4 | 3 | -0.724% | 5.659% | -6.1% | 0.1%
4 | 4 | 0.214% | 0.938% | -2.298% | 0.1%
4 | 5 | 0.258% | 0.044% | -0.084% | 0.1%
4 | 6 | 0.42% | 0.162% | 0.297% | 0.1%
4 | 7 | 0.25% | -0.17% | 0.309% | 0.1%
4 | 8 | 0.285% | 0.035% | 0.318% | 0.1%
4 | 9 | 0.473% | 0.188% | 0.336% | 0.1%
4 | 10 | 0.393% | -0.08% | 0.384% | 0.1%
4 | 11 | 0.241% | -0.152% | 0.369% | 0.1%
4 | 12 | 0.25% | 0.009% | 0.295% | 0.1%
4 | 13 | 0.134% | -0.116% | 0.208% | 0.1%
4 | 14 | 0.071% | -0.063% | 0.152% | 0.1%
4 | 15 | 0.107% | 0.036% | 0.104% | 0.1%
4 | 16 | 0.009% | -0.098% | 0.062% | 0.1%
Deviations of %error on test id #5. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -11.19% | 0% | 0% | 0.1%
5 | 1 | -4.919% | 6.274% | 0% | 0.1%
5 | 2 | -2.233% | 2.686% | -6.115% | 0.1%
5 | 3 | 0.473% | 2.706% | -2.226% | 0.1%
5 | 4 | 0.393% | -0.08% | -0.456% | 0.1%
5 | 5 | 0.321% | -0.072% | 0.396% | 0.1%
5 | 6 | 0.169% | -0.152% | 0.294% | 0.1%
5 | 7 | 0.151% | -0.018% | 0.214% | 0.1%
5 | 8 | 0.116% | -0.035% | 0.145% | 0.1%
5 | 9 | 0.357% | 0.241% | 0.208% | 0.1%
5 | 10 | 0.285% | -0.072% | 0.253% | 0.1%
5 | 11 | 0.241% | -0.044% | 0.294% | 0.1%
5 | 12 | 0.169% | -0.072% | 0.232% | 0.1%
Deviations of %error on test id #6. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | -14% | 0% | 0% | 0.1%
6 | 1 | -10.79% | 3.212% | 0% | 0.1%
6 | 2 | -4.402% | 6.383% | -9.728% | 0.1%
6 | 3 | -1.781% | 2.621% | -5.656% | 0.1%
6 | 4 | 0.276% | 2.057% | -1.969% | 0.1%
6 | 5 | 0.187% | -0.089% | -0.439% | 0.1%
6 | 6 | 0.151% | -0.036% | 0.205% | 0.1%
6 | 7 | 0.25% | 0.099% | 0.196% | 0.1%
Deviations of %error on test id #7. Total 18 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | -18.8% | 0% | 0% | 0.1%
7 | 1 | -15.74% | 3.059% | 0% | 0.1%
7 | 2 | -4.919% | 10.82% | -13.15% | 0.1%
7 | 3 | -2.9% | 2.019% | -7.854% | 0.1%
7 | 4 | 0.178% | 3.078% | -2.547% | 0.1%
7 | 5 | 0.196% | 0.018% | -0.842% | 0.1%
7 | 6 | 0.402% | 0.206% | 0.259% | 0.1%
7 | 7 | 0.339% | -0.063% | 0.312% | 0.1%
7 | 8 | 0.196% | -0.143% | 0.312% | 0.1%
7 | 9 | 0.312% | 0.116% | 0.282% | 0.1%
7 | 10 | 0.393% | 0.081% | 0.3% | 0.1%
7 | 11 | 0.267% | -0.126% | 0.324% | 0.1%
7 | 12 | 0.187% | -0.08% | 0.282% | 0.1%
7 | 13 | 0.169% | -0.018% | 0.208% | 0.1%
7 | 14 | 0.053% | -0.116% | 0.136% | 0.1%
7 | 15 | 0.053% | 0% | 0.092% | 0.1%
7 | 16 | 0.009% | -0.044% | 0.038% | 0.1%
7 | 17 | 0.044% | 0.035% | 0.035% | 0.1%
Deviations of %error on test id #8. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | -11.23% | 0% | 0% | 0.1%
8 | 1 | -4.1% | 7.128% | 0% | 0.1%
8 | 2 | -1.506% | 2.594% | -5.611% | 0.1%
8 | 3 | 0.018% | 1.524% | -1.863% | 0.1%
8 | 4 | 0.062% | 0.044% | -0.475% | 0.1%
8 | 5 | 0.062% | 0% | 0.047% | 0.1%
8 | 6 | 0.267% | 0.205% | 0.13% | 0.1%
8 | 7 | 0.384% | 0.117% | 0.238% | 0.1%
8 | 8 | 0.563% | 0.179% | 0.405% | 0.1%
8 | 9 | 0.482% | -0.081% | 0.476% | 0.1%
8 | 10 | 0.366% | -0.116% | 0.47% | 0.1%
8 | 11 | 0.223% | -0.143% | 0.357% | 0.1%
8 | 12 | 0.142% | -0.081% | 0.244% | 0.1%
8 | 13 | 0.018% | -0.124% | 0.128% | 0.1%
8 | 14 | 0.08% | 0.062% | 0.08% | 0.1%
8 | 15 | 0.053% | -0.027% | 0.05% | 0.1%
8 | 16 | 0.036% | -0.017% | 0.056% | 0.1%
Deviations of %error on test id #9. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | -12.28% | 0% | 0% | 0.1%
9 | 1 | -5.087% | 7.193% | 0% | 0.1%
9 | 2 | -2.799% | 2.288% | -6.722% | 0.1%
9 | 3 | 0.062% | 2.861% | -2.608% | 0.1%
9 | 4 | 0.044% | -0.018% | -0.898% | 0.1%
9 | 5 | 0.33% | 0.286% | 0.145% | 0.1%
9 | 6 | 0.303% | -0.027% | 0.226% | 0.1%
9 | 7 | 0.187% | -0.116% | 0.273% | 0.1%
9 | 8 | 0.393% | 0.206% | 0.294% | 0.1%
9 | 9 | 0.348% | -0.045% | 0.309% | 0.1%
9 | 10 | 0.285% | -0.063% | 0.342% | 0.1%
9 | 11 | 0.178% | -0.107% | 0.27% | 0.1%
9 | 12 | 0.187% | 0.009% | 0.217% | 0.1%
9 | 13 | 0.053% | -0.134% | 0.139% | 0.1%
9 | 14 | 0.08% | 0.027% | 0.107% | 0.1%
9 | 15 | -0.009% | -0.089% | 0.041% | 0.1%
9 | 16 | 0.08% | 0.089% | 0.05% | 0.1%
Deviations of %error on test id #10. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | -4.911% | 0% | 0% | 0.1%
10 | 1 | -2.25% | 2.661% | 0% | 0.1%
10 | 2 | 0.169% | 2.419% | -2.331% | 0.1%
10 | 3 | 0.134% | -0.035% | -0.649% | 0.1%
10 | 4 | 0.178% | 0.044% | 0.16% | 0.1%
10 | 5 | 0.303% | 0.125% | 0.205% | 0.1%
10 | 6 | 0.25% | -0.053% | 0.244% | 0.1%
10 | 7 | 0.348% | 0.098% | 0.3% | 0.1%
10 | 8 | 0.455% | 0.107% | 0.351% | 0.1%
10 | 9 | 0.285% | -0.17% | 0.363% | 0.1%
10 | 10 | 0.25% | -0.035% | 0.33% | 0.1%
10 | 11 | 0.125% | -0.125% | 0.22% | 0.1%
10 | 12 | 0.098% | -0.027% | 0.158% | 0.1%
10 | 13 | 0.009% | -0.089% | 0.077% | 0.1%
10 | 14 | 0.053% | 0.044% | 0.053% | 0.1%
Deviations of %error on test id #11. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | 0.205% | 0% | 0% | 0.1%
11 | 1 | 0.169% | -0.036% | 0% | 0.1%
11 | 2 | 0.187% | 0.018% | 0.187% | 0.1%
11 | 3 | 0.366% | 0.179% | 0.241% | 0.1%
11 | 4 | 0.366% | 0% | 0.306% | 0.1%
11 | 5 | 0.509% | 0.143% | 0.414% | 0.1%
11 | 6 | 0.393% | -0.116% | 0.423% | 0.1%
11 | 7 | 0.241% | -0.152% | 0.381% | 0.1%
11 | 8 | 0.151% | -0.09% | 0.262% | 0.1%
11 | 9 | 0.009% | -0.142% | 0.134% | 0.1%
11 | 10 | 0.036% | 0.027% | 0.065% | 0.1%
11 | 11 | 0% | -0.036% | 0.015% | 0.1%
11 | 12 | 0.036% | 0.036% | 0.024% | 0.1%
Deviations of %error on test id #12. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | 0.098% | 0% | 0% | 0.1%
12 | 1 | 0.312% | 0.214% | 0% | 0.1%
12 | 2 | 0.267% | -0.045% | 0.226% | 0.1%
12 | 3 | 0.241% | -0.026% | 0.273% | 0.1%
12 | 4 | 0.25% | 0.009% | 0.253% | 0.1%
Deviations of %error on test id #13. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | 0.116% | 0% | 0% | 0.1%
13 | 1 | -0.018% | -0.134% | 0% | 0.1%
13 | 2 | 0% | 0.018% | 0.033% | 0.1%
13 | 3 | 0.027% | 0.027% | 0.003% | 0.1%
13 | 4 | 0.044% | 0.017% | 0.024% | 0.1%
Deviations of %error on test id #14. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | 0.25% | 0% | 0% | 0.1%
14 | 1 | 0.178% | -0.072% | 0% | 0.1%
14 | 2 | 0.107% | -0.071% | 0.178% | 0.1%
14 | 3 | 0.036% | -0.071% | 0.107% | 0.1%
Deviations of %error on test id #15. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | 0.053% | 0% | 0% | 0.1%
15 | 1 | 0% | -0.053% | 0% | 0.1%
15 | 2 | 0.009% | 0.009% | 0.021% | 0.1%
15 | 3 | 0.018% | 0.009% | 0.009% | 0.1%
Deviations of %error on test id #16. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | 0.071% | 0% | 0% | 0.1%
16 | 1 | 0.018% | -0.053% | 0% | 0.1%
16 | 2 | -0.027% | -0.045% | 0.021% | 0.1%
16 | 3 | 0.036% | 0.063% | 0.009% | 0.1%
16 | 4 | -0.027% | -0.063% | -0.006% | 0.1%
Deviations of %error on test id #17. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | 0.232% | 0% | 0% | 0.1%
17 | 1 | 0.116% | -0.116% | 0% | 0.1%
17 | 2 | 0.062% | -0.054% | 0.137% | 0.1%
17 | 3 | 0% | -0.062% | 0.059% | 0.1%
17 | 4 | 0.018% | 0.018% | 0.027% | 0.1%
Deviations of %error on test id #18. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | -21.63% | 0% | 0% | 0.1%
18 | 1 | -10.28% | 11.35% | 0% | 0.1%
18 | 2 | -6.507% | 3.772% | -12.8% | 0.1%
18 | 3 | 0.42% | 6.927% | -5.455% | 0.1%
18 | 4 | 0.375% | -0.045% | -1.904% | 0.1%
18 | 5 | 0.294% | -0.081% | 0.363% | 0.1%
18 | 6 | 0.321% | 0.027% | 0.33% | 0.1%
Deviations of %error on test id #19. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | 0.053% | 0% | 0% | 0.1%
19 | 1 | 0.053% | 0% | 0% | 0.1%
19 | 2 | 0.036% | -0.017% | 0.047% | 0.1%
19 | 3 | 0.036% | 0% | 0.042% | 0.1%
Deviations of %error on test id #20. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | -10.54% | 0% | 0% | 0.1%
20 | 1 | -5.231% | 5.305% | 0% | 0.1%
20 | 2 | 0.267% | 5.498% | -5.167% | 0.1%
20 | 3 | 0.321% | 0.054% | -1.548% | 0.1%
20 | 4 | 0.258% | -0.063% | 0.282% | 0.1%
20 | 5 | 0.16% | -0.098% | 0.246% | 0.1%
Deviations of %error on test id #21. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | 0.009% | 0% | 0% | 0.1%
21 | 1 | 0.044% | 0.035% | 0% | 0.1%
21 | 2 | 0.044% | 0% | 0.032% | 0.1%
21 | 3 | 0.018% | -0.026% | 0.035% | 0.1%
Deviations of %error on test id #22. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | -0.009% | 0% | 0% | 0.1%
22 | 1 | 0.018% | 0.027% | 0% | 0.1%
22 | 2 | 0.009% | -0.009% | 0.006% | 0.1%
22 | 3 | -0.018% | -0.027% | 0.003% | 0.1%
Deviations of %error on test id #23. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | -0.036% | 0% | 0% | 0.1%
23 | 1 | 0.036% | 0.072% | 0% | 0.1%
23 | 2 | -0.044% | -0.08% | -0.015% | 0.1%
23 | 3 | 0.009% | 0.053% | 0% | 0.1%
Deviations of %error on test id #24. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | -0.027% | 0% | 0% | 0.1%
24 | 1 | -0.036% | -0.009% | 0% | 0.1%
24 | 2 | -0.027% | 0.009% | -0.03% | 0.1%
24 | 3 | 0.009% | 0.036% | -0.018% | 0.1%
Deviations of %error on test id #25. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | 0.018% | 0% | 0% | 0.1%
25 | 1 | -0.027% | -0.045% | 0% | 0.1%
25 | 2 | -0.053% | -0.026% | -0.021% | 0.1%
25 | 3 | 0% | 0.053% | -0.027% | 0.1%
25 | 4 | -0.009% | -0.009% | -0.021% | 0.1%
Deviations of %error on test id #26. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | -19.72% | 0% | 0% | 0.1%
26 | 1 | -13.77% | 5.951% | 0% | 0.1%
26 | 2 | -4.782% | 8.984% | -12.75% | 0.1%
26 | 3 | -0.767% | 4.015% | -6.438% | 0.1%
26 | 4 | 0.384% | 1.151% | -1.722% | 0.1%
26 | 5 | 0.303% | -0.081% | -0.027% | 0.1%
26 | 6 | 0.267% | -0.036% | 0.318% | 0.1%
26 | 7 | 0.5% | 0.233% | 0.357% | 0.1%
26 | 8 | 0.527% | 0.027% | 0.431% | 0.1%
26 | 9 | 0.482% | -0.045% | 0.503% | 0.1%
26 | 10 | 0.312% | -0.17% | 0.44% | 0.1%
26 | 11 | 0.089% | -0.223% | 0.294% | 0.1%
26 | 12 | 0.062% | -0.027% | 0.154% | 0.1%
26 | 13 | -0.08% | -0.142% | 0.024% | 0.1%
26 | 14 | 0% | 0.08% | -0.006% | 0.1%
26 | 15 | -0.027% | -0.027% | -0.036% | 0.1%
26 | 16 | -0.044% | -0.017% | -0.024% | 0.1%
Deviations of %error on test id #27. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | -21.44% | 0% | 0% | 0.1%
27 | 1 | -9.993% | 11.45% | 0% | 0.1%
27 | 2 | -2.547% | 7.446% | -11.33% | 0.1%
27 | 3 | 0.098% | 2.645% | -4.147% | 0.1%
27 | 4 | -0.027% | -0.125% | -0.825% | 0.1%
27 | 5 | 0.169% | 0.196% | 0.08% | 0.1%
27 | 6 | 0.375% | 0.206% | 0.172% | 0.1%
27 | 7 | 0.545% | 0.17% | 0.363% | 0.1%
27 | 8 | 0.42% | -0.125% | 0.447% | 0.1%
27 | 9 | 0.339% | -0.081% | 0.435% | 0.1%
27 | 10 | 0.107% | -0.232% | 0.289% | 0.1%
27 | 11 | 0.009% | -0.098% | 0.152% | 0.1%
27 | 12 | -0.018% | -0.027% | 0.033% | 0.1%
27 | 13 | -0.044% | -0.026% | -0.018% | 0.1%
Deviations of %error on test id #28. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | -0.08% | 0% | 0% | 0.1%
28 | 1 | -0.027% | 0.053% | 0% | 0.1%
28 | 2 | -0.027% | 0% | -0.045% | 0.1%
28 | 3 | -0.036% | -0.009% | -0.03% | 0.1%
28 | 4 | -0.053% | -0.017% | -0.039% | 0.1%
Deviations of %error on test id #29. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | -0.027% | 0% | 0% | 0.1%
29 | 1 | -0.053% | -0.026% | 0% | 0.1%
29 | 2 | -0.071% | -0.018% | -0.05% | 0.1%
29 | 3 | -0.027% | 0.044% | -0.05% | 0.1%
Deviations of %error on test id #30. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | -13.03% | 0% | 0% | 0.1%
30 | 1 | -6.894% | 6.132% | 0% | 0.1%
30 | 2 | 0.348% | 7.242% | -6.524% | 0.1%
30 | 3 | 0.285% | -0.063% | -2.087% | 0.1%
30 | 4 | 0.321% | 0.036% | 0.318% | 0.1%
30 | 5 | 0.5% | 0.179% | 0.369% | 0.1%
30 | 6 | 0.518% | 0.018% | 0.446% | 0.1%
30 | 7 | 0.545% | 0.027% | 0.521% | 0.1%
30 | 8 | 0.348% | -0.197% | 0.47% | 0.1%
30 | 9 | 0.151% | -0.197% | 0.348% | 0.1%
30 | 10 | 0.018% | -0.133% | 0.172% | 0.1%
30 | 11 | -0.027% | -0.045% | 0.047% | 0.1%
30 | 12 | -0.018% | 0.009% | -0.009% | 0.1%
30 | 13 | -0.027% | -0.009% | -0.024% | 0.1%
Deviations of %error on test id #31. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | -0.469% | 0% | 0% | 0.1%
31 | 1 | 0.232% | 0.701% | 0% | 0.1%
31 | 2 | 0.098% | -0.134% | -0.046% | 0.1%
31 | 3 | 0.178% | 0.08% | 0.169% | 0.1%
31 | 4 | 0.42% | 0.242% | 0.232% | 0.1%
31 | 5 | 0.554% | 0.134% | 0.384% | 0.1%
31 | 6 | 0.33% | -0.224% | 0.435% | 0.1%
31 | 7 | 0.241% | -0.089% | 0.375% | 0.1%
31 | 8 | 0.053% | -0.188% | 0.208% | 0.1%
31 | 9 | -0.018% | -0.071% | 0.092% | 0.1%
31 | 10 | -0.053% | -0.035% | -0.006% | 0.1%
31 | 11 | -0.036% | 0.017% | -0.036% | 0.1%
Deviations of %error on test id #32. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | -9.928% | 0% | 0% | 0.1%
32 | 1 | -2.825% | 7.103% | 0% | 0.1%
32 | 2 | 0.205% | 3.03% | -4.183% | 0.1%
32 | 3 | 0.089% | -0.116% | -0.844% | 0.1%
32 | 4 | 0.187% | 0.098% | 0.16% | 0.1%
32 | 5 | 0.339% | 0.152% | 0.205% | 0.1%
32 | 6 | 0.527% | 0.188% | 0.351% | 0.1%
32 | 7 | 0.411% | -0.116% | 0.426% | 0.1%
32 | 8 | 0.312% | -0.099% | 0.417% | 0.1%
32 | 9 | 0.036% | -0.276% | 0.253% | 0.1%
32 | 10 | 0.018% | -0.018% | 0.122% | 0.1%
32 | 11 | -0.053% | -0.071% | 0% | 0.1%
32 | 12 | -0.053% | 0% | -0.029% | 0.1%
32 | 13 | -0.036% | 0.017% | -0.047% | 0.1%
Deviations of %error on test id #33. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | -8.103% | 0% | 0% | 0.1%
33 | 1 | 0.581% | 8.684% | 0% | 0.1%
33 | 2 | 0.473% | -0.108% | -2.35% | 0.1%
33 | 3 | 0.321% | -0.152% | 0.458% | 0.1%
33 | 4 | 0.08% | -0.241% | 0.291% | 0.1%
33 | 5 | 0.142% | 0.062% | 0.181% | 0.1%
33 | 6 | 0.312% | 0.17% | 0.178% | 0.1%
33 | 7 | 0.214% | -0.098% | 0.223% | 0.1%
33 | 8 | 0.116% | -0.098% | 0.214% | 0.1%
33 | 9 | 0.053% | -0.063% | 0.128% | 0.1%
Deviations of %error on test id #34. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | -7.392% | 0% | 0% | 0.1%
34 | 1 | -2.833% | 4.559% | 0% | 0.1%
34 | 2 | 0.107% | 2.94% | -3.373% | 0.1%
34 | 3 | 0.25% | 0.143% | -0.825% | 0.1%
34 | 4 | 0.258% | 0.008% | 0.205% | 0.1%
34 | 5 | 0.196% | -0.062% | 0.235% | 0.1%
34 | 6 | 0.223% | 0.027% | 0.226% | 0.1%
34 | 7 | 0.267% | 0.044% | 0.229% | 0.1%
Deviations of %error on test id #35. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | 0.151% | 0% | 0% | 0.1%
35 | 1 | 0.339% | 0.188% | 0% | 0.1%
35 | 2 | 0.16% | -0.179% | 0.217% | 0.1%
35 | 3 | 0.169% | 0.009% | 0.223% | 0.1%
35 | 4 | 0.321% | 0.152% | 0.217% | 0.1%
35 | 5 | 0.178% | -0.143% | 0.223% | 0.1%
35 | 6 | 0.089% | -0.089% | 0.196% | 0.1%
35 | 7 | -0.009% | -0.098% | 0.086% | 0.1%
35 | 8 | -0.053% | -0.044% | 0.009% | 0.1%
Deviations of %error on test id #36. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | 0.276% | 0% | 0% | 0.1%
36 | 1 | 0.214% | -0.062% | 0% | 0.1%
36 | 2 | 0.116% | -0.098% | 0.202% | 0.1%
36 | 3 | 0.044% | -0.072% | 0.125% | 0.1%
Deviations of %error on test id #37. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | -0.089% | 0% | 0% | 0.1%
37 | 1 | 0% | 0.089% | 0% | 0.1%
37 | 2 | 0.303% | 0.303% | 0.071% | 0.1%
37 | 3 | 0.169% | -0.134% | 0.157% | 0.1%
37 | 4 | 0.151% | -0.018% | 0.208% | 0.1%
37 | 5 | 0.044% | -0.107% | 0.121% | 0.1%
37 | 6 | -0.089% | -0.133% | 0.035% | 0.1%
37 | 7 | -0.036% | 0.053% | -0.027% | 0.1%
37 | 8 | -0.106% | -0.07% | -0.077% | 0.1%
37 | 9 | -0.089% | 0.017% | -0.077% | 0.1%
37 | 10 | -0.08% | 0.009% | -0.092% | 0.1%
Deviations of %error on test id #38. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | -18.36% | 0% | 0% | 0.1%
38 | 1 | -9.667% | 8.692% | 0% | 0.1%
38 | 2 | 0.036% | 9.703% | -9.33% | 0.1%
38 | 3 | 0.178% | 0.142% | -3.151% | 0.1%
38 | 4 | 0.134% | -0.044% | 0.116% | 0.1%
38 | 5 | 0.25% | 0.116% | 0.187% | 0.1%
38 | 6 | 0.446% | 0.196% | 0.277% | 0.1%
38 | 7 | 0.527% | 0.081% | 0.408% | 0.1%
38 | 8 | 0.321% | -0.206% | 0.431% | 0.1%
38 | 9 | 0.125% | -0.196% | 0.324% | 0.1%
38 | 10 | -0.018% | -0.143% | 0.143% | 0.1%
38 | 11 | -0.062% | -0.044% | 0.015% | 0.1%
38 | 12 | -0.062% | 0% | -0.047% | 0.1%
38 | 13 | -0.106% | -0.044% | -0.077% | 0.1%
Deviations of %error on test id #39. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | 0.036% | 0% | 0% | 0.1%
39 | 1 | -0.106% | -0.142% | 0% | 0.1%
39 | 2 | -0.098% | 0.008% | -0.056% | 0.1%
39 | 3 | -0.098% | 0% | -0.101% | 0.1%
39 | 4 | -0.089% | 0.009% | -0.095% | 0.1%
39 | 5 | 0.303% | 0.392% | 0.039% | 0.1%
Deviations of %error on test id #40. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | -0.124% | 0% | 0% | 0.1%
40 | 1 | -0.08% | 0.044% | 0% | 0.1%
40 | 2 | -0.115% | -0.035% | -0.106% | 0.1%
40 | 3 | -0.071% | 0.044% | -0.089% | 0.1%

## 9. PF = 1.0
[<<]
Accuracy Test Report of 'PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:21:24 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 20 | 20 |  | ±0.1% | ±0.5% | 0.045% | PASSED | 2023-07-25 17:36
Yes | 1 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.5% | 0.037% | PASSED | 2023-07-25 17:37
Yes | 3 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.5% | 0.052% | PASSED | 2023-07-25 17:38
Yes | 4 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0.056% | PASSED | 2023-07-25 17:38
Yes | 7 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0.057% | PASSED | 2023-07-25 17:39
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 0.027% | 0% | 0% | 0.1%
0 | 1 | 0.071% | 0.044% | 0% | 0.1%
0 | 2 | 0.036% | -0.035% | 0.045% | 0.1%
0 | 3 | 0.027% | -0.009% | 0.045% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0.032% | 0% | 0% | 0.1%
1 | 1 | 0.039% | 0.007% | 0% | 0.1%
1 | 2 | 0.039% | 0% | 0.037% | 0.1%
1 | 3 | 0.032% | -0.007% | 0.037% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.036% | 0% | 0% | 0.1%
3 | 1 | 0.053% | 0.017% | 0% | 0.1%
3 | 2 | 0.058% | 0.005% | 0.049% | 0.1%
3 | 3 | 0.044% | -0.014% | 0.052% | 0.1%
Deviations of %error on test id #4. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.053% | 0% | 0% | 0.1%
4 | 1 | 0.036% | -0.017% | 0% | 0.1%
4 | 2 | 0.044% | 0.008% | 0.044% | 0.1%
4 | 3 | 0.08% | 0.036% | 0.053% | 0.1%
4 | 4 | 0.044% | -0.036% | 0.056% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.064% | 0% | 0% | 0.1%
7 | 1 | 0.057% | -0.007% | 0% | 0.1%
7 | 2 | 0.062% | 0.005% | 0.061% | 0.1%
7 | 3 | 0.052% | -0.01% | 0.057% | 0.1%

## 10. PF = 0.5L
[<<]
Accuracy Test Report of 'PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:21:28 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | -0.089% | PASSED | 2023-07-25 17:40
Yes | 1 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | 0.321% | PASSED | 2023-07-25 17:40
Yes | 3 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.033% | PASSED | 2023-07-25 17:41
Yes | 4 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.068% | PASSED | 2023-07-25 17:42
Yes | 7 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.072% | PASSED | 2023-07-25 17:43
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 19 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -32.12% | 0% | 0% | 0.1%
0 | 1 | -35.89% | -3.768% | 0% | 0.1%
0 | 2 | -27.47% | 8.415% | -31.83% | 0.1%
0 | 3 | -17.99% | 9.485% | -27.12% | 0.1%
0 | 4 | -2.816% | 15.17% | -16.09% | 0.1%
0 | 5 | -1.055% | 1.761% | -7.287% | 0.1%
0 | 6 | -0.018% | 1.037% | -1.296% | 0.1%
0 | 7 | 0.357% | 0.375% | -0.239% | 0.1%
0 | 8 | 0.232% | -0.125% | 0.19% | 0.1%
0 | 9 | 0.285% | 0.053% | 0.291% | 0.1%
0 | 10 | 0.68% | 0.395% | 0.399% | 0.1%
0 | 11 | 1.424% | 0.744% | 0.796% | 0.1%
0 | 12 | 1.224% | -0.2% | 1.109% | 0.1%
0 | 13 | 0.806% | -0.418% | 1.151% | 0.1%
0 | 14 | 0.393% | -0.413% | 0.808% | 0.1%
0 | 15 | -0.124% | -0.517% | 0.358% | 0.1%
0 | 16 | -0.071% | 0.053% | 0.066% | 0.1%
0 | 17 | -0.089% | -0.018% | -0.095% | 0.1%
0 | 18 | -0.106% | -0.017% | -0.089% | 0.1%
Deviations of %error on test id #1. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -7.716% | 0% | 0% | 0.1%
1 | 1 | 0.19% | 7.906% | 0% | 0.1%
1 | 2 | 0.387% | 0.197% | -2.38% | 0.1%
1 | 3 | 0.375% | -0.012% | 0.317% | 0.1%
1 | 4 | 0.291% | -0.084% | 0.351% | 0.1%
1 | 5 | 0.297% | 0.006% | 0.321% | 0.1%
Deviations of %error on test id #3. Total 21 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -15.63% | 0% | 0% | 0.1%
3 | 1 | -7.086% | 8.543% | 0% | 0.1%
3 | 2 | -2.707% | 4.379% | -8.474% | 0.1%
3 | 3 | 0.303% | 3.01% | -3.163% | 0.1%
3 | 4 | 0.473% | 0.17% | -0.644% | 0.1%
3 | 5 | 0.339% | -0.134% | 0.372% | 0.1%
3 | 6 | 0.303% | -0.036% | 0.372% | 0.1%
3 | 7 | 0.134% | -0.169% | 0.259% | 0.1%
3 | 8 | 0.142% | 0.008% | 0.193% | 0.1%
3 | 9 | 0.375% | 0.233% | 0.217% | 0.1%
3 | 10 | 0.267% | -0.108% | 0.261% | 0.1%
3 | 11 | 0.232% | -0.035% | 0.291% | 0.1%
3 | 12 | 0.491% | 0.259% | 0.33% | 0.1%
3 | 13 | 0.671% | 0.18% | 0.465% | 0.1%
3 | 14 | 0.5% | -0.171% | 0.554% | 0.1%
3 | 15 | 0.303% | -0.197% | 0.491% | 0.1%
3 | 16 | 0.276% | -0.027% | 0.36% | 0.1%
3 | 17 | 0.018% | -0.258% | 0.199% | 0.1%
3 | 18 | 0.018% | 0% | 0.104% | 0.1%
3 | 19 | 0.018% | 0% | 0.018% | 0.1%
3 | 20 | 0.062% | 0.044% | 0.033% | 0.1%
Deviations of %error on test id #4. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -21.49% | 0% | 0% | 0.1%
4 | 1 | -4.661% | 16.83% | 0% | 0.1%
4 | 2 | 0.036% | 4.697% | -8.704% | 0.1%
4 | 3 | 0.187% | 0.151% | -1.479% | 0.1%
4 | 4 | 0.437% | 0.25% | 0.22% | 0.1%
4 | 5 | 1.069% | 0.632% | 0.564% | 0.1%
4 | 6 | 0.825% | -0.244% | 0.777% | 0.1%
4 | 7 | 0.303% | -0.522% | 0.732% | 0.1%
4 | 8 | 0.053% | -0.25% | 0.394% | 0.1%
4 | 9 | 0.053% | 0% | 0.136% | 0.1%
4 | 10 | -0.098% | -0.151% | 0.003% | 0.1%
4 | 11 | -0.266% | -0.168% | -0.104% | 0.1%
4 | 12 | -0.124% | 0.142% | -0.163% | 0.1%
4 | 13 | 0.062% | 0.186% | -0.109% | 0.1%
4 | 14 | 0.062% | 0% | 0% | 0.1%
4 | 15 | 0.071% | 0.009% | 0.065% | 0.1%
4 | 16 | 0.071% | 0% | 0.068% | 0.1%
Deviations of %error on test id #7. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.244% | 0% | 0% | 0.1%
7 | 1 | 0.059% | -0.185% | 0% | 0.1%
7 | 2 | 0.073% | 0.014% | 0.125% | 0.1%
7 | 3 | 0.084% | 0.011% | 0.072% | 0.1%
7 | 4 | 0.059% | -0.025% | 0.072% | 0.1%

## 11. PF = 0.8C
[<<]
Accuracy Test Report of 'PF = 0.8C'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / PF = 0.8C'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:21:33 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | -0.1% | PASSED | 2023-07-25 17:44
Yes | 1 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | -0.144% | PASSED | 2023-07-25 17:44
Yes | 3 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | -0.08% | PASSED | 2023-07-25 17:44
Yes | 4 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | -0.089% | PASSED | 2023-07-25 17:45
Yes | 7 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.059% | PASSED | 2023-07-25 17:46
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -11.09% | 0% | 0% | 0.1%
0 | 1 | -9.172% | 1.923% | 0% | 0.1%
0 | 2 | -8.951% | 0.221% | -9.739% | 0.1%
0 | 3 | -5.7% | 3.251% | -7.941% | 0.1%
0 | 4 | -4.483% | 1.217% | -6.378% | 0.1%
0 | 5 | -2.917% | 1.566% | -4.367% | 0.1%
0 | 6 | -1.264% | 1.653% | -2.888% | 0.1%
0 | 7 | -0.933% | 0.331% | -1.705% | 0.1%
0 | 8 | -0.053% | 0.88% | -0.75% | 0.1%
0 | 9 | -0.124% | -0.071% | -0.37% | 0.1%
0 | 10 | -0.124% | 0% | -0.1% | 0.1%
0 | 11 | -0.053% | 0.071% | -0.1% | 0.1%
Deviations of %error on test id #1. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -5.462% | 0% | 0% | 0.1%
1 | 1 | -2.372% | 3.09% | 0% | 0.1%
1 | 2 | -0.805% | 1.567% | -2.88% | 0.1%
1 | 3 | -0.171% | 0.634% | -1.116% | 0.1%
1 | 4 | -0.136% | 0.035% | -0.371% | 0.1%
1 | 5 | -0.106% | 0.03% | -0.138% | 0.1%
1 | 6 | -0.201% | -0.095% | -0.148% | 0.1%
1 | 7 | -0.124% | 0.077% | -0.144% | 0.1%
Deviations of %error on test id #3. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -8.268% | 0% | 0% | 0.1%
3 | 1 | -6.273% | 1.995% | 0% | 0.1%
3 | 2 | -6.281% | -0.008% | -6.941% | 0.1%
3 | 3 | -2.648% | 3.633% | -5.067% | 0.1%
3 | 4 | -1.454% | 1.194% | -3.461% | 0.1%
3 | 5 | -1.514% | -0.06% | -1.872% | 0.1%
3 | 6 | -0.053% | 1.461% | -1.007% | 0.1%
3 | 7 | -0.16% | -0.107% | -0.576% | 0.1%
3 | 8 | -0.133% | 0.027% | -0.115% | 0.1%
3 | 9 | -0.089% | 0.044% | -0.127% | 0.1%
3 | 10 | -0.098% | -0.009% | -0.107% | 0.1%
3 | 11 | -0.053% | 0.045% | -0.08% | 0.1%
Deviations of %error on test id #4. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -10.71% | 0% | 0% | 0.1%
4 | 1 | -5.589% | 5.125% | 0% | 0.1%
4 | 2 | -2.992% | 2.597% | -6.432% | 0.1%
4 | 3 | -0.837% | 2.155% | -3.139% | 0.1%
4 | 4 | 0.044% | 0.881% | -1.262% | 0.1%
4 | 5 | 0.027% | -0.017% | -0.255% | 0.1%
4 | 6 | -0.08% | -0.107% | -0.003% | 0.1%
4 | 7 | -0.062% | 0.018% | -0.038% | 0.1%
4 | 8 | -0.115% | -0.053% | -0.086% | 0.1%
4 | 9 | -0.089% | 0.026% | -0.089% | 0.1%
Deviations of %error on test id #7. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | -0.064% | 0% | 0% | 0.1%
7 | 1 | 0.053% | 0.117% | 0% | 0.1%
7 | 2 | 0.06% | 0.007% | 0.016% | 0.1%
7 | 3 | 0.053% | -0.007% | 0.055% | 0.1%
7 | 4 | 0.064% | 0.011% | 0.059% | 0.1%

## 12. 5th Harmonics
[<<]
Accuracy Test Report of '5th Harmonics'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / 5th Harmonics'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:21:38 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Harmonics' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
1 | 0 | 0 | 1
3. Test Cases
Total 1 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 1 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° | {UcUbUa, 5th-10%, IcIbIa, 5th-40%} | 12 | 20 |  | ±0.1% | ±1% | 0.005% | PASSED | 2023-07-25 17:47
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #1. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0.041% | 0% | 0% | 0.1%
1 | 1 | -1.37% | -1.411% | 0% | 0.1%
1 | 2 | -6.286% | -4.916% | -2.538% | 0.1%
1 | 3 | -4.639% | 1.647% | -4.098% | 0.1%
1 | 4 | -2.264% | 2.375% | -4.396% | 0.1%
1 | 5 | 0.024% | 2.288% | -2.293% | 0.1%
1 | 6 | 0.485% | 0.461% | -0.585% | 0.1%
1 | 7 | 0.464% | -0.021% | 0.324% | 0.1%
1 | 8 | 0.276% | -0.188% | 0.408% | 0.1%
1 | 9 | 0.107% | -0.169% | 0.282% | 0.1%
1 | 10 | 0.024% | -0.083% | 0.136% | 0.1%
1 | 11 | -0.03% | -0.054% | 0.034% | 0.1%
1 | 12 | 0.021% | 0.051% | 0.005% | 0.1%

## 13. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:21:39 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RP | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | 0.043% | PASSED | 2023-07-25 17:50
Yes | 1 | RP | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.053% | PASSED | 2023-07-25 17:50
Yes | 2 | RP | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.065% | PASSED | 2023-07-25 17:51
Yes | 3 | RP | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.049% | PASSED | 2023-07-25 17:51
Yes | 4 | RP | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.067% | PASSED | 2023-07-25 17:51
Yes | 5 | RP | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | 0.061% | PASSED | 2023-07-25 17:52
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 0.058% | 0% | 0% | 0.1%
0 | 1 | 0.058% | 0% | 0% | 0.1%
0 | 2 | 0.036% | -0.022% | 0.051% | 0.1%
0 | 3 | 0.036% | 0% | 0.043% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0.049% | 0% | 0% | 0.1%
1 | 1 | 0.071% | 0.022% | 0% | 0.1%
1 | 2 | 0.049% | -0.022% | 0.056% | 0.1%
1 | 3 | 0.04% | -0.009% | 0.053% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | 0.053% | 0% | 0% | 0.1%
2 | 1 | 0.085% | 0.032% | 0% | 0.1%
2 | 2 | 0.062% | -0.023% | 0.067% | 0.1%
2 | 3 | 0.049% | -0.013% | 0.065% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.058% | 0% | 0% | 0.1%
3 | 1 | 0.049% | -0.009% | 0% | 0.1%
3 | 2 | 0.053% | 0.004% | 0.053% | 0.1%
3 | 3 | 0.044% | -0.009% | 0.049% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.044% | 0% | 0% | 0.1%
4 | 1 | 0.102% | 0.058% | 0% | 0.1%
4 | 2 | 0.058% | -0.044% | 0.068% | 0.1%
4 | 3 | 0.04% | -0.018% | 0.067% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | 0.076% | 0% | 0% | 0.1%
5 | 1 | 0.044% | -0.032% | 0% | 0.1%
5 | 2 | 0.036% | -0.008% | 0.052% | 0.1%
5 | 3 | 0.102% | 0.066% | 0.061% | 0.1%

## 14. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:21:43 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RP | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | 0.012% | PASSED | 2023-07-25 17:52
Yes | 1 | RP | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | -0.033% | PASSED | 2023-07-25 17:53
Yes | 2 | RP | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.041% | PASSED | 2023-07-25 17:54
Yes | 3 | RP | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.157% | PASSED | 2023-07-25 17:54
Yes | 4 | RP | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.151% | PASSED | 2023-07-25 17:55
Yes | 5 | RP | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | 0.041% | PASSED | 2023-07-25 17:55
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 19 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -22.45% | 0% | 0% | 0.1%
0 | 1 | -11.79% | 10.66% | 0% | 0.1%
0 | 2 | -4.206% | 7.579% | -12.81% | 0.1%
0 | 3 | 0% | 4.206% | -5.33% | 0.1%
0 | 4 | 0.036% | 0.036% | -1.39% | 0.1%
0 | 5 | 0.036% | 0% | 0.024% | 0.1%
0 | 6 | 0.375% | 0.339% | 0.149% | 0.1%
0 | 7 | 0.563% | 0.188% | 0.325% | 0.1%
0 | 8 | 0.437% | -0.126% | 0.458% | 0.1%
0 | 9 | 0.312% | -0.125% | 0.437% | 0.1%
0 | 10 | 0.098% | -0.214% | 0.282% | 0.1%
0 | 11 | 0.888% | 0.79% | 0.433% | 0.1%
0 | 12 | 1.115% | 0.227% | 0.7% | 0.1%
0 | 13 | 0.815% | -0.3% | 0.939% | 0.1%
0 | 14 | 0.491% | -0.324% | 0.807% | 0.1%
0 | 15 | 0.08% | -0.411% | 0.462% | 0.1%
0 | 16 | 0.009% | -0.071% | 0.193% | 0.1%
0 | 17 | 0.027% | 0.018% | 0.039% | 0.1%
0 | 18 | 0% | -0.027% | 0.012% | 0.1%
Deviations of %error on test id #1. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -13.2% | 0% | 0% | 0.1%
1 | 1 | 0.653% | 13.85% | 0% | 0.1%
1 | 2 | 0.779% | 0.126% | -3.923% | 0.1%
1 | 3 | 0.393% | -0.386% | 0.608% | 0.1%
1 | 4 | 0% | -0.393% | 0.391% | 0.1%
1 | 5 | -0.018% | -0.018% | 0.125% | 0.1%
1 | 6 | -0.044% | -0.026% | -0.021% | 0.1%
1 | 7 | -0.036% | 0.008% | -0.033% | 0.1%
Deviations of %error on test id #2. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -25.35% | 0% | 0% | 0.1%
2 | 1 | -18.31% | 7.036% | 0% | 0.1%
2 | 2 | -6.871% | 11.44% | -16.84% | 0.1%
2 | 3 | -3.862% | 3.009% | -9.682% | 0.1%
2 | 4 | 0.205% | 4.067% | -3.509% | 0.1%
2 | 5 | 0.518% | 0.313% | -1.046% | 0.1%
2 | 6 | 0.653% | 0.135% | 0.459% | 0.1%
2 | 7 | 0.428% | -0.225% | 0.533% | 0.1%
2 | 8 | 0.303% | -0.125% | 0.461% | 0.1%
2 | 9 | 0.428% | 0.125% | 0.386% | 0.1%
2 | 10 | 0.357% | -0.071% | 0.363% | 0.1%
2 | 11 | 0.241% | -0.116% | 0.342% | 0.1%
2 | 12 | 0.116% | -0.125% | 0.238% | 0.1%
2 | 13 | 0.062% | -0.054% | 0.14% | 0.1%
2 | 14 | 0.009% | -0.053% | 0.062% | 0.1%
2 | 15 | 0.053% | 0.044% | 0.041% | 0.1%
Deviations of %error on test id #3. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -20.46% | 0% | 0% | 0.1%
3 | 1 | -15.91% | 4.549% | 0% | 0.1%
3 | 2 | -4.32% | 11.59% | -13.56% | 0.1%
3 | 3 | -2.488% | 1.832% | -7.571% | 0.1%
3 | 4 | 0.25% | 2.738% | -2.186% | 0.1%
3 | 5 | 0.402% | 0.152% | -0.612% | 0.1%
3 | 6 | 0.232% | -0.17% | 0.295% | 0.1%
3 | 7 | 0.223% | -0.009% | 0.286% | 0.1%
3 | 8 | 0.142% | -0.081% | 0.199% | 0.1%
3 | 9 | 0.107% | -0.035% | 0.157% | 0.1%
Deviations of %error on test id #4. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -16.55% | 0% | 0% | 0.1%
4 | 1 | -7.285% | 9.27% | 0% | 0.1%
4 | 2 | -3.871% | 3.414% | -9.237% | 0.1%
4 | 3 | 0.473% | 4.344% | -3.561% | 0.1%
4 | 4 | 0.509% | 0.036% | -0.963% | 0.1%
4 | 5 | 0.68% | 0.171% | 0.554% | 0.1%
4 | 6 | 0.527% | -0.153% | 0.572% | 0.1%
4 | 7 | 0.348% | -0.179% | 0.518% | 0.1%
4 | 8 | 0.142% | -0.206% | 0.339% | 0.1%
4 | 9 | 0.42% | 0.278% | 0.303% | 0.1%
4 | 10 | 0.267% | -0.153% | 0.276% | 0.1%
4 | 11 | 0.196% | -0.071% | 0.294% | 0.1%
4 | 12 | 0.151% | -0.045% | 0.205% | 0.1%
4 | 13 | 0.107% | -0.044% | 0.151% | 0.1%
Deviations of %error on test id #5. Total 19 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -16.6% | 0% | 0% | 0.1%
5 | 1 | -12.46% | 4.134% | 0% | 0.1%
5 | 2 | -6.608% | 5.856% | -11.89% | 0.1%
5 | 3 | -2.488% | 4.12% | -7.187% | 0.1%
5 | 4 | -0.688% | 1.8% | -3.261% | 0.1%
5 | 5 | 0.276% | 0.964% | -0.967% | 0.1%
5 | 6 | 0.178% | -0.098% | -0.078% | 0.1%
5 | 7 | 0.125% | -0.053% | 0.193% | 0.1%
5 | 8 | 0.232% | 0.107% | 0.178% | 0.1%
5 | 9 | 0.788% | 0.556% | 0.382% | 0.1%
5 | 10 | 0.77% | -0.018% | 0.597% | 0.1%
5 | 11 | 0.87% | 0.1% | 0.809% | 0.1%
5 | 12 | 0.725% | -0.145% | 0.788% | 0.1%
5 | 13 | 0.536% | -0.189% | 0.71% | 0.1%
5 | 14 | 0.312% | -0.224% | 0.524% | 0.1%
5 | 15 | 0.151% | -0.161% | 0.333% | 0.1%
5 | 16 | 0.089% | -0.062% | 0.184% | 0.1%
5 | 17 | -0.009% | -0.098% | 0.077% | 0.1%
5 | 18 | 0.044% | 0.053% | 0.041% | 0.1%

## 15. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:21:49 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.01% | PASSED | 2023-07-25 17:58
Yes | 1 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0% | PASSED | 2023-07-25 17:58
Yes | 2 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.002% | PASSED | 2023-07-25 17:59
Yes | 3 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.016% | PASSED | 2023-07-25 17:59
Yes | 4 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.006% | PASSED | 2023-07-25 18:00
Yes | 5 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.025% | PASSED | 2023-07-25 18:00
Yes | 6 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.023% | PASSED | 2023-07-25 18:00
Yes | 7 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.024% | PASSED | 2023-07-25 18:01
Yes | 8 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.025% | PASSED | 2023-07-25 18:01
Yes | 9 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.027% | PASSED | 2023-07-25 18:02
Yes | 10 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.035% | PASSED | 2023-07-25 18:02
Yes | 11 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.037% | PASSED | 2023-07-25 18:02
Yes | 12 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.043% | PASSED | 2023-07-25 18:03
Yes | 13 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.045% | PASSED | 2023-07-25 18:03
Yes | 14 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.062% | PASSED | 2023-07-25 18:04
Yes | 15 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.053% | PASSED | 2023-07-25 18:04
Yes | 16 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.05% | PASSED | 2023-07-25 18:05
Yes | 17 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.061% | PASSED | 2023-07-25 18:05
Yes | 18 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.073% | PASSED | 2023-07-25 18:06
Yes | 19 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.064% | PASSED | 2023-07-25 18:06
Yes | 20 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.079% | PASSED | 2023-07-25 18:58
Yes | 21 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.07% | PASSED | 2023-07-25 18:07
Yes | 22 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.065% | PASSED | 2023-07-25 18:08
Yes | 23 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.083% | PASSED | 2023-07-25 18:08
Yes | 24 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.092% | PASSED | 2023-07-25 18:09
Yes | 25 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.073% | PASSED | 2023-07-25 18:09
Yes | 26 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.085% | PASSED | 2023-07-25 18:10
Yes | 27 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.092% | PASSED | 2023-07-25 18:11
Yes | 28 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.085% | PASSED | 2023-07-25 18:11
Yes | 29 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.094% | PASSED | 2023-07-25 18:12
Yes | 30 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.116% | PASSED | 2023-07-25 18:13
Yes | 31 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.113% | PASSED | 2023-07-25 18:14
Yes | 32 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.112% | PASSED | 2023-07-25 18:15
Yes | 33 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.104% | PASSED | 2023-07-25 18:16
Yes | 34 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.111% | PASSED | 2023-07-25 18:17
Yes | 35 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.131% | PASSED | 2023-07-25 18:18
Yes | 36 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.125% | PASSED | 2023-07-25 18:59
Yes | 37 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.126% | PASSED | 2023-07-25 18:59
Yes | 38 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.125% | PASSED | 2023-07-25 18:20
Yes | 39 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.114% | PASSED | 2023-07-25 18:22
Yes | 40 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.114% | PASSED | 2023-07-25 18:24
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 0.018% | 0% | 0% | 0.1%
0 | 1 | -0.031% | -0.049% | 0% | 0.1%
0 | 2 | 0.027% | 0.058% | 0.005% | 0.1%
0 | 3 | -0.027% | -0.054% | -0.01% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0.013% | 0% | 0% | 0.1%
1 | 1 | -0.004% | -0.017% | 0% | 0.1%
1 | 2 | 0.013% | 0.017% | 0.007% | 0.1%
1 | 3 | -0.009% | -0.022% | 0% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | 0.04% | 0% | 0% | 0.1%
2 | 1 | -0.027% | -0.067% | 0% | 0.1%
2 | 2 | 0.009% | 0.036% | 0.007% | 0.1%
2 | 3 | 0.013% | 0.004% | -0.002% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -0.013% | 0% | 0% | 0.1%
3 | 1 | 0.009% | 0.022% | 0% | 0.1%
3 | 2 | 0.022% | 0.013% | 0.006% | 0.1%
3 | 3 | 0.018% | -0.004% | 0.016% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.022% | 0% | 0% | 0.1%
4 | 1 | 0% | -0.022% | 0% | 0.1%
4 | 2 | 0.013% | 0.013% | 0.012% | 0.1%
4 | 3 | 0.004% | -0.009% | 0.006% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | 0.018% | 0% | 0% | 0.1%
5 | 1 | 0.013% | -0.005% | 0% | 0.1%
5 | 2 | 0.058% | 0.045% | 0.03% | 0.1%
5 | 3 | 0.004% | -0.054% | 0.025% | 0.1%
Deviations of %error on test id #6. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | 0.04% | 0% | 0% | 0.1%
6 | 1 | 0.013% | -0.027% | 0% | 0.1%
6 | 2 | 0.004% | -0.009% | 0.019% | 0.1%
6 | 3 | 0.053% | 0.049% | 0.023% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.062% | 0% | 0% | 0.1%
7 | 1 | 0.009% | -0.053% | 0% | 0.1%
7 | 2 | 0.058% | 0.049% | 0.043% | 0.1%
7 | 3 | 0.004% | -0.054% | 0.024% | 0.1%
Deviations of %error on test id #8. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | 0.04% | 0% | 0% | 0.1%
8 | 1 | 0% | -0.04% | 0% | 0.1%
8 | 2 | 0.044% | 0.044% | 0.028% | 0.1%
8 | 3 | 0.031% | -0.013% | 0.025% | 0.1%
Deviations of %error on test id #9. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | 0.049% | 0% | 0% | 0.1%
9 | 1 | 0% | -0.049% | 0% | 0.1%
9 | 2 | 0.044% | 0.044% | 0.031% | 0.1%
9 | 3 | 0.036% | -0.008% | 0.027% | 0.1%
Deviations of %error on test id #10. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | 0.053% | 0% | 0% | 0.1%
10 | 1 | 0% | -0.053% | 0% | 0.1%
10 | 2 | 0.053% | 0.053% | 0.035% | 0.1%
10 | 3 | 0.053% | 0% | 0.035% | 0.1%
Deviations of %error on test id #11. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | 0.08% | 0% | 0% | 0.1%
11 | 1 | 0.013% | -0.067% | 0% | 0.1%
11 | 2 | 0.04% | 0.027% | 0.044% | 0.1%
11 | 3 | 0.058% | 0.018% | 0.037% | 0.1%
Deviations of %error on test id #12. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | 0.076% | 0% | 0% | 0.1%
12 | 1 | 0.04% | -0.036% | 0% | 0.1%
12 | 2 | 0.053% | 0.013% | 0.056% | 0.1%
12 | 3 | 0.036% | -0.017% | 0.043% | 0.1%
Deviations of %error on test id #13. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | 0.049% | 0% | 0% | 0.1%
13 | 1 | 0.036% | -0.013% | 0% | 0.1%
13 | 2 | 0.058% | 0.022% | 0.048% | 0.1%
13 | 3 | 0.04% | -0.018% | 0.045% | 0.1%
Deviations of %error on test id #14. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | 0.04% | 0% | 0% | 0.1%
14 | 1 | 0.076% | 0.036% | 0% | 0.1%
14 | 2 | 0.049% | -0.027% | 0.055% | 0.1%
14 | 3 | 0.062% | 0.013% | 0.062% | 0.1%
Deviations of %error on test id #15. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | 0.049% | 0% | 0% | 0.1%
15 | 1 | 0.049% | 0% | 0% | 0.1%
15 | 2 | 0.031% | -0.018% | 0.043% | 0.1%
15 | 3 | 0.08% | 0.049% | 0.053% | 0.1%
Deviations of %error on test id #16. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | 0.067% | 0% | 0% | 0.1%
16 | 1 | 0.049% | -0.018% | 0% | 0.1%
16 | 2 | 0.022% | -0.027% | 0.046% | 0.1%
16 | 3 | 0.08% | 0.058% | 0.05% | 0.1%
Deviations of %error on test id #17. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | 0.089% | 0% | 0% | 0.1%
17 | 1 | 0.036% | -0.053% | 0% | 0.1%
17 | 2 | 0.089% | 0.053% | 0.071% | 0.1%
17 | 3 | 0.058% | -0.031% | 0.061% | 0.1%
Deviations of %error on test id #18. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | 0.036% | 0% | 0% | 0.1%
18 | 1 | 0.058% | 0.022% | 0% | 0.1%
18 | 2 | 0.08% | 0.022% | 0.058% | 0.1%
18 | 3 | 0.08% | 0% | 0.073% | 0.1%
Deviations of %error on test id #19. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | 0.049% | 0% | 0% | 0.1%
19 | 1 | 0.08% | 0.031% | 0% | 0.1%
19 | 2 | 0.031% | -0.049% | 0.053% | 0.1%
19 | 3 | 0.08% | 0.049% | 0.064% | 0.1%
Deviations of %error on test id #20. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | 0.062% | 0% | 0% | 0.1%
20 | 1 | 0.089% | 0.027% | 0% | 0.1%
20 | 2 | 0.071% | -0.018% | 0.074% | 0.1%
20 | 3 | 0.076% | 0.005% | 0.079% | 0.1%
Deviations of %error on test id #21. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | 0.089% | 0% | 0% | 0.1%
21 | 1 | 0.049% | -0.04% | 0% | 0.1%
21 | 2 | 0.071% | 0.022% | 0.07% | 0.1%
21 | 3 | 0.089% | 0.018% | 0.07% | 0.1%
Deviations of %error on test id #22. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | 0.093% | 0% | 0% | 0.1%
22 | 1 | 0.044% | -0.049% | 0% | 0.1%
22 | 2 | 0.071% | 0.027% | 0.069% | 0.1%
22 | 3 | 0.08% | 0.009% | 0.065% | 0.1%
Deviations of %error on test id #23. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | 0.107% | 0% | 0% | 0.1%
23 | 1 | 0.071% | -0.036% | 0% | 0.1%
23 | 2 | 0.071% | 0% | 0.083% | 0.1%
23 | 3 | 0.107% | 0.036% | 0.083% | 0.1%
Deviations of %error on test id #24. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | 0.08% | 0% | 0% | 0.1%
24 | 1 | 0.111% | 0.031% | 0% | 0.1%
24 | 2 | 0.076% | -0.035% | 0.089% | 0.1%
24 | 3 | 0.089% | 0.013% | 0.092% | 0.1%
Deviations of %error on test id #25. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | 0.076% | 0% | 0% | 0.1%
25 | 1 | 0.062% | -0.014% | 0% | 0.1%
25 | 2 | 0.071% | 0.009% | 0.07% | 0.1%
25 | 3 | 0.085% | 0.014% | 0.073% | 0.1%
Deviations of %error on test id #26. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | 0.085% | 0% | 0% | 0.1%
26 | 1 | 0.093% | 0.008% | 0% | 0.1%
26 | 2 | 0.08% | -0.013% | 0.086% | 0.1%
26 | 3 | 0.098% | 0.018% | 0.09% | 0.1%
26 | 4 | 0.076% | -0.022% | 0.085% | 0.1%
Deviations of %error on test id #27. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | 0.116% | 0% | 0% | 0.1%
27 | 1 | 0.076% | -0.04% | 0% | 0.1%
27 | 2 | 0.093% | 0.017% | 0.095% | 0.1%
27 | 3 | 0.107% | 0.014% | 0.092% | 0.1%
Deviations of %error on test id #28. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | 0.111% | 0% | 0% | 0.1%
28 | 1 | 0.067% | -0.044% | 0% | 0.1%
28 | 2 | 0.107% | 0.04% | 0.095% | 0.1%
28 | 3 | 0.08% | -0.027% | 0.085% | 0.1%
Deviations of %error on test id #29. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | 0.08% | 0% | 0% | 0.1%
29 | 1 | 0.111% | 0.031% | 0% | 0.1%
29 | 2 | 0.085% | -0.026% | 0.092% | 0.1%
29 | 3 | 0.085% | 0% | 0.094% | 0.1%
Deviations of %error on test id #30. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | 0.08% | 0% | 0% | 0.1%
30 | 1 | 0.12% | 0.04% | 0% | 0.1%
30 | 2 | 0.107% | -0.013% | 0.102% | 0.1%
30 | 3 | 0.12% | 0.013% | 0.116% | 0.1%
Deviations of %error on test id #31. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | 0.085% | 0% | 0% | 0.1%
31 | 1 | 0.12% | 0.035% | 0% | 0.1%
31 | 2 | 0.102% | -0.018% | 0.102% | 0.1%
31 | 3 | 0.116% | 0.014% | 0.113% | 0.1%
Deviations of %error on test id #32. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | 0.089% | 0% | 0% | 0.1%
32 | 1 | 0.116% | 0.027% | 0% | 0.1%
32 | 2 | 0.085% | -0.031% | 0.097% | 0.1%
32 | 3 | 0.134% | 0.049% | 0.112% | 0.1%
Deviations of %error on test id #33. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | 0.12% | 0% | 0% | 0.1%
33 | 1 | 0.08% | -0.04% | 0% | 0.1%
33 | 2 | 0.111% | 0.031% | 0.104% | 0.1%
33 | 3 | 0.12% | 0.009% | 0.104% | 0.1%
Deviations of %error on test id #34. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | 0.129% | 0% | 0% | 0.1%
34 | 1 | 0.093% | -0.036% | 0% | 0.1%
34 | 2 | 0.116% | 0.023% | 0.113% | 0.1%
34 | 3 | 0.125% | 0.009% | 0.111% | 0.1%
Deviations of %error on test id #35. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | 0.111% | 0% | 0% | 0.1%
35 | 1 | 0.156% | 0.045% | 0% | 0.1%
35 | 2 | 0.085% | -0.071% | 0.117% | 0.1%
35 | 3 | 0.151% | 0.066% | 0.131% | 0.1%
Deviations of %error on test id #36. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | 0.116% | 0% | 0% | 0.1%
36 | 1 | 0.125% | 0.009% | 0% | 0.1%
36 | 2 | 0.129% | 0.004% | 0.123% | 0.1%
36 | 3 | 0.12% | -0.009% | 0.125% | 0.1%
Deviations of %error on test id #37. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | 0.085% | 0% | 0% | 0.1%
37 | 1 | 0.125% | 0.04% | 0% | 0.1%
37 | 2 | 0.129% | 0.004% | 0.113% | 0.1%
37 | 3 | 0.125% | -0.004% | 0.126% | 0.1%
Deviations of %error on test id #38. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | 0.125% | 0% | 0% | 0.1%
38 | 1 | 0.138% | 0.013% | 0% | 0.1%
38 | 2 | 0.116% | -0.022% | 0.126% | 0.1%
38 | 3 | 0.12% | 0.004% | 0.125% | 0.1%
Deviations of %error on test id #39. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | 0.134% | 0% | 0% | 0.1%
39 | 1 | 0.093% | -0.041% | 0% | 0.1%
39 | 2 | 0.134% | 0.041% | 0.12% | 0.1%
39 | 3 | 0.116% | -0.018% | 0.114% | 0.1%
Deviations of %error on test id #40. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | 0.12% | 0% | 0% | 0.1%
40 | 1 | 0.134% | 0.014% | 0% | 0.1%
40 | 2 | 0.08% | -0.054% | 0.111% | 0.1%
40 | 3 | 0.129% | 0.049% | 0.114% | 0.1%

## 16. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:22:11 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.026% | PASSED | 2023-07-25 18:24
Yes | 1 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.027% | PASSED | 2023-07-25 18:25
Yes | 2 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.024% | PASSED | 2023-07-25 18:25
Yes | 3 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.151% | PASSED | 2023-07-25 18:26
Yes | 4 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.06% | PASSED | 2023-07-25 18:26
Yes | 5 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.229% | PASSED | 2023-07-25 18:26
Yes | 6 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.151% | PASSED | 2023-07-25 18:27
Yes | 7 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.178% | PASSED | 2023-07-25 19:00
Yes | 8 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.077% | PASSED | 2023-07-25 18:28
Yes | 9 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.181% | PASSED | 2023-07-25 18:28
Yes | 10 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.047% | PASSED | 2023-07-25 18:29
Yes | 11 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.178% | PASSED | 2023-07-25 18:29
Yes | 12 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.044% | PASSED | 2023-07-25 18:30
Yes | 13 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.042% | PASSED | 2023-07-25 18:31
Yes | 14 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.065% | PASSED | 2023-07-25 18:31
Yes | 15 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.05% | PASSED | 2023-07-25 18:32
Yes | 16 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.059% | PASSED | 2023-07-25 18:32
Yes | 17 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.065% | PASSED | 2023-07-25 18:33
Yes | 18 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.059% | PASSED | 2023-07-25 19:00
Yes | 19 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.053% | PASSED | 2023-07-25 18:34
Yes | 20 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.065% | PASSED | 2023-07-25 18:35
Yes | 21 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.071% | PASSED | 2023-07-25 18:37
Yes | 22 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.047% | PASSED | 2023-07-25 18:39
Yes | 23 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.074% | PASSED | 2023-07-25 18:40
Yes | 24 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.267% | PASSED | 2023-07-25 19:00
Yes | 25 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.229% | PASSED | 2023-07-25 19:01
Yes | 26 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.276% | PASSED | 2023-07-25 19:01
Yes | 27 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.065% | PASSED | 2023-07-25 18:42
Yes | 28 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.095% | PASSED | 2023-07-25 18:43
Yes | 29 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.05% | PASSED | 2023-07-25 18:44
Yes | 30 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.074% | PASSED | 2023-07-25 18:46
Yes | 31 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.059% | PASSED | 2023-07-25 18:47
Yes | 32 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.071% | PASSED | 2023-07-25 19:02
Yes | 33 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.05% | PASSED | 2023-07-25 18:51
Yes | 34 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.226% | PASSED | 2023-07-25 19:02
Yes | 35 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.065% | PASSED | 2023-07-25 19:03
Yes | 36 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.071% | PASSED | 2023-07-25 19:03
Yes | 37 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.083% | PASSED | 2023-07-25 19:04
Yes | 38 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.065% | PASSED | 2023-07-25 19:04
Yes | 39 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.166% | PASSED | 2023-07-25 19:05
Yes | 40 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 0° | 0.5L, cos(ϕ), ϕ=60° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.062% | PASSED | 2023-07-25 19:05
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 0.027% | 0% | 0% | 0.1%
0 | 1 | -0.009% | -0.036% | 0% | 0.1%
0 | 2 | 0.044% | 0.053% | 0.021% | 0.1%
0 | 3 | -0.027% | -0.071% | 0.003% | 0.1%
0 | 4 | 0.062% | 0.089% | 0.026% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -0.009% | 0% | 0% | 0.1%
1 | 1 | 0.018% | 0.027% | 0% | 0.1%
1 | 2 | 0.027% | 0.009% | 0.012% | 0.1%
1 | 3 | 0.036% | 0.009% | 0.027% | 0.1%
Deviations of %error on test id #2. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -0.027% | 0% | 0% | 0.1%
2 | 1 | 0.062% | 0.089% | 0% | 0.1%
2 | 2 | 0.036% | -0.026% | 0.024% | 0.1%
2 | 3 | 0.036% | 0% | 0.045% | 0.1%
2 | 4 | 0% | -0.036% | 0.024% | 0.1%
Deviations of %error on test id #3. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.294% | 0% | 0% | 0.1%
3 | 1 | 0.142% | -0.152% | 0% | 0.1%
3 | 2 | 0.214% | 0.072% | 0.217% | 0.1%
3 | 3 | 0.169% | -0.045% | 0.175% | 0.1%
3 | 4 | 0.071% | -0.098% | 0.151% | 0.1%
Deviations of %error on test id #4. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.044% | 0% | 0% | 0.1%
4 | 1 | 0.107% | 0.063% | 0% | 0.1%
4 | 2 | 0.151% | 0.044% | 0.101% | 0.1%
4 | 3 | 0.285% | 0.134% | 0.181% | 0.1%
4 | 4 | 0.241% | -0.044% | 0.226% | 0.1%
4 | 5 | 0.125% | -0.116% | 0.217% | 0.1%
4 | 6 | 0.205% | 0.08% | 0.19% | 0.1%
4 | 7 | 0.16% | -0.045% | 0.163% | 0.1%
4 | 8 | 0.053% | -0.107% | 0.139% | 0.1%
4 | 9 | 0.107% | 0.054% | 0.107% | 0.1%
4 | 10 | 0.036% | -0.071% | 0.065% | 0.1%
4 | 11 | 0.036% | 0% | 0.06% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | 0.062% | 0% | 0% | 0.1%
5 | 1 | 0.16% | 0.098% | 0% | 0.1%
5 | 2 | 0.232% | 0.072% | 0.151% | 0.1%
5 | 3 | 0.294% | 0.062% | 0.229% | 0.1%
Deviations of %error on test id #6. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | 0.276% | 0% | 0% | 0.1%
6 | 1 | 0.205% | -0.071% | 0% | 0.1%
6 | 2 | 0.151% | -0.054% | 0.211% | 0.1%
6 | 3 | 0.178% | 0.027% | 0.178% | 0.1%
6 | 4 | 0.125% | -0.053% | 0.151% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.205% | 0% | 0% | 0.1%
7 | 1 | 0.205% | 0% | 0% | 0.1%
7 | 2 | 0.214% | 0.009% | 0.208% | 0.1%
7 | 3 | 0.116% | -0.098% | 0.178% | 0.1%
Deviations of %error on test id #8. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | 0.053% | 0% | 0% | 0.1%
8 | 1 | 0.116% | 0.063% | 0% | 0.1%
8 | 2 | 0.08% | -0.036% | 0.083% | 0.1%
8 | 3 | 0.036% | -0.044% | 0.077% | 0.1%
Deviations of %error on test id #9. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | 0.125% | 0% | 0% | 0.1%
9 | 1 | 0.178% | 0.053% | 0% | 0.1%
9 | 2 | 0.294% | 0.116% | 0.199% | 0.1%
9 | 3 | 0.267% | -0.027% | 0.246% | 0.1%
9 | 4 | 0.214% | -0.053% | 0.258% | 0.1%
9 | 5 | 0.151% | -0.063% | 0.211% | 0.1%
9 | 6 | 0.178% | 0.027% | 0.181% | 0.1%
Deviations of %error on test id #10. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | 0.036% | 0% | 0% | 0.1%
10 | 1 | 0.044% | 0.008% | 0% | 0.1%
10 | 2 | 0.053% | 0.009% | 0.044% | 0.1%
10 | 3 | 0.044% | -0.009% | 0.047% | 0.1%
Deviations of %error on test id #11. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | 0.169% | 0% | 0% | 0.1%
11 | 1 | 0.303% | 0.134% | 0% | 0.1%
11 | 2 | 0.214% | -0.089% | 0.229% | 0.1%
11 | 3 | 0.205% | -0.009% | 0.241% | 0.1%
11 | 4 | 0.205% | 0% | 0.208% | 0.1%
11 | 5 | 0.125% | -0.08% | 0.178% | 0.1%
Deviations of %error on test id #12. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | 0.098% | 0% | 0% | 0.1%
12 | 1 | 0.036% | -0.062% | 0% | 0.1%
12 | 2 | 0.053% | 0.017% | 0.062% | 0.1%
12 | 3 | 0.044% | -0.009% | 0.044% | 0.1%
Deviations of %error on test id #13. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | 0.036% | 0% | 0% | 0.1%
13 | 1 | 0.036% | 0% | 0% | 0.1%
13 | 2 | 0.053% | 0.017% | 0.042% | 0.1%
13 | 3 | 0.036% | -0.017% | 0.042% | 0.1%
Deviations of %error on test id #14. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | 0.044% | 0% | 0% | 0.1%
14 | 1 | 0.116% | 0.072% | 0% | 0.1%
14 | 2 | 0.366% | 0.25% | 0.175% | 0.1%
14 | 3 | 0.563% | 0.197% | 0.348% | 0.1%
14 | 4 | 0.375% | -0.188% | 0.435% | 0.1%
14 | 5 | 0.375% | 0% | 0.438% | 0.1%
14 | 6 | 0.366% | -0.009% | 0.372% | 0.1%
14 | 7 | 0.205% | -0.161% | 0.315% | 0.1%
14 | 8 | 0.205% | 0% | 0.259% | 0.1%
14 | 9 | 0.151% | -0.054% | 0.187% | 0.1%
14 | 10 | 0.036% | -0.115% | 0.131% | 0.1%
14 | 11 | 0.036% | 0% | 0.074% | 0.1%
14 | 12 | 0.053% | 0.017% | 0.042% | 0.1%
14 | 13 | 0.107% | 0.054% | 0.065% | 0.1%
Deviations of %error on test id #15. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | 0.062% | 0% | 0% | 0.1%
15 | 1 | 0.053% | -0.009% | 0% | 0.1%
15 | 2 | 0.036% | -0.017% | 0.05% | 0.1%
15 | 3 | 0.071% | 0.035% | 0.053% | 0.1%
15 | 4 | 0.044% | -0.027% | 0.05% | 0.1%
Deviations of %error on test id #16. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | 0.053% | 0% | 0% | 0.1%
16 | 1 | 0.062% | 0.009% | 0% | 0.1%
16 | 2 | 0.044% | -0.018% | 0.053% | 0.1%
16 | 3 | 0.071% | 0.027% | 0.059% | 0.1%
Deviations of %error on test id #17. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | 0.071% | 0% | 0% | 0.1%
17 | 1 | 0.062% | -0.009% | 0% | 0.1%
17 | 2 | 0.08% | 0.018% | 0.071% | 0.1%
17 | 3 | 0.053% | -0.027% | 0.065% | 0.1%
Deviations of %error on test id #18. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | 0.027% | 0% | 0% | 0.1%
18 | 1 | 0.053% | 0.026% | 0% | 0.1%
18 | 2 | 0.098% | 0.045% | 0.059% | 0.1%
18 | 3 | 0.027% | -0.071% | 0.059% | 0.1%
Deviations of %error on test id #19. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | 0.116% | 0% | 0% | 0.1%
19 | 1 | 0.053% | -0.063% | 0% | 0.1%
19 | 2 | 0.044% | -0.009% | 0.071% | 0.1%
19 | 3 | 0.071% | 0.027% | 0.056% | 0.1%
19 | 4 | 0.044% | -0.027% | 0.053% | 0.1%
Deviations of %error on test id #20. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | 0.08% | 0% | 0% | 0.1%
20 | 1 | 0.062% | -0.018% | 0% | 0.1%
20 | 2 | 0.053% | -0.009% | 0.065% | 0.1%
20 | 3 | 0.08% | 0.027% | 0.065% | 0.1%
Deviations of %error on test id #21. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | 0.053% | 0% | 0% | 0.1%
21 | 1 | 0.071% | 0.018% | 0% | 0.1%
21 | 2 | 0.044% | -0.027% | 0.056% | 0.1%
21 | 3 | 0.125% | 0.081% | 0.08% | 0.1%
21 | 4 | 0.044% | -0.081% | 0.071% | 0.1%
Deviations of %error on test id #22. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | 0.062% | 0% | 0% | 0.1%
22 | 1 | 0.044% | -0.018% | 0% | 0.1%
22 | 2 | -0.009% | -0.053% | 0.032% | 0.1%
22 | 3 | 0.107% | 0.116% | 0.047% | 0.1%
22 | 4 | 0.044% | -0.063% | 0.047% | 0.1%
22 | 5 | 0.053% | 0.009% | 0.068% | 0.1%
22 | 6 | 0.044% | -0.009% | 0.047% | 0.1%
Deviations of %error on test id #23. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | 0.125% | 0% | 0% | 0.1%
23 | 1 | 0.053% | -0.072% | 0% | 0.1%
23 | 2 | 0.044% | -0.009% | 0.074% | 0.1%
23 | 3 | 0.125% | 0.081% | 0.074% | 0.1%
Deviations of %error on test id #24. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | 0.036% | 0% | 0% | 0.1%
24 | 1 | 0.098% | 0.062% | 0% | 0.1%
24 | 2 | 0.062% | -0.036% | 0.065% | 0.1%
24 | 3 | 0.232% | 0.17% | 0.131% | 0.1%
24 | 4 | 0.285% | 0.053% | 0.193% | 0.1%
24 | 5 | 0.25% | -0.035% | 0.256% | 0.1%
24 | 6 | 0.267% | 0.017% | 0.267% | 0.1%
Deviations of %error on test id #25. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | 0.036% | 0% | 0% | 0.1%
25 | 1 | 0.116% | 0.08% | 0% | 0.1%
25 | 2 | 0.294% | 0.178% | 0.149% | 0.1%
25 | 3 | 0.312% | 0.018% | 0.241% | 0.1%
25 | 4 | 0.276% | -0.036% | 0.294% | 0.1%
25 | 5 | 0.214% | -0.062% | 0.267% | 0.1%
25 | 6 | 0.196% | -0.018% | 0.229% | 0.1%
Deviations of %error on test id #26. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | 0.205% | 0% | 0% | 0.1%
26 | 1 | 0.33% | 0.125% | 0% | 0.1%
26 | 2 | 0.285% | -0.045% | 0.273% | 0.1%
26 | 3 | 0.232% | -0.053% | 0.282% | 0.1%
26 | 4 | 0.312% | 0.08% | 0.276% | 0.1%
Deviations of %error on test id #27. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | 0.071% | 0% | 0% | 0.1%
27 | 1 | 0.036% | -0.035% | 0% | 0.1%
27 | 2 | 0.107% | 0.071% | 0.071% | 0.1%
27 | 3 | 0.053% | -0.054% | 0.065% | 0.1%
27 | 4 | 0.036% | -0.017% | 0.065% | 0.1%
Deviations of %error on test id #28. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | 0.071% | 0% | 0% | 0.1%
28 | 1 | 0.107% | 0.036% | 0% | 0.1%
28 | 2 | 0.071% | -0.036% | 0.083% | 0.1%
28 | 3 | 0.107% | 0.036% | 0.095% | 0.1%
Deviations of %error on test id #29. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | 0.116% | 0% | 0% | 0.1%
29 | 1 | 0.044% | -0.072% | 0% | 0.1%
29 | 2 | 0.062% | 0.018% | 0.074% | 0.1%
29 | 3 | 0.044% | -0.018% | 0.05% | 0.1%
Deviations of %error on test id #30. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | 0.134% | 0% | 0% | 0.1%
30 | 1 | 0.044% | -0.09% | 0% | 0.1%
30 | 2 | 0.134% | 0.09% | 0.104% | 0.1%
30 | 3 | 0.044% | -0.09% | 0.074% | 0.1%
Deviations of %error on test id #31. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | 0.053% | 0% | 0% | 0.1%
31 | 1 | 0.116% | 0.063% | 0% | 0.1%
31 | 2 | 0.062% | -0.054% | 0.077% | 0.1%
31 | 3 | 0.044% | -0.018% | 0.074% | 0.1%
31 | 4 | 0.071% | 0.027% | 0.059% | 0.1%
Deviations of %error on test id #32. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | 0.25% | 0% | 0% | 0.1%
32 | 1 | 0.33% | 0.08% | 0% | 0.1%
32 | 2 | 0.303% | -0.027% | 0.294% | 0.1%
32 | 3 | 0.178% | -0.125% | 0.27% | 0.1%
32 | 4 | 0.276% | 0.098% | 0.252% | 0.1%
32 | 5 | 0.142% | -0.134% | 0.199% | 0.1%
32 | 6 | 0.142% | 0% | 0.187% | 0.1%
32 | 7 | 0.053% | -0.089% | 0.112% | 0.1%
32 | 8 | 0.036% | -0.017% | 0.077% | 0.1%
32 | 9 | 0.125% | 0.089% | 0.071% | 0.1%
Deviations of %error on test id #33. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | 0.125% | 0% | 0% | 0.1%
33 | 1 | 0.036% | -0.089% | 0% | 0.1%
33 | 2 | 0.044% | 0.008% | 0.068% | 0.1%
33 | 3 | 0.071% | 0.027% | 0.05% | 0.1%
Deviations of %error on test id #34. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | 0.071% | 0% | 0% | 0.1%
34 | 1 | 0.205% | 0.134% | 0% | 0.1%
34 | 2 | 0.321% | 0.116% | 0.199% | 0.1%
34 | 3 | 0.285% | -0.036% | 0.27% | 0.1%
34 | 4 | 0.232% | -0.053% | 0.279% | 0.1%
34 | 5 | 0.303% | 0.071% | 0.273% | 0.1%
34 | 6 | 0.142% | -0.161% | 0.226% | 0.1%
Deviations of %error on test id #35. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | 0.107% | 0% | 0% | 0.1%
35 | 1 | 0.16% | 0.053% | 0% | 0.1%
35 | 2 | 0.339% | 0.179% | 0.202% | 0.1%
35 | 3 | 0.366% | 0.027% | 0.288% | 0.1%
35 | 4 | 0.241% | -0.125% | 0.315% | 0.1%
35 | 5 | 0.25% | 0.009% | 0.286% | 0.1%
35 | 6 | 0.142% | -0.108% | 0.211% | 0.1%
35 | 7 | 0.142% | 0% | 0.178% | 0.1%
35 | 8 | 0.169% | 0.027% | 0.151% | 0.1%
35 | 9 | 0.044% | -0.125% | 0.118% | 0.1%
35 | 10 | 0.053% | 0.009% | 0.089% | 0.1%
35 | 11 | 0.044% | -0.009% | 0.047% | 0.1%
35 | 12 | 0.098% | 0.054% | 0.065% | 0.1%
Deviations of %error on test id #36. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | 0.071% | 0% | 0% | 0.1%
36 | 1 | 0.232% | 0.161% | 0% | 0.1%
36 | 2 | 0.321% | 0.089% | 0.208% | 0.1%
36 | 3 | 0.384% | 0.063% | 0.312% | 0.1%
36 | 4 | 0.232% | -0.152% | 0.312% | 0.1%
36 | 5 | 0.294% | 0.062% | 0.303% | 0.1%
36 | 6 | 0.187% | -0.107% | 0.238% | 0.1%
36 | 7 | 0.134% | -0.053% | 0.205% | 0.1%
36 | 8 | 0.178% | 0.044% | 0.166% | 0.1%
36 | 9 | 0.036% | -0.142% | 0.116% | 0.1%
36 | 10 | 0.036% | 0% | 0.083% | 0.1%
36 | 11 | 0.116% | 0.08% | 0.063% | 0.1%
36 | 12 | 0.036% | -0.08% | 0.063% | 0.1%
36 | 13 | 0.062% | 0.026% | 0.071% | 0.1%
Deviations of %error on test id #37. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | 0.053% | 0% | 0% | 0.1%
37 | 1 | 0.142% | 0.089% | 0% | 0.1%
37 | 2 | 0.053% | -0.089% | 0.083% | 0.1%
37 | 3 | 0.116% | 0.063% | 0.104% | 0.1%
37 | 4 | 0.08% | -0.036% | 0.083% | 0.1%
Deviations of %error on test id #38. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | 0.205% | 0% | 0% | 0.1%
38 | 1 | 0.089% | -0.116% | 0% | 0.1%
38 | 2 | 0.036% | -0.053% | 0.11% | 0.1%
38 | 3 | 0.053% | 0.017% | 0.059% | 0.1%
38 | 4 | 0.107% | 0.054% | 0.065% | 0.1%
Deviations of %error on test id #39. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | 0.134% | 0% | 0% | 0.1%
39 | 1 | 0.196% | 0.062% | 0% | 0.1%
39 | 2 | 0.169% | -0.027% | 0.166% | 0.1%
39 | 3 | 0.134% | -0.035% | 0.166% | 0.1%
Deviations of %error on test id #40. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | 0.107% | 0% | 0% | 0.1%
40 | 1 | 0.071% | -0.036% | 0% | 0.1%
40 | 2 | 0.053% | -0.018% | 0.077% | 0.1%
40 | 3 | 0.062% | 0.009% | 0.062% | 0.1%

## 17. PF = 1.0
[<<]
Accuracy Test Report of 'PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:22:37 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 20 | 20 |  | ±0.1% | ±0.5% | 0.045% | PASSED | 2023-07-25 19:31
Yes | 1 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 12 | 20 |  | ±0.1% | ±0.5% | 0.04% | PASSED | 2023-07-25 19:31
Yes | 3 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.5% | 0.04% | PASSED | 2023-07-25 19:31
Yes | 4 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0.047% | PASSED | 2023-07-25 19:32
Yes | 7 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0.073% | PASSED | 2023-07-25 19:32
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 0.027% | 0% | 0% | 0.1%
0 | 1 | 0.027% | 0% | 0% | 0.1%
0 | 2 | 0.071% | 0.044% | 0.042% | 0.1%
0 | 3 | 0.036% | -0.035% | 0.045% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -0.036% | 0% | 0% | 0.1%
1 | 1 | 0.027% | 0.063% | 0% | 0.1%
1 | 2 | 0.059% | 0.032% | 0.017% | 0.1%
1 | 3 | 0.033% | -0.026% | 0.04% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.036% | 0% | 0% | 0.1%
3 | 1 | -0.004% | -0.04% | 0% | 0.1%
3 | 2 | 0.04% | 0.044% | 0.024% | 0.1%
3 | 3 | 0.085% | 0.045% | 0.04% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.009% | 0% | 0% | 0.1%
4 | 1 | 0.009% | 0% | 0% | 0.1%
4 | 2 | 0.071% | 0.062% | 0.03% | 0.1%
4 | 3 | 0.062% | -0.009% | 0.047% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.071% | 0% | 0% | 0.1%
7 | 1 | 0.073% | 0.002% | 0% | 0.1%
7 | 2 | 0.087% | 0.014% | 0.077% | 0.1%
7 | 3 | 0.06% | -0.027% | 0.073% | 0.1%

## 18. PF = 0.5L
[<<]
Accuracy Test Report of 'PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:22:40 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | 0.273% | PASSED | 2023-07-25 19:33
Yes | 1 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | 0.303% | PASSED | 2023-07-25 19:34
Yes | 3 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.199% | PASSED | 2023-07-25 19:34
Yes | 4 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | -0.041% | PASSED | 2023-07-25 19:34
Yes | 7 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.169% | PASSED | 2023-07-25 19:36
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -12.01% | 0% | 0% | 0.1%
0 | 1 | -37.56% | -25.55% | 0% | 0.1%
0 | 2 | -33.08% | 4.479% | -27.55% | 0.1%
0 | 3 | -18.7% | 14.38% | -29.78% | 0.1%
0 | 4 | -2.411% | 16.29% | -18.07% | 0.1%
0 | 5 | 0.071% | 2.482% | -7.014% | 0.1%
0 | 6 | 0.267% | 0.196% | -0.691% | 0.1%
0 | 7 | 0.018% | -0.249% | 0.119% | 0.1%
0 | 8 | -0.036% | -0.054% | 0.083% | 0.1%
0 | 9 | -0.248% | -0.212% | -0.089% | 0.1%
0 | 10 | -0.916% | -0.668% | -0.4% | 0.1%
0 | 11 | -0.916% | 0% | -0.693% | 0.1%
0 | 12 | -0.337% | 0.579% | -0.723% | 0.1%
0 | 13 | 0.16% | 0.497% | -0.364% | 0.1%
0 | 14 | 0.25% | 0.09% | 0.024% | 0.1%
0 | 15 | 0.303% | 0.053% | 0.238% | 0.1%
0 | 16 | 0.267% | -0.036% | 0.273% | 0.1%
Deviations of %error on test id #1. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -36.29% | 0% | 0% | 0.1%
1 | 1 | -18.88% | 17.4% | 0% | 0.1%
1 | 2 | -2.128% | 16.76% | -19.1% | 0.1%
1 | 3 | 0.279% | 2.407% | -6.911% | 0.1%
1 | 4 | 0.214% | -0.065% | -0.545% | 0.1%
1 | 5 | 0% | -0.214% | 0.164% | 0.1%
1 | 6 | -0.101% | -0.101% | 0.038% | 0.1%
1 | 7 | -0.554% | -0.453% | -0.218% | 0.1%
1 | 8 | -0.354% | 0.2% | -0.336% | 0.1%
1 | 9 | 0.041% | 0.395% | -0.289% | 0.1%
1 | 10 | 0.238% | 0.197% | -0.025% | 0.1%
1 | 11 | 0.291% | 0.053% | 0.19% | 0.1%
1 | 12 | 0.261% | -0.03% | 0.263% | 0.1%
1 | 13 | 0.357% | 0.096% | 0.303% | 0.1%
Deviations of %error on test id #3. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -25.8% | 0% | 0% | 0.1%
3 | 1 | -16.42% | 9.382% | 0% | 0.1%
3 | 2 | -4.263% | 12.16% | -15.49% | 0.1%
3 | 3 | -0.504% | 3.759% | -7.062% | 0.1%
3 | 4 | 0.16% | 0.664% | -1.536% | 0.1%
3 | 5 | 0.107% | -0.053% | -0.079% | 0.1%
3 | 6 | -0.115% | -0.222% | 0.051% | 0.1%
3 | 7 | -0.027% | 0.088% | -0.012% | 0.1%
3 | 8 | -0.133% | -0.106% | -0.092% | 0.1%
3 | 9 | -0.354% | -0.221% | -0.171% | 0.1%
3 | 10 | -0.592% | -0.238% | -0.36% | 0.1%
3 | 11 | -0.46% | 0.132% | -0.469% | 0.1%
3 | 12 | -0.204% | 0.256% | -0.419% | 0.1%
3 | 13 | 0.134% | 0.338% | -0.177% | 0.1%
3 | 14 | 0.16% | 0.026% | 0.03% | 0.1%
3 | 15 | 0.223% | 0.063% | 0.172% | 0.1%
3 | 16 | 0.214% | -0.009% | 0.199% | 0.1%
Deviations of %error on test id #4. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -33.59% | 0% | 0% | 0.1%
4 | 1 | -11.26% | 22.32% | 0% | 0.1%
4 | 2 | 0.169% | 11.43% | -14.89% | 0.1%
4 | 3 | -0.195% | -0.364% | -3.763% | 0.1%
4 | 4 | -0.098% | 0.097% | -0.041% | 0.1%
4 | 5 | -0.053% | 0.045% | -0.115% | 0.1%
4 | 6 | 0.027% | 0.08% | -0.041% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.082% | 0% | 0% | 0.1%
7 | 1 | 0.175% | 0.093% | 0% | 0.1%
7 | 2 | 0.162% | -0.013% | 0.14% | 0.1%
7 | 3 | 0.169% | 0.007% | 0.169% | 0.1%

## 19. PF = 0.8C
[<<]
Accuracy Test Report of 'PF = 0.8C'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / PF = 0.8C'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:22:45 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | -0.047% | PASSED | 2023-07-25 19:36
Yes | 1 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | 0.008% | PASSED | 2023-07-25 19:36
Yes | 3 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | -0.024% | PASSED | 2023-07-25 19:37
Yes | 4 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.11% | PASSED | 2023-07-25 19:37
Yes | 7 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.032% | PASSED | 2023-07-25 19:38
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -44.71% | 0% | 0% | 0.1%
0 | 1 | 87.19% | 131.9% | 0% | 0.1%
0 | 2 | 5.614% | -81.57% | 16.03% | 0.1%
0 | 3 | -0.018% | -5.632% | 30.93% | 0.1%
0 | 4 | -0.053% | -0.035% | 1.848% | 0.1%
0 | 5 | -0.089% | -0.036% | -0.053% | 0.1%
0 | 6 | 0% | 0.089% | -0.047% | 0.1%
Deviations of %error on test id #1. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 730.1% | 0% | 0% | 0.1%
1 | 1 | 39.68% | -690.4% | 0% | 0.1%
1 | 2 | 13.43% | -26.25% | 261.1% | 0.1%
1 | 3 | -0.083% | -13.51% | 17.68% | 0.1%
1 | 4 | -0.053% | 0.03% | 4.431% | 0.1%
1 | 5 | -0.059% | -0.006% | -0.065% | 0.1%
1 | 6 | 0.154% | 0.213% | 0.014% | 0.1%
1 | 7 | 0.172% | 0.018% | 0.089% | 0.1%
1 | 8 | 0.214% | 0.042% | 0.18% | 0.1%
1 | 9 | 0.107% | -0.107% | 0.164% | 0.1%
1 | 10 | 0.053% | -0.054% | 0.125% | 0.1%
1 | 11 | -0.018% | -0.071% | 0.047% | 0.1%
1 | 12 | -0.012% | 0.006% | 0.008% | 0.1%
Deviations of %error on test id #3. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 132.1% | 0% | 0% | 0.1%
3 | 1 | 528.1% | 396% | 0% | 0.1%
3 | 2 | 153.5% | -374.6% | 271.3% | 0.1%
3 | 3 | 102.2% | -51.39% | 261.3% | 0.1%
3 | 4 | 43.4% | -58.75% | 99.7% | 0.1%
3 | 5 | 7.01% | -36.39% | 50.86% | 0.1%
3 | 6 | 5.505% | -1.505% | 18.64% | 0.1%
3 | 7 | 0.053% | -5.452% | 4.189% | 0.1%
3 | 8 | 0.053% | 0% | 1.87% | 0.1%
3 | 9 | -0.018% | -0.071% | 0.029% | 0.1%
3 | 10 | -0.036% | -0.018% | 0% | 0.1%
3 | 11 | -0.018% | 0.018% | -0.024% | 0.1%
Deviations of %error on test id #4. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 402.7% | 0% | 0% | 0.1%
4 | 1 | 108.1% | -294.6% | 0% | 0.1%
4 | 2 | 15.81% | -92.25% | 175.5% | 0.1%
4 | 3 | 0.098% | -15.71% | 41.32% | 0.1%
4 | 4 | 0.08% | -0.018% | 5.33% | 0.1%
4 | 5 | 0.134% | 0.054% | 0.104% | 0.1%
4 | 6 | 0.116% | -0.018% | 0.11% | 0.1%
Deviations of %error on test id #7. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.134% | 0% | 0% | 0.1%
7 | 1 | -0.004% | -0.138% | 0% | 0.1%
7 | 2 | 0.025% | 0.029% | 0.052% | 0.1%
7 | 3 | 0.037% | 0.012% | 0.019% | 0.1%
7 | 4 | 0.034% | -0.003% | 0.032% | 0.1%

## 20. 5th Harmonics
[<<]
Accuracy Test Report of '5th Harmonics'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / 5th Harmonics'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:22:50 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Harmonics' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
1 | 0 | 0 | 1
3. Test Cases
Total 1 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 1 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° | {UcUbUa, 5th-10%, IcIbIa, 5th-40%} | 12 | 20 |  | ±0.1% | ±1% | 0.021% | PASSED | 2023-07-25 19:39
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #1. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -0.248% | 0% | 0% | 0.1%
1 | 1 | -0.124% | 0.124% | 0% | 0.1%
1 | 2 | -1.952% | -1.828% | -0.775% | 0.1%
1 | 3 | -4.785% | -2.833% | -2.287% | 0.1%
1 | 4 | -3.353% | 1.432% | -3.363% | 0.1%
1 | 5 | -0.983% | 2.37% | -3.04% | 0.1%
1 | 6 | -0.077% | 0.906% | -1.471% | 0.1%
1 | 7 | 0.393% | 0.47% | -0.222% | 0.1%
1 | 8 | 0.434% | 0.041% | 0.25% | 0.1%
1 | 9 | 0.199% | -0.235% | 0.342% | 0.1%
1 | 10 | 0.047% | -0.152% | 0.227% | 0.1%
1 | 11 | 0.024% | -0.023% | 0.09% | 0.1%
1 | 12 | 0.009% | -0.015% | 0.027% | 0.1%
1 | 13 | 0.03% | 0.021% | 0.021% | 0.1%

## 21. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:22:52 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DQ | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | 0.052% | PASSED | 2023-07-25 19:45
Yes | 1 | DQ | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.048% | PASSED | 2023-07-25 19:45
Yes | 2 | DQ | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.055% | PASSED | 2023-07-25 19:45
Yes | 3 | DQ | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.058% | PASSED | 2023-07-25 19:46
Yes | 4 | DQ | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.064% | PASSED | 2023-07-25 19:46
Yes | 5 | DQ | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | 0.049% | PASSED | 2023-07-25 19:46
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 1.667% | 0% | 0% | 0.1%
0 | 1 | 0.415% | -1.252% | 0% | 0.1%
0 | 2 | 0.049% | -0.366% | 0.71% | 0.1%
0 | 3 | 0.076% | 0.027% | 0.18% | 0.1%
0 | 4 | 0.044% | -0.032% | 0.056% | 0.1%
0 | 5 | 0.036% | -0.008% | 0.052% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0.053% | 0% | 0% | 0.1%
1 | 1 | 0.018% | -0.035% | 0% | 0.1%
1 | 2 | 0.058% | 0.04% | 0.043% | 0.1%
1 | 3 | 0.067% | 0.009% | 0.048% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | 0.08% | 0% | 0% | 0.1%
2 | 1 | 0.049% | -0.031% | 0% | 0.1%
2 | 2 | 0.085% | 0.036% | 0.071% | 0.1%
2 | 3 | 0.031% | -0.054% | 0.055% | 0.1%
Deviations of %error on test id #3. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.027% | 0% | 0% | 0.1%
3 | 1 | 0.089% | 0.062% | 0% | 0.1%
3 | 2 | 0.044% | -0.045% | 0.053% | 0.1%
3 | 3 | 0.044% | 0% | 0.059% | 0.1%
3 | 4 | 0.085% | 0.041% | 0.058% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.089% | 0% | 0% | 0.1%
4 | 1 | 0.022% | -0.067% | 0% | 0.1%
4 | 2 | 0.089% | 0.067% | 0.067% | 0.1%
4 | 3 | 0.08% | -0.009% | 0.064% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | 0.071% | 0% | 0% | 0.1%
5 | 1 | 0.058% | -0.013% | 0% | 0.1%
5 | 2 | 0.071% | 0.013% | 0.067% | 0.1%
5 | 3 | 0.018% | -0.053% | 0.049% | 0.1%

## 22. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:22:56 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DQ | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | 0.172% | PASSED | 2023-07-25 19:47
Yes | 1 | DQ | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.211% | PASSED | 2023-07-25 19:47
Yes | 2 | DQ | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | -0.098% | PASSED | 2023-07-25 19:48
Yes | 3 | DQ | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.187% | PASSED | 2023-07-25 19:48
Yes | 4 | DQ | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.178% | PASSED | 2023-07-25 19:48
Yes | 5 | DQ | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | 0.187% | PASSED | 2023-07-25 19:49
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -34.85% | 0% | 0% | 0.1%
0 | 1 | -24.65% | 10.19% | 0% | 0.1%
0 | 2 | -10.01% | 14.65% | -23.17% | 0.1%
0 | 3 | 0.142% | 10.15% | -11.51% | 0.1%
0 | 4 | 0.151% | 0.009% | -3.238% | 0.1%
0 | 5 | 0.178% | 0.027% | 0.157% | 0.1%
0 | 6 | 0.187% | 0.009% | 0.172% | 0.1%
Deviations of %error on test id #1. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -25.49% | 0% | 0% | 0.1%
1 | 1 | -13.88% | 11.6% | 0% | 0.1%
1 | 2 | -3.342% | 10.54% | -14.24% | 0.1%
1 | 3 | -0.169% | 3.173% | -5.799% | 0.1%
1 | 4 | -0.044% | 0.125% | -1.185% | 0.1%
1 | 5 | 0% | 0.044% | -0.071% | 0.1%
1 | 6 | -0.16% | -0.16% | -0.068% | 0.1%
1 | 7 | -0.186% | -0.026% | -0.115% | 0.1%
1 | 8 | -0.186% | 0% | -0.177% | 0.1%
1 | 9 | -0.027% | 0.159% | -0.133% | 0.1%
1 | 10 | 0.142% | 0.169% | -0.024% | 0.1%
1 | 11 | 0.187% | 0.045% | 0.101% | 0.1%
1 | 12 | 0.196% | 0.009% | 0.175% | 0.1%
1 | 13 | 0.25% | 0.054% | 0.211% | 0.1%
Deviations of %error on test id #2. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -25.46% | 0% | 0% | 0.1%
2 | 1 | -10.91% | 14.54% | 0% | 0.1%
2 | 2 | -4.231% | 6.681% | -13.53% | 0.1%
2 | 3 | 0.134% | 4.365% | -5.003% | 0.1%
2 | 4 | 0.062% | -0.072% | -1.345% | 0.1%
2 | 5 | -0.16% | -0.222% | 0.012% | 0.1%
2 | 6 | -0.071% | 0.089% | -0.056% | 0.1%
2 | 7 | -0.133% | -0.062% | -0.121% | 0.1%
2 | 8 | -0.089% | 0.044% | -0.098% | 0.1%
Deviations of %error on test id #3. Total 18 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -29.43% | 0% | 0% | 0.1%
3 | 1 | -21.82% | 7.611% | 0% | 0.1%
3 | 2 | -9.863% | 11.96% | -20.37% | 0.1%
3 | 3 | -1.978% | 7.885% | -11.22% | 0.1%
3 | 4 | 0.205% | 2.183% | -3.879% | 0.1%
3 | 5 | 0.053% | -0.152% | -0.573% | 0.1%
3 | 6 | -0.062% | -0.115% | 0.065% | 0.1%
3 | 7 | -0.169% | -0.107% | -0.059% | 0.1%
3 | 8 | -0.142% | 0.027% | -0.124% | 0.1%
3 | 9 | -0.142% | 0% | -0.151% | 0.1%
3 | 10 | -0.759% | -0.617% | -0.348% | 0.1%
3 | 11 | -0.574% | 0.185% | -0.492% | 0.1%
3 | 12 | -0.266% | 0.308% | -0.533% | 0.1%
3 | 13 | -0.053% | 0.213% | -0.298% | 0.1%
3 | 14 | 0.107% | 0.16% | -0.071% | 0.1%
3 | 15 | 0.196% | 0.089% | 0.083% | 0.1%
3 | 16 | 0.16% | -0.036% | 0.154% | 0.1%
3 | 17 | 0.205% | 0.045% | 0.187% | 0.1%
Deviations of %error on test id #4. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -25.72% | 0% | 0% | 0.1%
4 | 1 | -18.51% | 7.205% | 0% | 0.1%
4 | 2 | -3.417% | 15.1% | -15.88% | 0.1%
4 | 3 | -1.807% | 1.61% | -7.912% | 0.1%
4 | 4 | 0.196% | 2.003% | -1.676% | 0.1%
4 | 5 | 0.205% | 0.009% | -0.469% | 0.1%
4 | 6 | 0.169% | -0.036% | 0.19% | 0.1%
4 | 7 | 0.16% | -0.009% | 0.178% | 0.1%
Deviations of %error on test id #5. Total 18 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -30.1% | 0% | 0% | 0.1%
5 | 1 | -21.26% | 8.845% | 0% | 0.1%
5 | 2 | -15.36% | 5.895% | -22.24% | 0.1%
5 | 3 | -2.327% | 13.04% | -12.98% | 0.1%
5 | 4 | -1.419% | 0.908% | -6.369% | 0.1%
5 | 5 | 0.062% | 1.481% | -1.228% | 0.1%
5 | 6 | -0.16% | -0.222% | -0.506% | 0.1%
5 | 7 | -0.124% | 0.036% | -0.074% | 0.1%
5 | 8 | -0.044% | 0.08% | -0.109% | 0.1%
5 | 9 | -0.23% | -0.186% | -0.133% | 0.1%
5 | 10 | -0.222% | 0.008% | -0.165% | 0.1%
5 | 11 | -0.266% | -0.044% | -0.239% | 0.1%
5 | 12 | -0.16% | 0.106% | -0.216% | 0.1%
5 | 13 | -0.027% | 0.133% | -0.151% | 0.1%
5 | 14 | 0.116% | 0.143% | -0.024% | 0.1%
5 | 15 | 0.16% | 0.044% | 0.083% | 0.1%
5 | 16 | 0.205% | 0.045% | 0.16% | 0.1%
5 | 17 | 0.196% | -0.009% | 0.187% | 0.1%

## 23. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:23:02 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.017% | PASSED | 2023-07-25 19:49
Yes | 1 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.001% | PASSED | 2023-07-25 19:50
Yes | 2 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.043% | PASSED | 2023-07-25 19:50
Yes | 3 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.021% | PASSED | 2023-07-25 19:50
Yes | 4 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.018% | PASSED | 2023-07-25 19:51
Yes | 5 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.022% | PASSED | 2023-07-25 19:51
Yes | 6 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.04% | PASSED | 2023-07-25 19:51
Yes | 7 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.076% | PASSED | 2023-07-25 19:52
Yes | 8 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.05% | PASSED | 2023-07-25 19:52
Yes | 9 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.025% | PASSED | 2023-07-25 19:52
Yes | 10 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.076% | PASSED | 2023-07-25 19:52
Yes | 11 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.084% | PASSED | 2023-07-25 19:53
Yes | 12 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.068% | PASSED | 2023-07-25 19:53
Yes | 13 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.055% | PASSED | 2023-07-25 19:53
Yes | 14 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.044% | PASSED | 2023-07-25 19:54
Yes | 15 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.038% | PASSED | 2023-07-25 19:54
Yes | 16 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.065% | PASSED | 2023-07-25 19:54
Yes | 17 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.119% | PASSED | 2023-07-25 19:55
Yes | 18 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.111% | PASSED | 2023-07-25 19:55
Yes | 19 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.098% | PASSED | 2023-07-25 19:55
Yes | 20 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.1% | PASSED | 2023-07-25 19:55
Yes | 21 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.086% | PASSED | 2023-07-25 19:56
Yes | 22 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.094% | PASSED | 2023-07-25 19:56
Yes | 23 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.117% | PASSED | 2023-07-25 19:56
Yes | 24 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.117% | PASSED | 2023-07-25 19:57
Yes | 25 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.107% | PASSED | 2023-07-25 19:57
Yes | 26 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.094% | PASSED | 2023-07-25 19:57
Yes | 27 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.102% | PASSED | 2023-07-25 19:58
Yes | 28 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.139% | PASSED | 2023-07-25 19:58
Yes | 29 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.108% | PASSED | 2023-07-25 19:58
Yes | 30 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.126% | PASSED | 2023-07-25 19:58
Yes | 31 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.157% | PASSED | 2023-07-25 19:59
Yes | 32 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.123% | PASSED | 2023-07-25 19:59
Yes | 33 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.103% | PASSED | 2023-07-25 19:59
Yes | 34 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.114% | PASSED | 2023-07-25 20:00
Yes | 35 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.101% | PASSED | 2023-07-25 20:00
Yes | 36 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.126% | PASSED | 2023-07-25 20:00
Yes | 37 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.156% | PASSED | 2023-07-25 20:01
Yes | 38 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.141% | PASSED | 2023-07-25 20:01
Yes | 39 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.139% | PASSED | 2023-07-25 20:01
Yes | 40 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.147% | PASSED | 2023-07-25 20:02
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -0.031% | 0% | 0% | 0.1%
0 | 1 | -0.004% | 0.027% | 0% | 0.1%
0 | 2 | -0.004% | 0% | -0.013% | 0.1%
0 | 3 | -0.044% | -0.04% | -0.017% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0% | 0% | 0% | 0.1%
1 | 1 | 0.004% | 0.004% | 0% | 0.1%
1 | 2 | 0.004% | 0% | 0.003% | 0.1%
1 | 3 | -0.004% | -0.008% | 0.001% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | 0.044% | 0% | 0% | 0.1%
2 | 1 | 0.031% | -0.013% | 0% | 0.1%
2 | 2 | 0.049% | 0.018% | 0.041% | 0.1%
2 | 3 | 0.049% | 0% | 0.043% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.018% | 0% | 0% | 0.1%
3 | 1 | 0.009% | -0.009% | 0% | 0.1%
3 | 2 | 0.031% | 0.022% | 0.019% | 0.1%
3 | 3 | 0.022% | -0.009% | 0.021% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.004% | 0% | 0% | 0.1%
4 | 1 | 0.049% | 0.045% | 0% | 0.1%
4 | 2 | -0.004% | -0.053% | 0.016% | 0.1%
4 | 3 | 0.009% | 0.013% | 0.018% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | 0.018% | 0% | 0% | 0.1%
5 | 1 | 0.004% | -0.014% | 0% | 0.1%
5 | 2 | 0.009% | 0.005% | 0.01% | 0.1%
5 | 3 | 0.053% | 0.044% | 0.022% | 0.1%
Deviations of %error on test id #6. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | 0.044% | 0% | 0% | 0.1%
6 | 1 | 0.04% | -0.004% | 0% | 0.1%
6 | 2 | 0.044% | 0.004% | 0.043% | 0.1%
6 | 3 | 0.036% | -0.008% | 0.04% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.049% | 0% | 0% | 0.1%
7 | 1 | 0.089% | 0.04% | 0% | 0.1%
7 | 2 | 0.058% | -0.031% | 0.065% | 0.1%
7 | 3 | 0.08% | 0.022% | 0.076% | 0.1%
Deviations of %error on test id #8. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | 0.04% | 0% | 0% | 0.1%
8 | 1 | 0.036% | -0.004% | 0% | 0.1%
8 | 2 | 0.071% | 0.035% | 0.049% | 0.1%
8 | 3 | 0.044% | -0.027% | 0.05% | 0.1%
Deviations of %error on test id #9. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | 0.04% | 0% | 0% | 0.1%
9 | 1 | 0.044% | 0.004% | 0% | 0.1%
9 | 2 | 0% | -0.044% | 0.028% | 0.1%
9 | 3 | 0.031% | 0.031% | 0.025% | 0.1%
Deviations of %error on test id #10. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | 0.062% | 0% | 0% | 0.1%
10 | 1 | 0.085% | 0.023% | 0% | 0.1%
10 | 2 | 0.049% | -0.036% | 0.065% | 0.1%
10 | 3 | 0.093% | 0.044% | 0.076% | 0.1%
Deviations of %error on test id #11. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | 0.085% | 0% | 0% | 0.1%
11 | 1 | 0.093% | 0.008% | 0% | 0.1%
11 | 2 | 0.093% | 0% | 0.09% | 0.1%
11 | 3 | 0.067% | -0.026% | 0.084% | 0.1%
Deviations of %error on test id #12. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | 0.085% | 0% | 0% | 0.1%
12 | 1 | 0.076% | -0.009% | 0% | 0.1%
12 | 2 | 0.044% | -0.032% | 0.068% | 0.1%
12 | 3 | 0.085% | 0.041% | 0.068% | 0.1%
Deviations of %error on test id #13. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | 0.08% | 0% | 0% | 0.1%
13 | 1 | 0.044% | -0.036% | 0% | 0.1%
13 | 2 | 0.071% | 0.027% | 0.065% | 0.1%
13 | 3 | 0.049% | -0.022% | 0.055% | 0.1%
Deviations of %error on test id #14. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | 0.04% | 0% | 0% | 0.1%
14 | 1 | 0.022% | -0.018% | 0% | 0.1%
14 | 2 | 0.071% | 0.049% | 0.044% | 0.1%
14 | 3 | 0.04% | -0.031% | 0.044% | 0.1%
Deviations of %error on test id #15. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | 0.044% | 0% | 0% | 0.1%
15 | 1 | 0.044% | 0% | 0% | 0.1%
15 | 2 | 0.04% | -0.004% | 0.043% | 0.1%
15 | 3 | 0.031% | -0.009% | 0.038% | 0.1%
Deviations of %error on test id #16. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | 0.089% | 0% | 0% | 0.1%
16 | 1 | 0.044% | -0.045% | 0% | 0.1%
16 | 2 | 0.111% | 0.067% | 0.081% | 0.1%
16 | 3 | 0.04% | -0.071% | 0.065% | 0.1%
Deviations of %error on test id #17. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | 0.116% | 0% | 0% | 0.1%
17 | 1 | 0.12% | 0.004% | 0% | 0.1%
17 | 2 | 0.129% | 0.009% | 0.122% | 0.1%
17 | 3 | 0.107% | -0.022% | 0.119% | 0.1%
Deviations of %error on test id #18. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | 0.125% | 0% | 0% | 0.1%
18 | 1 | 0.085% | -0.04% | 0% | 0.1%
18 | 2 | 0.151% | 0.066% | 0.12% | 0.1%
18 | 3 | 0.098% | -0.053% | 0.111% | 0.1%
Deviations of %error on test id #19. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | 0.089% | 0% | 0% | 0.1%
19 | 1 | 0.116% | 0.027% | 0% | 0.1%
19 | 2 | 0.093% | -0.023% | 0.099% | 0.1%
19 | 3 | 0.085% | -0.008% | 0.098% | 0.1%
Deviations of %error on test id #20. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | 0.076% | 0% | 0% | 0.1%
20 | 1 | 0.085% | 0.009% | 0% | 0.1%
20 | 2 | 0.125% | 0.04% | 0.095% | 0.1%
20 | 3 | 0.089% | -0.036% | 0.1% | 0.1%
Deviations of %error on test id #21. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | 0.049% | 0% | 0% | 0.1%
21 | 1 | 0.089% | 0.04% | 0% | 0.1%
21 | 2 | 0.089% | 0% | 0.076% | 0.1%
21 | 3 | 0.08% | -0.009% | 0.086% | 0.1%
Deviations of %error on test id #22. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | 0.085% | 0% | 0% | 0.1%
22 | 1 | 0.085% | 0% | 0% | 0.1%
22 | 2 | 0.125% | 0.04% | 0.098% | 0.1%
22 | 3 | 0.071% | -0.054% | 0.094% | 0.1%
Deviations of %error on test id #23. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | 0.111% | 0% | 0% | 0.1%
23 | 1 | 0.12% | 0.009% | 0% | 0.1%
23 | 2 | 0.12% | 0% | 0.117% | 0.1%
23 | 3 | 0.111% | -0.009% | 0.117% | 0.1%
Deviations of %error on test id #24. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | 0.116% | 0% | 0% | 0.1%
24 | 1 | 0.125% | 0.009% | 0% | 0.1%
24 | 2 | 0.107% | -0.018% | 0.116% | 0.1%
24 | 3 | 0.12% | 0.013% | 0.117% | 0.1%
Deviations of %error on test id #25. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | 0.093% | 0% | 0% | 0.1%
25 | 1 | 0.129% | 0.036% | 0% | 0.1%
25 | 2 | 0.098% | -0.031% | 0.107% | 0.1%
25 | 3 | 0.093% | -0.005% | 0.107% | 0.1%
Deviations of %error on test id #26. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | 0.134% | 0% | 0% | 0.1%
26 | 1 | 0.076% | -0.058% | 0% | 0.1%
26 | 2 | 0.12% | 0.044% | 0.11% | 0.1%
26 | 3 | 0.085% | -0.035% | 0.094% | 0.1%
Deviations of %error on test id #27. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | 0.134% | 0% | 0% | 0.1%
27 | 1 | 0.08% | -0.054% | 0% | 0.1%
27 | 2 | 0.093% | 0.013% | 0.102% | 0.1%
27 | 3 | 0.134% | 0.041% | 0.102% | 0.1%
Deviations of %error on test id #28. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | 0.085% | 0% | 0% | 0.1%
28 | 1 | 0.169% | 0.084% | 0% | 0.1%
28 | 2 | 0.089% | -0.08% | 0.114% | 0.1%
28 | 3 | 0.16% | 0.071% | 0.139% | 0.1%
Deviations of %error on test id #29. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | 0.129% | 0% | 0% | 0.1%
29 | 1 | 0.093% | -0.036% | 0% | 0.1%
29 | 2 | 0.134% | 0.041% | 0.119% | 0.1%
29 | 3 | 0.098% | -0.036% | 0.108% | 0.1%
Deviations of %error on test id #30. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | 0.12% | 0% | 0% | 0.1%
30 | 1 | 0.125% | 0.005% | 0% | 0.1%
30 | 2 | 0.129% | 0.004% | 0.125% | 0.1%
30 | 3 | 0.125% | -0.004% | 0.126% | 0.1%
Deviations of %error on test id #31. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | 0.12% | 0% | 0% | 0.1%
31 | 1 | 0.169% | 0.049% | 0% | 0.1%
31 | 2 | 0.129% | -0.04% | 0.139% | 0.1%
31 | 3 | 0.174% | 0.045% | 0.157% | 0.1%
Deviations of %error on test id #32. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | 0.116% | 0% | 0% | 0.1%
32 | 1 | 0.129% | 0.013% | 0% | 0.1%
32 | 2 | 0.116% | -0.013% | 0.12% | 0.1%
32 | 3 | 0.125% | 0.009% | 0.123% | 0.1%
Deviations of %error on test id #33. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | 0.089% | 0% | 0% | 0.1%
33 | 1 | 0.134% | 0.045% | 0% | 0.1%
33 | 2 | 0.085% | -0.049% | 0.103% | 0.1%
33 | 3 | 0.089% | 0.004% | 0.103% | 0.1%
Deviations of %error on test id #34. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | 0.08% | 0% | 0% | 0.1%
34 | 1 | 0.116% | 0.036% | 0% | 0.1%
34 | 2 | 0.12% | 0.004% | 0.105% | 0.1%
34 | 3 | 0.107% | -0.013% | 0.114% | 0.1%
Deviations of %error on test id #35. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | 0.08% | 0% | 0% | 0.1%
35 | 1 | 0.125% | 0.045% | 0% | 0.1%
35 | 2 | 0.085% | -0.04% | 0.097% | 0.1%
35 | 3 | 0.093% | 0.008% | 0.101% | 0.1%
Deviations of %error on test id #36. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | 0.111% | 0% | 0% | 0.1%
36 | 1 | 0.125% | 0.014% | 0% | 0.1%
36 | 2 | 0.129% | 0.004% | 0.122% | 0.1%
36 | 3 | 0.125% | -0.004% | 0.126% | 0.1%
Deviations of %error on test id #37. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | 0.165% | 0% | 0% | 0.1%
37 | 1 | 0.169% | 0.004% | 0% | 0.1%
37 | 2 | 0.134% | -0.035% | 0.156% | 0.1%
37 | 3 | 0.165% | 0.031% | 0.156% | 0.1%
Deviations of %error on test id #38. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | 0.125% | 0% | 0% | 0.1%
38 | 1 | 0.151% | 0.026% | 0% | 0.1%
38 | 2 | 0.12% | -0.031% | 0.132% | 0.1%
38 | 3 | 0.151% | 0.031% | 0.141% | 0.1%
Deviations of %error on test id #39. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | 0.151% | 0% | 0% | 0.1%
39 | 1 | 0.129% | -0.022% | 0% | 0.1%
39 | 2 | 0.16% | 0.031% | 0.147% | 0.1%
39 | 3 | 0.129% | -0.031% | 0.139% | 0.1%
Deviations of %error on test id #40. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | 0.134% | 0% | 0% | 0.1%
40 | 1 | 0.138% | 0.004% | 0% | 0.1%
40 | 2 | 0.174% | 0.036% | 0.149% | 0.1%
40 | 3 | 0.129% | -0.045% | 0.147% | 0.1%

## 24. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:23:23 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.163% | PASSED | 2023-07-25 20:02
Yes | 1 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.077% | PASSED | 2023-07-25 20:02
Yes | 2 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.113% | PASSED | 2023-07-25 20:03
Yes | 3 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.077% | PASSED | 2023-07-25 20:03
Yes | 4 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.074% | PASSED | 2023-07-25 20:04
Yes | 5 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.101% | PASSED | 2023-07-25 20:04
Yes | 6 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.145% | PASSED | 2023-07-25 20:05
Yes | 7 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.149% | PASSED | 2023-07-25 20:05
Yes | 8 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.092% | PASSED | 2023-07-25 20:06
Yes | 9 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.116% | PASSED | 2023-07-25 20:06
Yes | 10 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.184% | PASSED | 2023-07-25 20:06
Yes | 11 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.172% | PASSED | 2023-07-25 20:07
Yes | 12 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.184% | PASSED | 2023-07-25 20:07
Yes | 13 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.05% | PASSED | 2023-07-25 20:08
Yes | 14 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.238% | PASSED | 2023-07-25 20:08
Yes | 15 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.211% | PASSED | 2023-07-25 20:09
Yes | 16 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.199% | PASSED | 2023-07-25 20:09
Yes | 17 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.036% | PASSED | 2023-07-25 20:09
Yes | 18 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.306% | PASSED | 2023-07-25 20:10
Yes | 19 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.044% | PASSED | 2023-07-25 20:10
Yes | 20 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.255% | PASSED | 2023-07-25 20:11
Yes | 21 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.309% | PASSED | 2023-07-25 20:11
Yes | 22 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.258% | PASSED | 2023-07-25 20:12
Yes | 23 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.095% | PASSED | 2023-07-25 20:12
Yes | 24 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.366% | PASSED | 2023-07-25 20:13
Yes | 25 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.345% | PASSED | 2023-07-25 20:13
Yes | 26 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.089% | PASSED | 2023-07-25 20:13
Yes | 27 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.363% | PASSED | 2023-07-25 20:14
Yes | 28 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.3% | PASSED | 2023-07-25 20:14
Yes | 29 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.33% | PASSED | 2023-07-25 20:15
Yes | 30 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.402% | PASSED | 2023-07-25 20:15
Yes | 31 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.3% | PASSED | 2023-07-25 20:16
Yes | 32 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.044% | PASSED | 2023-07-25 20:16
Yes | 33 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.333% | PASSED | 2023-07-25 20:17
Yes | 34 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.339% | PASSED | 2023-07-25 20:17
Yes | 35 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.336% | PASSED | 2023-07-25 20:18
Yes | 36 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.366% | PASSED | 2023-07-25 20:18
Yes | 37 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.426% | PASSED | 2023-07-25 20:19
Yes | 38 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.375% | PASSED | 2023-07-25 20:19
Yes | 39 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.434% | PASSED | 2023-07-25 20:20
Yes | 40 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.42% | PASSED | 2023-07-25 20:20
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -27.88% | 0% | 0% | 0.1%
0 | 1 | -21.52% | 6.364% | 0% | 0.1%
0 | 2 | -9.106% | 12.41% | -19.5% | 0.1%
0 | 3 | -2.623% | 6.483% | -11.08% | 0.1%
0 | 4 | -0.16% | 2.463% | -3.963% | 0.1%
0 | 5 | -0.239% | -0.079% | -1.007% | 0.1%
0 | 6 | -0.16% | 0.079% | -0.186% | 0.1%
0 | 7 | -0.089% | 0.071% | -0.163% | 0.1%
Deviations of %error on test id #1. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -24.53% | 0% | 0% | 0.1%
1 | 1 | -14.4% | 10.14% | 0% | 0.1%
1 | 2 | -1.721% | 12.68% | -13.55% | 0.1%
1 | 3 | 0.125% | 1.846% | -5.331% | 0.1%
1 | 4 | 0.071% | -0.054% | -0.508% | 0.1%
1 | 5 | 0.116% | 0.045% | 0.104% | 0.1%
1 | 6 | 0.062% | -0.054% | 0.083% | 0.1%
1 | 7 | 0.053% | -0.009% | 0.077% | 0.1%
Deviations of %error on test id #2. Total 19 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -32.48% | 0% | 0% | 0.1%
2 | 1 | -27.4% | 5.086% | 0% | 0.1%
2 | 2 | -19.34% | 8.052% | -26.41% | 0.1%
2 | 3 | -4.157% | 15.19% | -16.96% | 0.1%
2 | 4 | -1.935% | 2.222% | -8.478% | 0.1%
2 | 5 | 0.125% | 2.06% | -1.989% | 0.1%
2 | 6 | 0.062% | -0.063% | -0.583% | 0.1%
2 | 7 | -0.106% | -0.168% | 0.027% | 0.1%
2 | 8 | -0.151% | -0.045% | -0.065% | 0.1%
2 | 9 | -0.053% | 0.098% | -0.103% | 0.1%
2 | 10 | -0.204% | -0.151% | -0.136% | 0.1%
2 | 11 | -0.548% | -0.344% | -0.268% | 0.1%
2 | 12 | -0.566% | -0.018% | -0.439% | 0.1%
2 | 13 | -0.204% | 0.362% | -0.439% | 0.1%
2 | 14 | -0.062% | 0.142% | -0.277% | 0.1%
2 | 15 | 0.098% | 0.16% | -0.056% | 0.1%
2 | 16 | 0.062% | -0.036% | 0.033% | 0.1%
2 | 17 | 0.116% | 0.054% | 0.092% | 0.1%
2 | 18 | 0.16% | 0.044% | 0.113% | 0.1%
Deviations of %error on test id #3. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -39.07% | 0% | 0% | 0.1%
3 | 1 | -25.44% | 13.63% | 0% | 0.1%
3 | 2 | -17.29% | 8.157% | -27.27% | 0.1%
3 | 3 | -8.729% | 8.556% | -17.15% | 0.1%
3 | 4 | 0.044% | 8.773% | -8.657% | 0.1%
3 | 5 | 0.053% | 0.009% | -2.877% | 0.1%
3 | 6 | 0.08% | 0.027% | 0.059% | 0.1%
3 | 7 | 0.098% | 0.018% | 0.077% | 0.1%
Deviations of %error on test id #4. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -34.8% | 0% | 0% | 0.1%
4 | 1 | -27.71% | 7.097% | 0% | 0.1%
4 | 2 | -12.79% | 14.92% | -25.1% | 0.1%
4 | 3 | -5.375% | 7.415% | -15.29% | 0.1%
4 | 4 | 0.062% | 5.437% | -6.034% | 0.1%
4 | 5 | -0.018% | -0.08% | -1.777% | 0.1%
4 | 6 | -0.248% | -0.23% | -0.068% | 0.1%
4 | 7 | -0.169% | 0.079% | -0.145% | 0.1%
4 | 8 | -0.036% | 0.133% | -0.151% | 0.1%
4 | 9 | -0.248% | -0.212% | -0.151% | 0.1%
4 | 10 | -0.204% | 0.044% | -0.163% | 0.1%
4 | 11 | -0.222% | -0.018% | -0.225% | 0.1%
4 | 12 | -0.601% | -0.379% | -0.342% | 0.1%
4 | 13 | -0.539% | 0.062% | -0.454% | 0.1%
4 | 14 | -0.266% | 0.273% | -0.469% | 0.1%
4 | 15 | -0.098% | 0.168% | -0.301% | 0.1%
4 | 16 | 0.018% | 0.116% | -0.115% | 0.1%
4 | 17 | 0.062% | 0.044% | -0.006% | 0.1%
4 | 18 | 0.08% | 0.018% | 0.053% | 0.1%
4 | 19 | 0.08% | 0% | 0.074% | 0.1%
Deviations of %error on test id #5. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -23.91% | 0% | 0% | 0.1%
5 | 1 | -17.27% | 6.647% | 0% | 0.1%
5 | 2 | -8.544% | 8.723% | -16.57% | 0.1%
5 | 3 | -1.892% | 6.652% | -9.234% | 0.1%
5 | 4 | -0.16% | 1.732% | -3.532% | 0.1%
5 | 5 | -0.186% | -0.026% | -0.746% | 0.1%
5 | 6 | -0.036% | 0.15% | -0.127% | 0.1%
5 | 7 | -0.16% | -0.124% | -0.127% | 0.1%
5 | 8 | -0.319% | -0.159% | -0.172% | 0.1%
5 | 9 | -0.31% | 0.009% | -0.263% | 0.1%
5 | 10 | -0.186% | 0.124% | -0.272% | 0.1%
5 | 11 | -0.08% | 0.106% | -0.192% | 0.1%
5 | 12 | 0.08% | 0.16% | -0.062% | 0.1%
5 | 13 | 0.071% | -0.009% | 0.024% | 0.1%
5 | 14 | 0.116% | 0.045% | 0.089% | 0.1%
5 | 15 | 0.116% | 0% | 0.101% | 0.1%
Deviations of %error on test id #6. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | -34.02% | 0% | 0% | 0.1%
6 | 1 | -26.14% | 7.88% | 0% | 0.1%
6 | 2 | -17.15% | 8.986% | -25.77% | 0.1%
6 | 3 | -2.656% | 14.49% | -15.32% | 0.1%
6 | 4 | -1.359% | 1.297% | -7.055% | 0.1%
6 | 5 | -0.16% | 1.199% | -1.392% | 0.1%
6 | 6 | -0.106% | 0.054% | -0.542% | 0.1%
6 | 7 | -0.177% | -0.071% | -0.148% | 0.1%
6 | 8 | -0.151% | 0.026% | -0.145% | 0.1%
Deviations of %error on test id #7. Total 19 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | -30.99% | 0% | 0% | 0.1%
7 | 1 | -22.24% | 8.748% | 0% | 0.1%
7 | 2 | -7.84% | 14.4% | -20.36% | 0.1%
7 | 3 | -3.117% | 4.723% | -11.07% | 0.1%
7 | 4 | 0.187% | 3.304% | -3.59% | 0.1%
7 | 5 | 0.151% | -0.036% | -0.926% | 0.1%
7 | 6 | 0.16% | 0.009% | 0.166% | 0.1%
7 | 7 | 0.027% | -0.133% | 0.113% | 0.1%
7 | 8 | -0.248% | -0.275% | -0.02% | 0.1%
7 | 9 | -0.266% | -0.018% | -0.162% | 0.1%
7 | 10 | -0.151% | 0.115% | -0.222% | 0.1%
7 | 11 | -0.248% | -0.097% | -0.222% | 0.1%
7 | 12 | -0.513% | -0.265% | -0.304% | 0.1%
7 | 13 | -0.381% | 0.132% | -0.381% | 0.1%
7 | 14 | -0.16% | 0.221% | -0.351% | 0.1%
7 | 15 | 0.036% | 0.196% | -0.168% | 0.1%
7 | 16 | 0.134% | 0.098% | 0.003% | 0.1%
7 | 17 | 0.196% | 0.062% | 0.122% | 0.1%
7 | 18 | 0.116% | -0.08% | 0.149% | 0.1%
Deviations of %error on test id #8. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | -30.51% | 0% | 0% | 0.1%
8 | 1 | -16.52% | 13.98% | 0% | 0.1%
8 | 2 | -7.757% | 8.767% | -18.26% | 0.1%
8 | 3 | -1.986% | 5.771% | -8.756% | 0.1%
8 | 4 | 0.134% | 2.12% | -3.203% | 0.1%
8 | 5 | 0.107% | -0.027% | -0.582% | 0.1%
8 | 6 | 0.125% | 0.018% | 0.122% | 0.1%
8 | 7 | 0.044% | -0.081% | 0.092% | 0.1%
Deviations of %error on test id #9. Total 19 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | -33.55% | 0% | 0% | 0.1%
9 | 1 | -24.89% | 8.664% | 0% | 0.1%
9 | 2 | -9.617% | 15.27% | -22.69% | 0.1%
9 | 3 | -3.458% | 6.159% | -12.65% | 0.1%
9 | 4 | 0.178% | 3.636% | -4.299% | 0.1%
9 | 5 | 0.16% | -0.018% | -1.04% | 0.1%
9 | 6 | 0.178% | 0.018% | 0.172% | 0.1%
9 | 7 | 0.062% | -0.116% | 0.133% | 0.1%
9 | 8 | -0.195% | -0.257% | 0.015% | 0.1%
9 | 9 | -0.301% | -0.106% | -0.145% | 0.1%
9 | 10 | -0.151% | 0.15% | -0.216% | 0.1%
9 | 11 | -0.504% | -0.353% | -0.319% | 0.1%
9 | 12 | -0.522% | -0.018% | -0.392% | 0.1%
9 | 13 | -0.301% | 0.221% | -0.442% | 0.1%
9 | 14 | -0.169% | 0.132% | -0.331% | 0.1%
9 | 15 | 0.044% | 0.213% | -0.142% | 0.1%
9 | 16 | 0.107% | 0.063% | -0.006% | 0.1%
9 | 17 | 0.125% | 0.018% | 0.092% | 0.1%
9 | 18 | 0.116% | -0.009% | 0.116% | 0.1%
Deviations of %error on test id #10. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | -21.72% | 0% | 0% | 0.1%
10 | 1 | -14.04% | 7.681% | 0% | 0.1%
10 | 2 | -2.665% | 11.37% | -12.81% | 0.1%
10 | 3 | -0.478% | 2.187% | -5.726% | 0.1%
10 | 4 | -0.151% | 0.327% | -1.098% | 0.1%
10 | 5 | -0.195% | -0.044% | -0.275% | 0.1%
10 | 6 | -0.018% | 0.177% | -0.121% | 0.1%
10 | 7 | -0.106% | -0.088% | -0.106% | 0.1%
10 | 8 | -0.337% | -0.231% | -0.154% | 0.1%
10 | 9 | -0.23% | 0.107% | -0.224% | 0.1%
10 | 10 | -0.062% | 0.168% | -0.21% | 0.1%
10 | 11 | 0.018% | 0.08% | -0.091% | 0.1%
10 | 12 | 0.196% | 0.178% | 0.051% | 0.1%
10 | 13 | 0.187% | -0.009% | 0.134% | 0.1%
10 | 14 | 0.187% | 0% | 0.19% | 0.1%
10 | 15 | 0.178% | -0.009% | 0.184% | 0.1%
Deviations of %error on test id #11. Total 19 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | -31.57% | 0% | 0% | 0.1%
11 | 1 | -22.49% | 9.08% | 0% | 0.1%
11 | 2 | -9.755% | 12.74% | -21.27% | 0.1%
11 | 3 | -2.521% | 7.234% | -11.59% | 0.1%
11 | 4 | 0.16% | 2.681% | -4.039% | 0.1%
11 | 5 | 0.071% | -0.089% | -0.763% | 0.1%
11 | 6 | -0.062% | -0.133% | 0.056% | 0.1%
11 | 7 | -0.053% | 0.009% | -0.015% | 0.1%
11 | 8 | -0.195% | -0.142% | -0.103% | 0.1%
11 | 9 | -0.337% | -0.142% | -0.195% | 0.1%
11 | 10 | -0.732% | -0.395% | -0.421% | 0.1%
11 | 11 | -0.486% | 0.246% | -0.518% | 0.1%
11 | 12 | -0.213% | 0.273% | -0.477% | 0.1%
11 | 13 | 0.116% | 0.329% | -0.194% | 0.1%
11 | 14 | 0.098% | -0.018% | 0% | 0.1%
11 | 15 | 0.205% | 0.107% | 0.14% | 0.1%
11 | 16 | 0.151% | -0.054% | 0.151% | 0.1%
11 | 17 | 0.169% | 0.018% | 0.175% | 0.1%
11 | 18 | 0.196% | 0.027% | 0.172% | 0.1%
Deviations of %error on test id #12. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | -20.56% | 0% | 0% | 0.1%
12 | 1 | -9.508% | 11.05% | 0% | 0.1%
12 | 2 | -3.276% | 6.232% | -11.11% | 0.1%
12 | 3 | 0.196% | 3.472% | -4.196% | 0.1%
12 | 4 | 0.205% | 0.009% | -0.958% | 0.1%
12 | 5 | -0.009% | -0.214% | 0.131% | 0.1%
12 | 6 | -0.062% | -0.053% | 0.045% | 0.1%
12 | 7 | 0% | 0.062% | -0.024% | 0.1%
12 | 8 | -0.23% | -0.23% | -0.097% | 0.1%
12 | 9 | -0.609% | -0.379% | -0.28% | 0.1%
12 | 10 | -0.601% | 0.008% | -0.48% | 0.1%
12 | 11 | -0.239% | 0.362% | -0.483% | 0.1%
12 | 12 | 0% | 0.239% | -0.28% | 0.1%
12 | 13 | 0.116% | 0.116% | -0.041% | 0.1%
12 | 14 | 0.187% | 0.071% | 0.101% | 0.1%
12 | 15 | 0.169% | -0.018% | 0.157% | 0.1%
12 | 16 | 0.196% | 0.027% | 0.184% | 0.1%
Deviations of %error on test id #13. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | -35.21% | 0% | 0% | 0.1%
13 | 1 | -23.02% | 12.19% | 0% | 0.1%
13 | 2 | -10.29% | 12.72% | -22.84% | 0.1%
13 | 3 | -4.214% | 6.08% | -12.51% | 0.1%
13 | 4 | -0.062% | 4.152% | -4.857% | 0.1%
13 | 5 | -0.169% | -0.107% | -1.482% | 0.1%
13 | 6 | -0.062% | 0.107% | -0.098% | 0.1%
13 | 7 | -0.018% | 0.044% | -0.083% | 0.1%
13 | 8 | -0.018% | 0% | -0.033% | 0.1%
13 | 9 | -0.115% | -0.097% | -0.05% | 0.1%
Deviations of %error on test id #14. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | -20.01% | 0% | 0% | 0.1%
14 | 1 | -10.93% | 9.088% | 0% | 0.1%
14 | 2 | -2.833% | 8.093% | -11.26% | 0.1%
14 | 3 | 0.205% | 3.038% | -4.518% | 0.1%
14 | 4 | 0.062% | -0.143% | -0.855% | 0.1%
14 | 5 | -0.053% | -0.115% | 0.071% | 0.1%
14 | 6 | 0.062% | 0.115% | 0.024% | 0.1%
14 | 7 | 0.187% | 0.125% | 0.065% | 0.1%
14 | 8 | -0.133% | -0.32% | 0.039% | 0.1%
14 | 9 | -0.53% | -0.397% | -0.159% | 0.1%
14 | 10 | -0.557% | -0.027% | -0.407% | 0.1%
14 | 11 | -0.177% | 0.38% | -0.421% | 0.1%
14 | 12 | 0% | 0.177% | -0.245% | 0.1%
14 | 13 | 0.178% | 0.178% | 0% | 0.1%
14 | 14 | 0.232% | 0.054% | 0.137% | 0.1%
14 | 15 | 0.241% | 0.009% | 0.217% | 0.1%
14 | 16 | 0.241% | 0% | 0.238% | 0.1%
Deviations of %error on test id #15. Total 18 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | -30.74% | 0% | 0% | 0.1%
15 | 1 | -20.62% | 10.12% | 0% | 0.1%
15 | 2 | -5.802% | 14.82% | -19.06% | 0.1%
15 | 3 | -1.798% | 4.004% | -9.408% | 0.1%
15 | 4 | 0.196% | 1.994% | -2.468% | 0.1%
15 | 5 | 0.044% | -0.152% | -0.519% | 0.1%
15 | 6 | -0.124% | -0.168% | 0.039% | 0.1%
15 | 7 | 0.036% | 0.16% | -0.015% | 0.1%
15 | 8 | -0.018% | -0.054% | -0.035% | 0.1%
15 | 9 | -0.478% | -0.46% | -0.153% | 0.1%
15 | 10 | -0.627% | -0.149% | -0.374% | 0.1%
15 | 11 | -0.372% | 0.255% | -0.492% | 0.1%
15 | 12 | -0.151% | 0.221% | -0.383% | 0.1%
15 | 13 | 0.107% | 0.258% | -0.139% | 0.1%
15 | 14 | 0.214% | 0.107% | 0.057% | 0.1%
15 | 15 | 0.214% | 0% | 0.178% | 0.1%
15 | 16 | 0.223% | 0.009% | 0.217% | 0.1%
15 | 17 | 0.196% | -0.027% | 0.211% | 0.1%
Deviations of %error on test id #16. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | -31.08% | 0% | 0% | 0.1%
16 | 1 | -21.54% | 9.546% | 0% | 0.1%
16 | 2 | -10.92% | 10.62% | -21.18% | 0.1%
16 | 3 | 0.258% | 11.18% | -10.73% | 0.1%
16 | 4 | 0.214% | -0.044% | -3.482% | 0.1%
16 | 5 | 0.178% | -0.036% | 0.217% | 0.1%
16 | 6 | 0.205% | 0.027% | 0.199% | 0.1%
Deviations of %error on test id #17. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | -31.43% | 0% | 0% | 0.1%
17 | 1 | -22.05% | 9.378% | 0% | 0.1%
17 | 2 | -8.833% | 13.22% | -20.77% | 0.1%
17 | 3 | -1.815% | 7.018% | -10.9% | 0.1%
17 | 4 | 0.294% | 2.109% | -3.451% | 0.1%
17 | 5 | 0.062% | -0.232% | -0.486% | 0.1%
17 | 6 | 0% | -0.062% | 0.119% | 0.1%
17 | 7 | 0.071% | 0.071% | 0.044% | 0.1%
17 | 8 | 0.036% | -0.035% | 0.036% | 0.1%
Deviations of %error on test id #18. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | -25.15% | 0% | 0% | 0.1%
18 | 1 | -14.09% | 11.06% | 0% | 0.1%
18 | 2 | -4.32% | 9.769% | -14.52% | 0.1%
18 | 3 | -1.454% | 2.866% | -6.621% | 0.1%
18 | 4 | -0.018% | 1.436% | -1.931% | 0.1%
18 | 5 | 0.098% | 0.116% | -0.458% | 0.1%
18 | 6 | 0.241% | 0.143% | 0.107% | 0.1%
18 | 7 | 0.107% | -0.134% | 0.149% | 0.1%
18 | 8 | -0.133% | -0.24% | 0.072% | 0.1%
18 | 9 | -0.115% | 0.018% | -0.047% | 0.1%
18 | 10 | 0.142% | 0.257% | -0.035% | 0.1%
18 | 11 | 0.25% | 0.108% | 0.092% | 0.1%
18 | 12 | 0.321% | 0.071% | 0.238% | 0.1%
18 | 13 | 0.294% | -0.027% | 0.288% | 0.1%
18 | 14 | 0.303% | 0.009% | 0.306% | 0.1%
Deviations of %error on test id #19. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | -27.07% | 0% | 0% | 0.1%
19 | 1 | -17.76% | 9.309% | 0% | 0.1%
19 | 2 | -7.163% | 10.59% | -17.33% | 0.1%
19 | 3 | 0.205% | 7.368% | -8.238% | 0.1%
19 | 4 | 0.053% | -0.152% | -2.302% | 0.1%
19 | 5 | 0.08% | 0.027% | 0.113% | 0.1%
19 | 6 | 0.089% | 0.009% | 0.074% | 0.1%
19 | 7 | 0.062% | -0.027% | 0.077% | 0.1%
19 | 8 | -0.018% | -0.08% | 0.044% | 0.1%
Deviations of %error on test id #20. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | -26.02% | 0% | 0% | 0.1%
20 | 1 | -15.26% | 10.76% | 0% | 0.1%
20 | 2 | -2.942% | 12.32% | -14.74% | 0.1%
20 | 3 | -0.82% | 2.122% | -6.341% | 0.1%
20 | 4 | 0.16% | 0.98% | -1.201% | 0.1%
20 | 5 | -0.027% | -0.187% | -0.229% | 0.1%
20 | 6 | 0.062% | 0.089% | 0.065% | 0.1%
20 | 7 | 0.169% | 0.107% | 0.068% | 0.1%
20 | 8 | -0.036% | -0.205% | 0.065% | 0.1%
20 | 9 | -0.062% | -0.026% | 0.024% | 0.1%
20 | 10 | -0.486% | -0.424% | -0.195% | 0.1%
20 | 11 | -0.301% | 0.185% | -0.283% | 0.1%
20 | 12 | 0.018% | 0.319% | -0.256% | 0.1%
20 | 13 | 0.178% | 0.16% | -0.035% | 0.1%
20 | 14 | 0.258% | 0.08% | 0.151% | 0.1%
20 | 15 | 0.241% | -0.017% | 0.226% | 0.1%
20 | 16 | 0.267% | 0.026% | 0.255% | 0.1%
Deviations of %error on test id #21. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | -25.7% | 0% | 0% | 0.1%
21 | 1 | -13.77% | 11.93% | 0% | 0.1%
21 | 2 | -5.359% | 8.414% | -14.95% | 0.1%
21 | 3 | 0.33% | 5.689% | -6.267% | 0.1%
21 | 4 | 0.25% | -0.08% | -1.593% | 0.1%
21 | 5 | 0.107% | -0.143% | 0.229% | 0.1%
21 | 6 | 0.009% | -0.098% | 0.122% | 0.1%
21 | 7 | -0.124% | -0.133% | -0.003% | 0.1%
21 | 8 | -0.204% | -0.08% | -0.106% | 0.1%
21 | 9 | -0.389% | -0.185% | -0.239% | 0.1%
21 | 10 | -0.257% | 0.132% | -0.283% | 0.1%
21 | 11 | 0.044% | 0.301% | -0.201% | 0.1%
21 | 12 | 0.25% | 0.206% | 0.012% | 0.1%
21 | 13 | 0.303% | 0.053% | 0.199% | 0.1%
21 | 14 | 0.312% | 0.009% | 0.288% | 0.1%
21 | 15 | 0.312% | 0% | 0.309% | 0.1%
Deviations of %error on test id #22. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | -26.26% | 0% | 0% | 0.1%
22 | 1 | -16.41% | 9.857% | 0% | 0.1%
22 | 2 | -4.361% | 12.04% | -15.68% | 0.1%
22 | 3 | -0.363% | 3.998% | -7.043% | 0.1%
22 | 4 | 0.107% | 0.47% | -1.539% | 0.1%
22 | 5 | 0% | -0.107% | -0.085% | 0.1%
22 | 6 | 0.036% | 0.036% | 0.048% | 0.1%
22 | 7 | -0.053% | -0.089% | -0.006% | 0.1%
22 | 8 | -0.204% | -0.151% | -0.074% | 0.1%
22 | 9 | -0.053% | 0.151% | -0.103% | 0.1%
22 | 10 | 0.142% | 0.195% | -0.038% | 0.1%
22 | 11 | 0.241% | 0.099% | 0.11% | 0.1%
22 | 12 | 0.267% | 0.026% | 0.217% | 0.1%
22 | 13 | 0.267% | 0% | 0.258% | 0.1%
Deviations of %error on test id #23. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | -22.14% | 0% | 0% | 0.1%
23 | 1 | -12.24% | 9.906% | 0% | 0.1%
23 | 2 | -2.58% | 9.659% | -12.32% | 0.1%
23 | 3 | 0.285% | 2.865% | -4.845% | 0.1%
23 | 4 | 0.107% | -0.178% | -0.729% | 0.1%
23 | 5 | 0.08% | -0.027% | 0.157% | 0.1%
23 | 6 | 0.134% | 0.054% | 0.107% | 0.1%
23 | 7 | 0.071% | -0.063% | 0.095% | 0.1%
Deviations of %error on test id #24. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | -4.991% | 0% | 0% | 0.1%
24 | 1 | -0.124% | 4.867% | 0% | 0.1%
24 | 2 | 0.187% | 0.311% | -1.643% | 0.1%
24 | 3 | 0.08% | -0.107% | 0.048% | 0.1%
24 | 4 | 0.214% | 0.134% | 0.16% | 0.1%
24 | 5 | 0.276% | 0.062% | 0.19% | 0.1%
24 | 6 | 0.018% | -0.258% | 0.169% | 0.1%
24 | 7 | 0.098% | 0.08% | 0.131% | 0.1%
24 | 8 | 0.205% | 0.107% | 0.107% | 0.1%
24 | 9 | -0.292% | -0.497% | 0.004% | 0.1%
24 | 10 | -0.68% | -0.388% | -0.256% | 0.1%
24 | 11 | -0.389% | 0.291% | -0.454% | 0.1%
24 | 12 | -0.044% | 0.345% | -0.371% | 0.1%
24 | 13 | 0.285% | 0.329% | -0.049% | 0.1%
24 | 14 | 0.375% | 0.09% | 0.205% | 0.1%
24 | 15 | 0.339% | -0.036% | 0.333% | 0.1%
24 | 16 | 0.384% | 0.045% | 0.366% | 0.1%
Deviations of %error on test id #25. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | -18.78% | 0% | 0% | 0.1%
25 | 1 | -9.566% | 9.218% | 0% | 0.1%
25 | 2 | -2.114% | 7.452% | -10.15% | 0.1%
25 | 3 | 0.142% | 2.256% | -3.846% | 0.1%
25 | 4 | -0.009% | -0.151% | -0.66% | 0.1%
25 | 5 | 0.098% | 0.107% | 0.077% | 0.1%
25 | 6 | 0.142% | 0.044% | 0.077% | 0.1%
25 | 7 | -0.133% | -0.275% | 0.036% | 0.1%
25 | 8 | -0.071% | 0.062% | -0.021% | 0.1%
25 | 9 | 0.125% | 0.196% | -0.026% | 0.1%
25 | 10 | 0.294% | 0.169% | 0.116% | 0.1%
25 | 11 | 0.321% | 0.027% | 0.247% | 0.1%
25 | 12 | 0.348% | 0.027% | 0.321% | 0.1%
25 | 13 | 0.339% | -0.009% | 0.336% | 0.1%
25 | 14 | 0.348% | 0.009% | 0.345% | 0.1%
Deviations of %error on test id #26. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | -11.27% | 0% | 0% | 0.1%
26 | 1 | -1.909% | 9.361% | 0% | 0.1%
26 | 2 | -0.027% | 1.882% | -4.402% | 0.1%
26 | 3 | 0.027% | 0.054% | -0.636% | 0.1%
26 | 4 | 0.134% | 0.107% | 0.045% | 0.1%
26 | 5 | 0.098% | -0.036% | 0.086% | 0.1%
26 | 6 | 0.036% | -0.062% | 0.089% | 0.1%
26 | 7 | 0.134% | 0.098% | 0.089% | 0.1%
Deviations of %error on test id #27. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | -15.36% | 0% | 0% | 0.1%
27 | 1 | -6.771% | 8.585% | 0% | 0.1%
27 | 2 | -1.419% | 5.352% | -7.849% | 0.1%
27 | 3 | 0.107% | 1.526% | -2.694% | 0.1%
27 | 4 | 0% | -0.107% | -0.437% | 0.1%
27 | 5 | 0.116% | 0.116% | 0.074% | 0.1%
27 | 6 | -0.036% | -0.152% | 0.027% | 0.1%
27 | 7 | -0.204% | -0.168% | -0.041% | 0.1%
27 | 8 | -0.036% | 0.168% | -0.092% | 0.1%
27 | 9 | 0.187% | 0.223% | -0.018% | 0.1%
27 | 10 | 0.294% | 0.107% | 0.148% | 0.1%
27 | 11 | 0.366% | 0.072% | 0.282% | 0.1%
27 | 12 | 0.339% | -0.027% | 0.333% | 0.1%
27 | 13 | 0.384% | 0.045% | 0.363% | 0.1%
Deviations of %error on test id #28. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | -9.718% | 0% | 0% | 0.1%
28 | 1 | -2.182% | 7.536% | 0% | 0.1%
28 | 2 | -0.124% | 2.058% | -4.008% | 0.1%
28 | 3 | -0.062% | 0.062% | -0.789% | 0.1%
28 | 4 | 0.134% | 0.196% | -0.017% | 0.1%
28 | 5 | 0.196% | 0.062% | 0.089% | 0.1%
28 | 6 | 0.053% | -0.143% | 0.128% | 0.1%
28 | 7 | 0.098% | 0.045% | 0.116% | 0.1%
28 | 8 | 0.169% | 0.071% | 0.107% | 0.1%
28 | 9 | 0.294% | 0.125% | 0.187% | 0.1%
28 | 10 | 0.303% | 0.009% | 0.255% | 0.1%
28 | 11 | 0.276% | -0.027% | 0.291% | 0.1%
28 | 12 | 0.321% | 0.045% | 0.3% | 0.1%
Deviations of %error on test id #29. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | -21.63% | 0% | 0% | 0.1%
29 | 1 | -11.86% | 9.774% | 0% | 0.1%
29 | 2 | -2.992% | 8.869% | -12.16% | 0.1%
29 | 3 | 0.178% | 3.17% | -4.892% | 0.1%
29 | 4 | 0.303% | 0.125% | -0.837% | 0.1%
29 | 5 | 0.384% | 0.081% | 0.288% | 0.1%
29 | 6 | 0.25% | -0.134% | 0.312% | 0.1%
29 | 7 | -0.044% | -0.294% | 0.197% | 0.1%
29 | 8 | -0.08% | -0.036% | 0.042% | 0.1%
29 | 9 | 0.151% | 0.231% | 0.009% | 0.1%
29 | 10 | 0.267% | 0.116% | 0.113% | 0.1%
29 | 11 | 0.303% | 0.036% | 0.24% | 0.1%
29 | 12 | 0.366% | 0.063% | 0.312% | 0.1%
29 | 13 | 0.321% | -0.045% | 0.33% | 0.1%
Deviations of %error on test id #30. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | -15.34% | 0% | 0% | 0.1%
30 | 1 | -5.51% | 9.833% | 0% | 0.1%
30 | 2 | -1.471% | 4.039% | -7.441% | 0.1%
30 | 3 | 0.107% | 1.578% | -2.291% | 0.1%
30 | 4 | 0.241% | 0.134% | -0.374% | 0.1%
30 | 5 | 0.178% | -0.063% | 0.175% | 0.1%
30 | 6 | 0.018% | -0.16% | 0.146% | 0.1%
30 | 7 | -0.098% | -0.116% | 0.033% | 0.1%
30 | 8 | 0.098% | 0.196% | 0.006% | 0.1%
30 | 9 | 0.241% | 0.143% | 0.08% | 0.1%
30 | 10 | 0.393% | 0.152% | 0.244% | 0.1%
30 | 11 | 0.366% | -0.027% | 0.333% | 0.1%
30 | 12 | 0.42% | 0.054% | 0.393% | 0.1%
30 | 13 | 0.42% | 0% | 0.402% | 0.1%
Deviations of %error on test id #31. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | -4.166% | 0% | 0% | 0.1%
31 | 1 | 0.196% | 4.362% | 0% | 0.1%
31 | 2 | -0.133% | -0.329% | -1.368% | 0.1%
31 | 3 | -0.248% | -0.115% | -0.062% | 0.1%
31 | 4 | -0.257% | -0.009% | -0.213% | 0.1%
31 | 5 | 0% | 0.257% | -0.168% | 0.1%
31 | 6 | 0.205% | 0.205% | -0.017% | 0.1%
31 | 7 | 0.267% | 0.062% | 0.157% | 0.1%
31 | 8 | 0.312% | 0.045% | 0.261% | 0.1%
31 | 9 | 0.321% | 0.009% | 0.3% | 0.1%
Deviations of %error on test id #32. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | -9.928% | 0% | 0% | 0.1%
32 | 1 | -3.632% | 6.296% | 0% | 0.1%
32 | 2 | 0.33% | 3.962% | -4.41% | 0.1%
32 | 3 | 0.125% | -0.205% | -1.059% | 0.1%
32 | 4 | -0.018% | -0.143% | 0.146% | 0.1%
32 | 5 | 0.062% | 0.08% | 0.056% | 0.1%
32 | 6 | 0.044% | -0.018% | 0.029% | 0.1%
32 | 7 | 0.027% | -0.017% | 0.044% | 0.1%
Deviations of %error on test id #33. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | -20.81% | 0% | 0% | 0.1%
33 | 1 | -10.93% | 9.887% | 0% | 0.1%
33 | 2 | -3.092% | 7.834% | -11.61% | 0.1%
33 | 3 | -0.009% | 3.083% | -4.676% | 0.1%
33 | 4 | 0.125% | 0.134% | -0.992% | 0.1%
33 | 5 | 0.134% | 0.009% | 0.083% | 0.1%
33 | 6 | 0% | -0.134% | 0.086% | 0.1%
33 | 7 | -0.08% | -0.08% | 0.018% | 0.1%
33 | 8 | 0.009% | 0.089% | -0.024% | 0.1%
33 | 9 | 0.16% | 0.151% | 0.03% | 0.1%
33 | 10 | 0.33% | 0.17% | 0.166% | 0.1%
33 | 11 | 0.357% | 0.027% | 0.282% | 0.1%
33 | 12 | 0.321% | -0.036% | 0.336% | 0.1%
33 | 13 | 0.321% | 0% | 0.333% | 0.1%
Deviations of %error on test id #34. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | -14.42% | 0% | 0% | 0.1%
34 | 1 | -3.292% | 11.12% | 0% | 0.1%
34 | 2 | -0.478% | 2.814% | -6.062% | 0.1%
34 | 3 | 0.089% | 0.567% | -1.227% | 0.1%
34 | 4 | 0.044% | -0.045% | -0.115% | 0.1%
34 | 5 | -0.071% | -0.115% | 0.021% | 0.1%
34 | 6 | -0.151% | -0.08% | -0.059% | 0.1%
34 | 7 | -0.142% | 0.009% | -0.121% | 0.1%
34 | 8 | 0.089% | 0.231% | -0.068% | 0.1%
34 | 9 | 0.241% | 0.152% | 0.063% | 0.1%
34 | 10 | 0.312% | 0.071% | 0.214% | 0.1%
34 | 11 | 0.348% | 0.036% | 0.3% | 0.1%
34 | 12 | 0.357% | 0.009% | 0.339% | 0.1%
Deviations of %error on test id #35. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | -15.25% | 0% | 0% | 0.1%
35 | 1 | -5.771% | 9.476% | 0% | 0.1%
35 | 2 | 0.036% | 5.807% | -6.994% | 0.1%
35 | 3 | 0.125% | 0.089% | -1.87% | 0.1%
35 | 4 | 0.151% | 0.026% | 0.104% | 0.1%
35 | 5 | -0.053% | -0.204% | 0.074% | 0.1%
35 | 6 | -0.151% | -0.098% | -0.018% | 0.1%
35 | 7 | 0.044% | 0.195% | -0.053% | 0.1%
35 | 8 | 0.196% | 0.152% | 0.03% | 0.1%
35 | 9 | 0.294% | 0.098% | 0.178% | 0.1%
35 | 10 | 0.357% | 0.063% | 0.282% | 0.1%
35 | 11 | 0.285% | -0.072% | 0.312% | 0.1%
35 | 12 | 0.366% | 0.081% | 0.336% | 0.1%
Deviations of %error on test id #36. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | -13.96% | 0% | 0% | 0.1%
36 | 1 | -6.156% | 7.808% | 0% | 0.1%
36 | 2 | -1.394% | 4.762% | -7.171% | 0.1%
36 | 3 | 0.25% | 1.644% | -2.433% | 0.1%
36 | 4 | 0.393% | 0.143% | -0.25% | 0.1%
36 | 5 | 0.214% | -0.179% | 0.286% | 0.1%
36 | 6 | 0.018% | -0.196% | 0.208% | 0.1%
36 | 7 | -0.089% | -0.107% | 0.048% | 0.1%
36 | 8 | 0.08% | 0.169% | 0.003% | 0.1%
36 | 9 | 0.303% | 0.223% | 0.098% | 0.1%
36 | 10 | 0.357% | 0.054% | 0.247% | 0.1%
36 | 11 | 0.375% | 0.018% | 0.345% | 0.1%
36 | 12 | 0.366% | -0.009% | 0.366% | 0.1%
Deviations of %error on test id #37. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | -12.21% | 0% | 0% | 0.1%
37 | 1 | -3.375% | 8.837% | 0% | 0.1%
37 | 2 | 0.214% | 3.589% | -5.124% | 0.1%
37 | 3 | 0.321% | 0.107% | -0.947% | 0.1%
37 | 4 | 0.33% | 0.009% | 0.288% | 0.1%
37 | 5 | 0.044% | -0.286% | 0.232% | 0.1%
37 | 6 | -0.027% | -0.071% | 0.116% | 0.1%
37 | 7 | 0.009% | 0.036% | 0.009% | 0.1%
37 | 8 | 0.223% | 0.214% | 0.068% | 0.1%
37 | 9 | 0.366% | 0.143% | 0.199% | 0.1%
37 | 10 | 0.411% | 0.045% | 0.333% | 0.1%
37 | 11 | 0.455% | 0.044% | 0.411% | 0.1%
37 | 12 | 0.411% | -0.044% | 0.426% | 0.1%
Deviations of %error on test id #38. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | -13.13% | 0% | 0% | 0.1%
38 | 1 | -4.45% | 8.677% | 0% | 0.1%
38 | 2 | 0.232% | 4.682% | -5.782% | 0.1%
38 | 3 | 0.339% | 0.107% | -1.293% | 0.1%
38 | 4 | 0.321% | -0.018% | 0.297% | 0.1%
38 | 5 | 0.125% | -0.196% | 0.262% | 0.1%
38 | 6 | -0.089% | -0.214% | 0.119% | 0.1%
38 | 7 | 0.089% | 0.178% | 0.042% | 0.1%
38 | 8 | 0.276% | 0.187% | 0.092% | 0.1%
38 | 9 | 0.357% | 0.081% | 0.241% | 0.1%
38 | 10 | 0.411% | 0.054% | 0.348% | 0.1%
38 | 11 | 0.357% | -0.054% | 0.375% | 0.1%
Deviations of %error on test id #39. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | -9.106% | 0% | 0% | 0.1%
39 | 1 | -1.901% | 7.205% | 0% | 0.1%
39 | 2 | 0.33% | 2.231% | -3.559% | 0.1%
39 | 3 | 0.42% | 0.09% | -0.384% | 0.1%
39 | 4 | 0.232% | -0.188% | 0.327% | 0.1%
39 | 5 | -0.027% | -0.259% | 0.208% | 0.1%
39 | 6 | -0.036% | -0.009% | 0.056% | 0.1%
39 | 7 | 0.232% | 0.268% | 0.056% | 0.1%
39 | 8 | 0.375% | 0.143% | 0.19% | 0.1%
39 | 9 | 0.402% | 0.027% | 0.336% | 0.1%
39 | 10 | 0.455% | 0.053% | 0.411% | 0.1%
39 | 11 | 0.446% | -0.009% | 0.434% | 0.1%
Deviations of %error on test id #40. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | -8.058% | 0% | 0% | 0.1%
40 | 1 | 0.098% | 8.156% | 0% | 0.1%
40 | 2 | 0.178% | 0.08% | -2.594% | 0.1%
40 | 3 | 0.339% | 0.161% | 0.205% | 0.1%
40 | 4 | 0.205% | -0.134% | 0.241% | 0.1%
40 | 5 | -0.009% | -0.214% | 0.178% | 0.1%
40 | 6 | -0.018% | -0.009% | 0.059% | 0.1%
40 | 7 | 0.241% | 0.259% | 0.071% | 0.1%
40 | 8 | 0.384% | 0.143% | 0.202% | 0.1%
40 | 9 | 0.393% | 0.009% | 0.339% | 0.1%
40 | 10 | 0.446% | 0.053% | 0.408% | 0.1%
40 | 11 | 0.42% | -0.026% | 0.42% | 0.1%

## 25. PF = 1.0
[<<]
Accuracy Test Report of 'PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:24:07 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 20 | 20 |  | ±0.1% | ±0.5% | 0.024% | PASSED | 2023-07-25 20:21
Yes | 1 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.5% | 0.007% | PASSED | 2023-07-25 20:21
Yes | 3 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.5% | 0.039% | PASSED | 2023-07-25 20:22
Yes | 4 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0.038% | PASSED | 2023-07-25 20:22
Yes | 7 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0.056% | PASSED | 2023-07-25 20:23
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 0.027% | 0% | 0% | 0.1%
0 | 1 | 0.027% | 0% | 0% | 0.1%
0 | 2 | 0.018% | -0.009% | 0.024% | 0.1%
0 | 3 | 0.027% | 0.009% | 0.024% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -0.064% | 0% | 0% | 0.1%
1 | 1 | 0% | 0.064% | 0% | 0.1%
1 | 2 | 0.011% | 0.011% | -0.018% | 0.1%
1 | 3 | 0.011% | 0% | 0.007% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.058% | 0% | 0% | 0.1%
3 | 1 | 0.027% | -0.031% | 0% | 0.1%
3 | 2 | 0.018% | -0.009% | 0.034% | 0.1%
3 | 3 | 0.071% | 0.053% | 0.039% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.018% | 0% | 0% | 0.1%
4 | 1 | 0.044% | 0.026% | 0% | 0.1%
4 | 2 | 0.053% | 0.009% | 0.038% | 0.1%
4 | 3 | 0.018% | -0.035% | 0.038% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.053% | 0% | 0% | 0.1%
7 | 1 | 0.055% | 0.002% | 0% | 0.1%
7 | 2 | 0.064% | 0.009% | 0.057% | 0.1%
7 | 3 | 0.05% | -0.014% | 0.056% | 0.1%

## 26. PF = 0.5L
[<<]
Accuracy Test Report of 'PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:24:11 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
No | 0 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | 0.232% | PASSED | 2023-07-25 20:23
Yes | 1 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | 0.248% | PASSED | 2023-07-25 20:31
No | 3 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | -0.127% | PASSED | 2023-07-25 20:24
No | 4 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.142% | PASSED | 2023-07-25 20:25
No | 7 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.119% | PASSED | 2023-07-25 20:26
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -15.82% | 0% | 0% | 0.1%
0 | 1 | -32.26% | -16.44% | 0% | 0.1%
0 | 2 | -33.92% | -1.663% | -27.33% | 0.1%
0 | 3 | -20.45% | 13.48% | -28.88% | 0.1%
0 | 4 | -6.109% | 14.34% | -20.16% | 0.1%
0 | 5 | 0.267% | 6.376% | -8.764% | 0.1%
0 | 6 | 0.232% | -0.035% | -1.87% | 0.1%
0 | 7 | 0.232% | 0% | 0.244% | 0.1%
0 | 8 | 0.232% | 0% | 0.232% | 0.1%
Deviations of %error on test id #1. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -22.31% | 0% | 0% | 0.1%
1 | 1 | -5.345% | 16.97% | 0% | 0.1%
1 | 2 | 0.975% | 6.32% | -8.894% | 0.1%
1 | 3 | 0.065% | -0.91% | -1.435% | 0.1%
1 | 4 | -0.077% | -0.142% | 0.321% | 0.1%
1 | 5 | -0.177% | -0.1% | -0.063% | 0.1%
1 | 6 | -0.036% | 0.141% | -0.097% | 0.1%
1 | 7 | 0.125% | 0.161% | -0.029% | 0.1%
1 | 8 | -0.112% | -0.237% | -0.008% | 0.1%
1 | 9 | -0.16% | -0.048% | -0.049% | 0.1%
1 | 10 | 0.131% | 0.291% | -0.047% | 0.1%
1 | 11 | 0.232% | 0.101% | 0.068% | 0.1%
1 | 12 | 0.25% | 0.018% | 0.204% | 0.1%
1 | 13 | 0.244% | -0.006% | 0.242% | 0.1%
1 | 14 | 0.25% | 0.006% | 0.248% | 0.1%
Deviations of %error on test id #3. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -34.25% | 0% | 0% | 0.1%
3 | 1 | -25.75% | 8.505% | 0% | 0.1%
3 | 2 | -11.93% | 13.82% | -23.98% | 0.1%
3 | 3 | -3.599% | 8.331% | -13.76% | 0.1%
3 | 4 | 0.214% | 3.813% | -5.105% | 0.1%
3 | 5 | -0.036% | -0.25% | -1.14% | 0.1%
3 | 6 | -0.151% | -0.115% | 0.009% | 0.1%
3 | 7 | -0.142% | 0.009% | -0.11% | 0.1%
3 | 8 | -0.089% | 0.053% | -0.127% | 0.1%
3 | 9 | -0.151% | -0.062% | -0.127% | 0.1%
Deviations of %error on test id #4. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -31.18% | 0% | 0% | 0.1%
4 | 1 | -8.744% | 22.44% | 0% | 0.1%
4 | 2 | 0.142% | 8.886% | -13.26% | 0.1%
4 | 3 | 0.018% | -0.124% | -2.861% | 0.1%
4 | 4 | -0.23% | -0.248% | -0.023% | 0.1%
4 | 5 | -0.381% | -0.151% | -0.198% | 0.1%
4 | 6 | -0.398% | -0.017% | -0.336% | 0.1%
4 | 7 | 0.027% | 0.425% | -0.251% | 0.1%
4 | 8 | 0.16% | 0.133% | -0.07% | 0.1%
4 | 9 | 0.16% | 0% | 0.116% | 0.1%
4 | 10 | 0.384% | 0.224% | 0.235% | 0.1%
4 | 11 | 0.411% | 0.027% | 0.318% | 0.1%
4 | 12 | 0.214% | -0.197% | 0.336% | 0.1%
4 | 13 | 0.134% | -0.08% | 0.253% | 0.1%
4 | 14 | 0.142% | 0.008% | 0.163% | 0.1%
4 | 15 | 0.151% | 0.009% | 0.142% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.062% | 0% | 0% | 0.1%
7 | 1 | 0.119% | 0.057% | 0% | 0.1%
7 | 2 | 0.114% | -0.005% | 0.098% | 0.1%
7 | 3 | 0.125% | 0.011% | 0.119% | 0.1%

## 27. PF = 0.8C
[<<]
Accuracy Test Report of 'PF = 0.8C'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / PF = 0.8C'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:24:17 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | 0.136% | PASSED | 2023-07-25 20:26
Yes | 1 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | -0.075% | PASSED | 2023-07-25 20:27
Yes | 3 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.065% | PASSED | 2023-07-25 20:27
Yes | 4 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.178% | PASSED | 2023-07-25 20:28
Yes | 7 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.028% | PASSED | 2023-07-25 20:29
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -42.29% | 0% | 0% | 0.1%
0 | 1 | 260.8% | 303.1% | 0% | 0.1%
0 | 2 | 553.3% | 292.5% | 257.3% | 0.1%
0 | 3 | 62.81% | -490.5% | 292.3% | 0.1%
0 | 4 | 41.23% | -21.58% | 219.1% | 0.1%
0 | 5 | -0.071% | -41.3% | 34.65% | 0.1%
0 | 6 | 0.036% | 0.107% | 13.73% | 0.1%
0 | 7 | 0.089% | 0.053% | 0.018% | 0.1%
0 | 8 | 0.16% | 0.071% | 0.095% | 0.1%
0 | 9 | 0.16% | 0% | 0.136% | 0.1%
Deviations of %error on test id #1. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 150.2% | 0% | 0% | 0.1%
1 | 1 | 187.1% | 36.91% | 0% | 0.1%
1 | 2 | 45.81% | -141.3% | 127.7% | 0.1%
1 | 3 | 8.111% | -37.7% | 80.35% | 0.1%
1 | 4 | -0.124% | -8.235% | 17.93% | 0.1%
1 | 5 | -0.065% | 0.059% | 2.641% | 0.1%
1 | 6 | -0.089% | -0.024% | -0.093% | 0.1%
1 | 7 | -0.071% | 0.018% | -0.075% | 0.1%
Deviations of %error on test id #3. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 102.2% | 0% | 0% | 0.1%
3 | 1 | 396% | 293.9% | 0% | 0.1%
3 | 2 | 396.5% | 0.438% | 298.2% | 0.1%
3 | 3 | 34.94% | -361.5% | 275.8% | 0.1%
3 | 4 | 30.1% | -4.838% | 153.8% | 0.1%
3 | 5 | 7.841% | -22.26% | 24.3% | 0.1%
3 | 6 | 0.009% | -7.832% | 12.65% | 0.1%
3 | 7 | 0.009% | 0% | 2.62% | 0.1%
3 | 8 | 0.089% | 0.08% | 0.036% | 0.1%
3 | 9 | 0.098% | 0.009% | 0.065% | 0.1%
Deviations of %error on test id #4. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 235.6% | 0% | 0% | 0.1%
4 | 1 | 112.5% | -123.1% | 0% | 0.1%
4 | 2 | 8.131% | -104.4% | 118.8% | 0.1%
4 | 3 | 0.996% | -7.135% | 40.56% | 0.1%
4 | 4 | 0.071% | -0.925% | 3.066% | 0.1%
4 | 5 | 0.116% | 0.045% | 0.394% | 0.1%
4 | 6 | 0.196% | 0.08% | 0.128% | 0.1%
4 | 7 | 0.223% | 0.027% | 0.178% | 0.1%
Deviations of %error on test id #7. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.134% | 0% | 0% | 0.1%
7 | 1 | -0.055% | -0.189% | 0% | 0.1%
7 | 2 | 0.023% | 0.078% | 0.034% | 0.1%
7 | 3 | 0.034% | 0.011% | 0.001% | 0.1%
7 | 4 | 0.027% | -0.007% | 0.028% | 0.1%

## 28. 5th Harmonics
[<<]
Accuracy Test Report of '5th Harmonics'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / 5th Harmonics'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:24:21 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Harmonics' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
1 | 0 | 0 | 1
3. Test Cases
Total 1 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 1 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° | {UcUbUa, 5th-10%, IcIbIa, 5th-40%} | 12 | 20 |  | ±0.1% | ±1% | 0.009% | PASSED | 2023-07-25 20:29
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #1. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -0.177% | 0% | 0% | 0.1%
1 | 1 | -2.564% | -2.387% | 0% | 0.1%
1 | 2 | -5.663% | -3.099% | -2.801% | 0.1%
1 | 3 | -2.825% | 2.838% | -3.684% | 0.1%
1 | 4 | -0.442% | 2.383% | -2.977% | 0.1%
1 | 5 | 0.015% | 0.457% | -1.084% | 0.1%
1 | 6 | 0.524% | 0.509% | 0.032% | 0.1%
1 | 7 | 0.345% | -0.179% | 0.295% | 0.1%
1 | 8 | 0.122% | -0.223% | 0.33% | 0.1%
1 | 9 | 0.024% | -0.098% | 0.164% | 0.1%
1 | 10 | -0.021% | -0.045% | 0.042% | 0.1%
1 | 11 | 0.024% | 0.045% | 0.009% | 0.1%

## 29. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:24:23 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RQ | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | 0.037% | PASSED | 2023-07-25 20:32
Yes | 1 | RQ | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.059% | PASSED | 2023-07-25 20:33
Yes | 2 | RQ | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.055% | PASSED | 2023-07-25 20:33
Yes | 3 | RQ | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.046% | PASSED | 2023-07-25 20:33
Yes | 4 | RQ | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.054% | PASSED | 2023-07-25 20:33
Yes | 5 | RQ | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | 0.037% | PASSED | 2023-07-25 20:34
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 0.022% | 0% | 0% | 0.1%
0 | 1 | 0.031% | 0.009% | 0% | 0.1%
0 | 2 | 0.027% | -0.004% | 0.027% | 0.1%
0 | 3 | 0.053% | 0.026% | 0.037% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0.049% | 0% | 0% | 0.1%
1 | 1 | 0.058% | 0.009% | 0% | 0.1%
1 | 2 | 0.062% | 0.004% | 0.056% | 0.1%
1 | 3 | 0.058% | -0.004% | 0.059% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | 0.031% | 0% | 0% | 0.1%
2 | 1 | 0.049% | 0.018% | 0% | 0.1%
2 | 2 | 0.036% | -0.013% | 0.039% | 0.1%
2 | 3 | 0.08% | 0.044% | 0.055% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.027% | 0% | 0% | 0.1%
3 | 1 | 0.036% | 0.009% | 0% | 0.1%
3 | 2 | 0.036% | 0% | 0.033% | 0.1%
3 | 3 | 0.067% | 0.031% | 0.046% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.027% | 0% | 0% | 0.1%
4 | 1 | 0.027% | 0% | 0% | 0.1%
4 | 2 | 0.076% | 0.049% | 0.043% | 0.1%
4 | 3 | 0.058% | -0.018% | 0.054% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | 0.062% | 0% | 0% | 0.1%
5 | 1 | 0.004% | -0.058% | 0% | 0.1%
5 | 2 | 0.053% | 0.049% | 0.04% | 0.1%
5 | 3 | 0.053% | 0% | 0.037% | 0.1%

## 30. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:24:26 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RQ | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | 0.136% | PASSED | 2023-07-25 20:34
Yes | 1 | RQ | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.146% | PASSED | 2023-07-25 20:35
Yes | 2 | RQ | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | -0.189% | PASSED | 2023-07-25 20:35
Yes | 3 | RQ | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.071% | PASSED | 2023-07-25 20:36
Yes | 4 | RQ | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | -0.139% | PASSED | 2023-07-25 20:36
Yes | 5 | RQ | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | -0.195% | PASSED | 2023-07-25 20:36
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -15.14% | 0% | 0% | 0.1%
0 | 1 | -2.639% | 12.5% | 0% | 0.1%
0 | 2 | -0.239% | 2.4% | -6.006% | 0.1%
0 | 3 | -0.089% | 0.15% | -0.989% | 0.1%
0 | 4 | 0.018% | 0.107% | -0.103% | 0.1%
0 | 5 | -0.151% | -0.169% | -0.074% | 0.1%
0 | 6 | -0.248% | -0.097% | -0.127% | 0.1%
0 | 7 | -0.195% | 0.053% | -0.198% | 0.1%
0 | 8 | 0% | 0.195% | -0.148% | 0.1%
0 | 9 | 0.107% | 0.107% | -0.029% | 0.1%
0 | 10 | 0.151% | 0.044% | 0.086% | 0.1%
0 | 11 | 0.142% | -0.009% | 0.133% | 0.1%
0 | 12 | 0.116% | -0.026% | 0.136% | 0.1%
Deviations of %error on test id #1. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -31.64% | 0% | 0% | 0.1%
1 | 1 | -21.64% | 10% | 0% | 0.1%
1 | 2 | -10.17% | 11.48% | -21.15% | 0.1%
1 | 3 | -0.151% | 10.01% | -10.65% | 0.1%
1 | 4 | 0.16% | 0.311% | -3.385% | 0.1%
1 | 5 | 0.196% | 0.036% | 0.068% | 0.1%
1 | 6 | -0.133% | -0.329% | 0.074% | 0.1%
1 | 7 | -0.089% | 0.044% | -0.009% | 0.1%
1 | 8 | -0.027% | 0.062% | -0.083% | 0.1%
1 | 9 | -0.301% | -0.274% | -0.139% | 0.1%
1 | 10 | -0.706% | -0.405% | -0.345% | 0.1%
1 | 11 | -0.522% | 0.184% | -0.51% | 0.1%
1 | 12 | -0.177% | 0.345% | -0.468% | 0.1%
1 | 13 | 0.062% | 0.239% | -0.212% | 0.1%
1 | 14 | 0.134% | 0.072% | 0.006% | 0.1%
1 | 15 | 0.169% | 0.035% | 0.122% | 0.1%
1 | 16 | 0.134% | -0.035% | 0.146% | 0.1%
Deviations of %error on test id #2. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -25.14% | 0% | 0% | 0.1%
2 | 1 | -11.08% | 14.06% | 0% | 0.1%
2 | 2 | -4.231% | 6.85% | -13.48% | 0.1%
2 | 3 | 0.196% | 4.427% | -5.039% | 0.1%
2 | 4 | 0.018% | -0.178% | -1.339% | 0.1%
2 | 5 | -0.239% | -0.257% | -0.008% | 0.1%
2 | 6 | -0.151% | 0.088% | -0.124% | 0.1%
2 | 7 | -0.248% | -0.097% | -0.213% | 0.1%
2 | 8 | -0.169% | 0.079% | -0.189% | 0.1%
Deviations of %error on test id #3. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -31.61% | 0% | 0% | 0.1%
3 | 1 | -20.75% | 10.86% | 0% | 0.1%
3 | 2 | -6.825% | 13.93% | -19.73% | 0.1%
3 | 3 | -2.766% | 4.059% | -10.11% | 0.1%
3 | 4 | 0.16% | 2.926% | -3.144% | 0.1%
3 | 5 | 0.125% | -0.035% | -0.827% | 0.1%
3 | 6 | 0.009% | -0.116% | 0.098% | 0.1%
3 | 7 | -0.292% | -0.301% | -0.053% | 0.1%
3 | 8 | -0.416% | -0.124% | -0.233% | 0.1%
3 | 9 | -0.451% | -0.035% | -0.386% | 0.1%
3 | 10 | -0.266% | 0.185% | -0.378% | 0.1%
3 | 11 | -0.062% | 0.204% | -0.26% | 0.1%
3 | 12 | 0.027% | 0.089% | -0.1% | 0.1%
3 | 13 | 0.062% | 0.035% | 0.009% | 0.1%
3 | 14 | 0.125% | 0.063% | 0.071% | 0.1%
Deviations of %error on test id #4. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -26.14% | 0% | 0% | 0.1%
4 | 1 | -15.85% | 10.29% | 0% | 0.1%
4 | 2 | -8.811% | 7.039% | -16.93% | 0.1%
4 | 3 | 0.151% | 8.962% | -8.17% | 0.1%
4 | 4 | 0.16% | 0.009% | -2.833% | 0.1%
4 | 5 | -0.027% | -0.187% | 0.095% | 0.1%
4 | 6 | -0.106% | -0.079% | 0.009% | 0.1%
4 | 7 | -0.08% | 0.026% | -0.071% | 0.1%
4 | 8 | -0.142% | -0.062% | -0.109% | 0.1%
4 | 9 | -0.195% | -0.053% | -0.139% | 0.1%
Deviations of %error on test id #5. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -33.04% | 0% | 0% | 0.1%
5 | 1 | -27.44% | 5.601% | 0% | 0.1%
5 | 2 | -11.68% | 15.76% | -24.05% | 0.1%
5 | 3 | -7.726% | 3.955% | -15.62% | 0.1%
5 | 4 | 0.16% | 7.886% | -6.416% | 0.1%
5 | 5 | 0.178% | 0.018% | -2.463% | 0.1%
5 | 6 | 0.018% | -0.16% | 0.119% | 0.1%
5 | 7 | -0.151% | -0.169% | 0.015% | 0.1%
5 | 8 | -0.151% | 0% | -0.095% | 0.1%
5 | 9 | -0.195% | -0.044% | -0.166% | 0.1%
5 | 10 | -0.239% | -0.044% | -0.195% | 0.1%

## 31. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:24:33 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.027% | PASSED | 2023-07-25 20:38
Yes | 1 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.01% | PASSED | 2023-07-25 20:38
Yes | 2 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.024% | PASSED | 2023-07-25 20:38
Yes | 3 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.007% | PASSED | 2023-07-25 20:39
Yes | 4 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.012% | PASSED | 2023-07-25 20:39
Yes | 5 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.012% | PASSED | 2023-07-25 20:39
Yes | 6 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.037% | PASSED | 2023-07-25 20:40
Yes | 7 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.036% | PASSED | 2023-07-25 20:40
Yes | 8 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.036% | PASSED | 2023-07-25 20:40
Yes | 9 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.015% | PASSED | 2023-07-25 20:40
Yes | 10 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.047% | PASSED | 2023-07-25 20:41
Yes | 11 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.058% | PASSED | 2023-07-25 20:41
Yes | 12 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.052% | PASSED | 2023-07-25 20:41
Yes | 13 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.045% | PASSED | 2023-07-25 20:42
Yes | 14 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.036% | PASSED | 2023-07-25 20:42
Yes | 15 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.04% | PASSED | 2023-07-25 20:42
Yes | 16 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.061% | PASSED | 2023-07-25 20:43
Yes | 17 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.089% | PASSED | 2023-07-25 20:43
Yes | 18 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.11% | PASSED | 2023-07-25 20:43
Yes | 19 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.074% | PASSED | 2023-07-25 20:43
Yes | 20 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.095% | PASSED | 2023-07-25 20:44
Yes | 21 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.073% | PASSED | 2023-07-25 20:44
Yes | 22 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.088% | PASSED | 2023-07-25 20:44
Yes | 23 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.111% | PASSED | 2023-07-25 20:45
Yes | 24 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.112% | PASSED | 2023-07-25 20:45
Yes | 25 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.095% | PASSED | 2023-07-25 20:45
Yes | 26 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.095% | PASSED | 2023-07-25 20:46
Yes | 27 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.086% | PASSED | 2023-07-25 20:46
Yes | 28 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.09% | PASSED | 2023-07-25 20:46
Yes | 29 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.102% | PASSED | 2023-07-25 20:47
Yes | 30 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.141% | PASSED | 2023-07-25 20:47
Yes | 31 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.137% | PASSED | 2023-07-25 20:47
Yes | 32 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.104% | PASSED | 2023-07-25 20:48
Yes | 33 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.099% | PASSED | 2023-07-25 20:48
Yes | 34 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.108% | PASSED | 2023-07-25 20:48
Yes | 35 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.102% | PASSED | 2023-07-25 20:48
Yes | 36 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.116% | PASSED | 2023-07-25 20:49
Yes | 37 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.147% | PASSED | 2023-07-25 20:49
Yes | 38 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.125% | PASSED | 2023-07-25 20:49
Yes | 39 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.151% | PASSED | 2023-07-25 20:50
Yes | 40 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.12% | PASSED | 2023-07-25 20:50
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -0.027% | 0% | 0% | 0.1%
0 | 1 | -0.027% | 0% | 0% | 0.1%
0 | 2 | -0.049% | -0.022% | -0.034% | 0.1%
0 | 3 | -0.004% | 0.045% | -0.027% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -0.044% | 0% | 0% | 0.1%
1 | 1 | -0.004% | 0.04% | 0% | 0.1%
1 | 2 | -0.022% | -0.018% | -0.023% | 0.1%
1 | 3 | -0.004% | 0.018% | -0.01% | 0.1%
Deviations of %error on test id #2. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -3.977% | 0% | 0% | 0.1%
2 | 1 | -1.841% | 2.136% | 0% | 0.1%
2 | 2 | 0.022% | 1.863% | -1.932% | 0.1%
2 | 3 | -0.013% | -0.035% | -0.611% | 0.1%
2 | 4 | 0.071% | 0.084% | 0.027% | 0.1%
2 | 5 | 0.013% | -0.058% | 0.024% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0% | 0% | 0% | 0.1%
3 | 1 | 0.009% | 0.009% | 0% | 0.1%
3 | 2 | 0.013% | 0.004% | 0.007% | 0.1%
3 | 3 | 0% | -0.013% | 0.007% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -0.031% | 0% | 0% | 0.1%
4 | 1 | 0.013% | 0.044% | 0% | 0.1%
4 | 2 | 0.004% | -0.009% | -0.005% | 0.1%
4 | 3 | 0.018% | 0.014% | 0.012% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | 0.004% | 0% | 0% | 0.1%
5 | 1 | 0.004% | 0% | 0% | 0.1%
5 | 2 | 0.036% | 0.032% | 0.015% | 0.1%
5 | 3 | -0.004% | -0.04% | 0.012% | 0.1%
Deviations of %error on test id #6. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | 0.04% | 0% | 0% | 0.1%
6 | 1 | 0.04% | 0% | 0% | 0.1%
6 | 2 | 0.036% | -0.004% | 0.039% | 0.1%
6 | 3 | 0.036% | 0% | 0.037% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.031% | 0% | 0% | 0.1%
7 | 1 | 0.04% | 0.009% | 0% | 0.1%
7 | 2 | 0.036% | -0.004% | 0.036% | 0.1%
7 | 3 | 0.031% | -0.005% | 0.036% | 0.1%
Deviations of %error on test id #8. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | 0.067% | 0% | 0% | 0.1%
8 | 1 | 0.036% | -0.031% | 0% | 0.1%
8 | 2 | 0.031% | -0.005% | 0.045% | 0.1%
8 | 3 | 0.04% | 0.009% | 0.036% | 0.1%
Deviations of %error on test id #9. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | 0.044% | 0% | 0% | 0.1%
9 | 1 | 0.04% | -0.004% | 0% | 0.1%
9 | 2 | 0.013% | -0.027% | 0.032% | 0.1%
9 | 3 | -0.009% | -0.022% | 0.015% | 0.1%
Deviations of %error on test id #10. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | 0.036% | 0% | 0% | 0.1%
10 | 1 | 0.076% | 0.04% | 0% | 0.1%
10 | 2 | 0.022% | -0.054% | 0.045% | 0.1%
10 | 3 | 0.044% | 0.022% | 0.047% | 0.1%
Deviations of %error on test id #11. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | 0.076% | 0% | 0% | 0.1%
11 | 1 | 0.031% | -0.045% | 0% | 0.1%
11 | 2 | 0.071% | 0.04% | 0.059% | 0.1%
11 | 3 | 0.071% | 0% | 0.058% | 0.1%
Deviations of %error on test id #12. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | 0.085% | 0% | 0% | 0.1%
12 | 1 | 0.076% | -0.009% | 0% | 0.1%
12 | 2 | 0.031% | -0.045% | 0.064% | 0.1%
12 | 3 | 0.049% | 0.018% | 0.052% | 0.1%
Deviations of %error on test id #13. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | 0.036% | 0% | 0% | 0.1%
13 | 1 | 0.036% | 0% | 0% | 0.1%
13 | 2 | 0.062% | 0.026% | 0.045% | 0.1%
13 | 3 | 0.036% | -0.026% | 0.045% | 0.1%
Deviations of %error on test id #14. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | 0.031% | 0% | 0% | 0.1%
14 | 1 | 0.04% | 0.009% | 0% | 0.1%
14 | 2 | 0.036% | -0.004% | 0.036% | 0.1%
14 | 3 | 0.031% | -0.005% | 0.036% | 0.1%
Deviations of %error on test id #15. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | 0.036% | 0% | 0% | 0.1%
15 | 1 | 0.031% | -0.005% | 0% | 0.1%
15 | 2 | 0.049% | 0.018% | 0.039% | 0.1%
15 | 3 | 0.04% | -0.009% | 0.04% | 0.1%
Deviations of %error on test id #16. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | 0.058% | 0% | 0% | 0.1%
16 | 1 | 0.036% | -0.022% | 0% | 0.1%
16 | 2 | 0.08% | 0.044% | 0.058% | 0.1%
16 | 3 | 0.067% | -0.013% | 0.061% | 0.1%
Deviations of %error on test id #17. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | 0.08% | 0% | 0% | 0.1%
17 | 1 | 0.12% | 0.04% | 0% | 0.1%
17 | 2 | 0.071% | -0.049% | 0.09% | 0.1%
17 | 3 | 0.076% | 0.005% | 0.089% | 0.1%
Deviations of %error on test id #18. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | 0.089% | 0% | 0% | 0.1%
18 | 1 | 0.076% | -0.013% | 0% | 0.1%
18 | 2 | 0.125% | 0.049% | 0.097% | 0.1%
18 | 3 | 0.129% | 0.004% | 0.11% | 0.1%
Deviations of %error on test id #19. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | 0.093% | 0% | 0% | 0.1%
19 | 1 | 0.08% | -0.013% | 0% | 0.1%
19 | 2 | 0.085% | 0.005% | 0.086% | 0.1%
19 | 3 | 0.058% | -0.027% | 0.074% | 0.1%
Deviations of %error on test id #20. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | 0.12% | 0% | 0% | 0.1%
20 | 1 | 0.085% | -0.035% | 0% | 0.1%
20 | 2 | 0.125% | 0.04% | 0.11% | 0.1%
20 | 3 | 0.076% | -0.049% | 0.095% | 0.1%
Deviations of %error on test id #21. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | 0.071% | 0% | 0% | 0.1%
21 | 1 | 0.085% | 0.014% | 0% | 0.1%
21 | 2 | 0.044% | -0.041% | 0.067% | 0.1%
21 | 3 | 0.089% | 0.045% | 0.073% | 0.1%
Deviations of %error on test id #22. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | 0.089% | 0% | 0% | 0.1%
22 | 1 | 0.089% | 0% | 0% | 0.1%
22 | 2 | 0.134% | 0.045% | 0.104% | 0.1%
22 | 3 | 0.04% | -0.094% | 0.088% | 0.1%
Deviations of %error on test id #23. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | 0.116% | 0% | 0% | 0.1%
23 | 1 | 0.116% | 0% | 0% | 0.1%
23 | 2 | 0.111% | -0.005% | 0.114% | 0.1%
23 | 3 | 0.107% | -0.004% | 0.111% | 0.1%
Deviations of %error on test id #24. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | 0.093% | 0% | 0% | 0.1%
24 | 1 | 0.125% | 0.032% | 0% | 0.1%
24 | 2 | 0.134% | 0.009% | 0.117% | 0.1%
24 | 3 | 0.076% | -0.058% | 0.112% | 0.1%
Deviations of %error on test id #25. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | 0.129% | 0% | 0% | 0.1%
25 | 1 | 0.067% | -0.062% | 0% | 0.1%
25 | 2 | 0.129% | 0.062% | 0.108% | 0.1%
25 | 3 | 0.089% | -0.04% | 0.095% | 0.1%
Deviations of %error on test id #26. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | 0.085% | 0% | 0% | 0.1%
26 | 1 | 0.111% | 0.026% | 0% | 0.1%
26 | 2 | 0.08% | -0.031% | 0.092% | 0.1%
26 | 3 | 0.093% | 0.013% | 0.095% | 0.1%
Deviations of %error on test id #27. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | 0.12% | 0% | 0% | 0.1%
27 | 1 | 0.062% | -0.058% | 0% | 0.1%
27 | 2 | 0.116% | 0.054% | 0.099% | 0.1%
27 | 3 | 0.08% | -0.036% | 0.086% | 0.1%
Deviations of %error on test id #28. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | 0.129% | 0% | 0% | 0.1%
28 | 1 | 0.071% | -0.058% | 0% | 0.1%
28 | 2 | 0.12% | 0.049% | 0.107% | 0.1%
28 | 3 | 0.08% | -0.04% | 0.09% | 0.1%
Deviations of %error on test id #29. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | 0.067% | 0% | 0% | 0.1%
29 | 1 | 0.111% | 0.044% | 0% | 0.1%
29 | 2 | 0.116% | 0.005% | 0.098% | 0.1%
29 | 3 | 0.08% | -0.036% | 0.102% | 0.1%
Deviations of %error on test id #30. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | 0.151% | 0% | 0% | 0.1%
30 | 1 | 0.178% | 0.027% | 0% | 0.1%
30 | 2 | 0.138% | -0.04% | 0.156% | 0.1%
30 | 3 | 0.107% | -0.031% | 0.141% | 0.1%
Deviations of %error on test id #31. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | 0.129% | 0% | 0% | 0.1%
31 | 1 | 0.129% | 0% | 0% | 0.1%
31 | 2 | 0.116% | -0.013% | 0.125% | 0.1%
31 | 3 | 0.165% | 0.049% | 0.137% | 0.1%
Deviations of %error on test id #32. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | 0.12% | 0% | 0% | 0.1%
32 | 1 | 0.089% | -0.031% | 0% | 0.1%
32 | 2 | 0.134% | 0.045% | 0.114% | 0.1%
32 | 3 | 0.089% | -0.045% | 0.104% | 0.1%
Deviations of %error on test id #33. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | 0.125% | 0% | 0% | 0.1%
33 | 1 | 0.085% | -0.04% | 0% | 0.1%
33 | 2 | 0.116% | 0.031% | 0.109% | 0.1%
33 | 3 | 0.071% | -0.045% | 0.091% | 0.1%
33 | 4 | 0.111% | 0.04% | 0.099% | 0.1%
Deviations of %error on test id #34. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | 0.125% | 0% | 0% | 0.1%
34 | 1 | 0.111% | -0.014% | 0% | 0.1%
34 | 2 | 0.089% | -0.022% | 0.108% | 0.1%
34 | 3 | 0.125% | 0.036% | 0.108% | 0.1%
Deviations of %error on test id #35. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | 0.129% | 0% | 0% | 0.1%
35 | 1 | 0.093% | -0.036% | 0% | 0.1%
35 | 2 | 0.12% | 0.027% | 0.114% | 0.1%
35 | 3 | 0.093% | -0.027% | 0.102% | 0.1%
Deviations of %error on test id #36. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | 0.111% | 0% | 0% | 0.1%
36 | 1 | 0.116% | 0.005% | 0% | 0.1%
36 | 2 | 0.116% | 0% | 0.114% | 0.1%
36 | 3 | 0.116% | 0% | 0.116% | 0.1%
Deviations of %error on test id #37. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | 0.12% | 0% | 0% | 0.1%
37 | 1 | 0.16% | 0.04% | 0% | 0.1%
37 | 2 | 0.12% | -0.04% | 0.133% | 0.1%
37 | 3 | 0.16% | 0.04% | 0.147% | 0.1%
Deviations of %error on test id #38. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | 0.147% | 0% | 0% | 0.1%
38 | 1 | 0.116% | -0.031% | 0% | 0.1%
38 | 2 | 0.147% | 0.031% | 0.137% | 0.1%
38 | 3 | 0.111% | -0.036% | 0.125% | 0.1%
Deviations of %error on test id #39. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | 0.116% | 0% | 0% | 0.1%
39 | 1 | 0.169% | 0.053% | 0% | 0.1%
39 | 2 | 0.129% | -0.04% | 0.138% | 0.1%
39 | 3 | 0.156% | 0.027% | 0.151% | 0.1%
Deviations of %error on test id #40. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | 0.165% | 0% | 0% | 0.1%
40 | 1 | 0.12% | -0.045% | 0% | 0.1%
40 | 2 | 0.12% | 0% | 0.135% | 0.1%
40 | 3 | 0.12% | 0% | 0.12% | 0.1%

## 32. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\CS\20230725 SMAT 3PH4W NPL CS SM Active Reactive Accuracy.zip'
Generated on 7/27/2023 2:24:56 PM by version Alpha5 - Build 2023 Jul. 26
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.065% | PASSED | 2023-07-25 20:51
Yes | 1 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.05% | PASSED | 2023-07-25 20:51
Yes | 2 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.062% | PASSED | 2023-07-25 20:52
Yes | 3 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.044% | PASSED | 2023-07-25 20:53
Yes | 4 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.03% | PASSED | 2023-07-25 20:53
Yes | 5 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.071% | PASSED | 2023-07-25 20:54
Yes | 6 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.122% | PASSED | 2023-07-25 20:55
Yes | 7 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.146% | PASSED | 2023-07-25 20:55
Yes | 8 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.125% | PASSED | 2023-07-25 20:56
Yes | 9 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.148% | PASSED | 2023-07-25 20:57
Yes | 10 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.142% | PASSED | 2023-07-25 20:57
Yes | 11 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.131% | PASSED | 2023-07-25 20:58
Yes | 12 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.163% | PASSED | 2023-07-25 20:59
Yes | 13 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.119% | PASSED | 2023-07-25 20:59
Yes | 14 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.19% | PASSED | 2023-07-25 21:00
Yes | 15 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.148% | PASSED | 2023-07-25 21:01
Yes | 16 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.193% | PASSED | 2023-07-25 21:01
Yes | 17 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.247% | PASSED | 2023-07-25 21:02
Yes | 18 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.261% | PASSED | 2023-07-25 21:03
Yes | 19 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.232% | PASSED | 2023-07-25 21:03
Yes | 20 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.226% | PASSED | 2023-07-25 21:04
Yes | 21 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.285% | PASSED | 2023-07-25 21:05
Yes | 22 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.217% | PASSED | 2023-07-25 21:05
Yes | 23 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.324% | PASSED | 2023-07-25 21:06
Yes | 24 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.3% | PASSED | 2023-07-25 21:07
Yes | 25 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.327% | PASSED | 2023-07-25 21:07
Yes | 26 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.485% | PASSED | 2023-07-25 21:08
Yes | 27 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.369% | PASSED | 2023-07-25 21:09
Yes | 28 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.36% | PASSED | 2023-07-25 21:10
Yes | 29 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.336% | PASSED | 2023-07-25 21:10
Yes | 30 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.518% | PASSED | 2023-07-25 21:11
Yes | 31 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.36% | PASSED | 2023-07-25 21:12
Yes | 32 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.324% | PASSED | 2023-07-25 21:12
Yes | 33 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.282% | PASSED | 2023-07-25 21:13
Yes | 34 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.306% | PASSED | 2023-07-25 21:13
Yes | 35 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.59% | PASSED | 2023-07-25 21:20
Yes | 36 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.345% | PASSED | 2023-07-25 21:15
Yes | 37 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.527% | PASSED | 2023-07-25 21:16
Yes | 38 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.393% | PASSED | 2023-07-25 21:16
Yes | 39 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.746% | PASSED | 2023-07-25 21:17
Yes | 40 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 0° | 0.5L, sin(ϕ), ϕ=30° | 0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.743% | PASSED | 2023-07-25 21:21
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -23.79% | 0% | 0% | 0.1%
0 | 1 | -0.697% | 23.09% | 0% | 0.1%
0 | 2 | -1.047% | -0.35% | -8.511% | 0.1%
0 | 3 | -0.653% | 0.394% | -0.799% | 0.1%
0 | 4 | 0.044% | 0.697% | -0.552% | 0.1%
0 | 5 | 0.062% | 0.018% | -0.182% | 0.1%
0 | 6 | 0.071% | 0.009% | 0.059% | 0.1%
0 | 7 | 0.062% | -0.009% | 0.065% | 0.1%
Deviations of %error on test id #1. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -70.28% | 0% | 0% | 0.1%
1 | 1 | -13.23% | 57.05% | 0% | 0.1%
1 | 2 | 0.142% | 13.37% | -27.79% | 0.1%
1 | 3 | -1.134% | -1.276% | -4.74% | 0.1%
1 | 4 | -0.425% | 0.709% | -0.472% | 0.1%
1 | 5 | -1.09% | -0.665% | -0.883% | 0.1%
1 | 6 | -0.177% | 0.913% | -0.564% | 0.1%
1 | 7 | 0.107% | 0.284% | -0.387% | 0.1%
1 | 8 | 0.312% | 0.205% | 0.081% | 0.1%
1 | 9 | 0.59% | 0.278% | 0.336% | 0.1%
1 | 10 | 0.071% | -0.519% | 0.324% | 0.1%
1 | 11 | 0.062% | -0.009% | 0.241% | 0.1%
1 | 12 | 0.044% | -0.018% | 0.059% | 0.1%
1 | 13 | 0.044% | 0% | 0.05% | 0.1%
Deviations of %error on test id #2. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -51.85% | 0% | 0% | 0.1%
2 | 1 | -4.41% | 47.44% | 0% | 0.1%
2 | 2 | -0.345% | 4.065% | -18.87% | 0.1%
2 | 3 | -1.055% | -0.71% | -1.937% | 0.1%
2 | 4 | -0.089% | 0.966% | -0.496% | 0.1%
2 | 5 | 0.071% | 0.16% | -0.358% | 0.1%
2 | 6 | 0.044% | -0.027% | 0.009% | 0.1%
2 | 7 | 0.062% | 0.018% | 0.059% | 0.1%
2 | 8 | 0.08% | 0.018% | 0.062% | 0.1%
Deviations of %error on test id #3. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -28.85% | 0% | 0% | 0.1%
3 | 1 | -0.292% | 28.56% | 0% | 0.1%
3 | 2 | -0.994% | -0.702% | -10.04% | 0.1%
3 | 3 | -1.926% | -0.932% | -1.071% | 0.1%
3 | 4 | -0.266% | 1.66% | -1.062% | 0.1%
3 | 5 | 0.08% | 0.346% | -0.704% | 0.1%
3 | 6 | 0.608% | 0.528% | 0.141% | 0.1%
3 | 7 | 0.536% | -0.072% | 0.408% | 0.1%
3 | 8 | 0.044% | -0.492% | 0.396% | 0.1%
3 | 9 | 0.044% | 0% | 0.208% | 0.1%
3 | 10 | 0.053% | 0.009% | 0.047% | 0.1%
3 | 11 | 0.036% | -0.017% | 0.044% | 0.1%
Deviations of %error on test id #4. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -39.9% | 0% | 0% | 0.1%
4 | 1 | -0.133% | 39.76% | 0% | 0.1%
4 | 2 | -0.863% | -0.73% | -13.63% | 0.1%
4 | 3 | -1.995% | -1.132% | -0.997% | 0.1%
4 | 4 | -0.319% | 1.676% | -1.059% | 0.1%
4 | 5 | 0.053% | 0.372% | -0.754% | 0.1%
4 | 6 | 0.482% | 0.429% | 0.072% | 0.1%
4 | 7 | 0.527% | 0.045% | 0.354% | 0.1%
4 | 8 | 0.044% | -0.483% | 0.351% | 0.1%
4 | 9 | 0.036% | -0.008% | 0.202% | 0.1%
4 | 10 | 0.027% | -0.009% | 0.036% | 0.1%
4 | 11 | 0.027% | 0% | 0.03% | 0.1%
Deviations of %error on test id #5. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -52.51% | 0% | 0% | 0.1%
5 | 1 | -0.151% | 52.36% | 0% | 0.1%
5 | 2 | -0.732% | -0.581% | -17.8% | 0.1%
5 | 3 | -0.469% | 0.263% | -0.451% | 0.1%
5 | 4 | -1.003% | -0.534% | -0.735% | 0.1%
5 | 5 | -0.213% | 0.79% | -0.562% | 0.1%
5 | 6 | 0.134% | 0.347% | -0.361% | 0.1%
5 | 7 | 0.357% | 0.223% | 0.093% | 0.1%
5 | 8 | 0.536% | 0.179% | 0.342% | 0.1%
5 | 9 | 0.08% | -0.456% | 0.324% | 0.1%
5 | 10 | 0.071% | -0.009% | 0.229% | 0.1%
5 | 11 | 0.071% | 0% | 0.074% | 0.1%
5 | 12 | 0.071% | 0% | 0.071% | 0.1%
Deviations of %error on test id #6. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | -47.25% | 0% | 0% | 0.1%
6 | 1 | -1.264% | 45.98% | 0% | 0.1%
6 | 2 | -1.142% | 0.122% | -16.55% | 0.1%
6 | 3 | -0.697% | 0.445% | -1.034% | 0.1%
6 | 4 | 0.053% | 0.75% | -0.595% | 0.1%
6 | 5 | 0.125% | 0.072% | -0.173% | 0.1%
6 | 6 | 0.125% | 0% | 0.101% | 0.1%
6 | 7 | 0.116% | -0.009% | 0.122% | 0.1%
Deviations of %error on test id #7. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | -6.894% | 0% | 0% | 0.1%
7 | 1 | -0.354% | 6.54% | 0% | 0.1%
7 | 2 | -1.402% | -1.048% | -2.883% | 0.1%
7 | 3 | -0.16% | 1.242% | -0.639% | 0.1%
7 | 4 | 0.151% | 0.311% | -0.47% | 0.1%
7 | 5 | 0.169% | 0.018% | 0.053% | 0.1%
7 | 6 | 0.134% | -0.035% | 0.151% | 0.1%
7 | 7 | 0.134% | 0% | 0.146% | 0.1%
Deviations of %error on test id #8. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | -50.44% | 0% | 0% | 0.1%
8 | 1 | 0.455% | 50.89% | 0% | 0.1%
8 | 2 | -0.451% | -0.906% | -16.81% | 0.1%
8 | 3 | -1.523% | -1.072% | -0.506% | 0.1%
8 | 4 | -0.837% | 0.686% | -0.937% | 0.1%
8 | 5 | 0.142% | 0.979% | -0.739% | 0.1%
8 | 6 | 0.357% | 0.215% | -0.113% | 0.1%
8 | 7 | 0.888% | 0.531% | 0.462% | 0.1%
8 | 8 | 0.151% | -0.737% | 0.465% | 0.1%
8 | 9 | 0.125% | -0.026% | 0.388% | 0.1%
8 | 10 | 0.134% | 0.009% | 0.137% | 0.1%
8 | 11 | 0.116% | -0.018% | 0.125% | 0.1%
Deviations of %error on test id #9. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | -58.37% | 0% | 0% | 0.1%
9 | 1 | 0.375% | 58.75% | 0% | 0.1%
9 | 2 | -0.266% | -0.641% | -19.42% | 0.1%
9 | 3 | -0.671% | -0.405% | -0.187% | 0.1%
9 | 4 | -0.75% | -0.079% | -0.562% | 0.1%
9 | 5 | -2.182% | -1.432% | -1.201% | 0.1%
9 | 6 | 0.018% | 2.2% | -0.971% | 0.1%
9 | 7 | 0.888% | 0.87% | -0.425% | 0.1%
9 | 8 | 1.69% | 0.802% | 0.865% | 0.1%
9 | 9 | 0.698% | -0.992% | 1.092% | 0.1%
9 | 10 | 0.151% | -0.547% | 0.846% | 0.1%
9 | 11 | 0.142% | -0.009% | 0.33% | 0.1%
9 | 12 | 0.151% | 0.009% | 0.148% | 0.1%
9 | 13 | 0.151% | 0% | 0.148% | 0.1%
Deviations of %error on test id #10. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | -31.73% | 0% | 0% | 0.1%
10 | 1 | 0.25% | 31.98% | 0% | 0.1%
10 | 2 | -1.116% | -1.366% | -10.87% | 0.1%
10 | 3 | -0.275% | 0.841% | -0.38% | 0.1%
10 | 4 | -0.899% | -0.624% | -0.763% | 0.1%
10 | 5 | 0.178% | 1.077% | -0.332% | 0.1%
10 | 6 | 0.241% | 0.063% | -0.16% | 0.1%
10 | 7 | 0.734% | 0.493% | 0.384% | 0.1%
10 | 8 | 0.25% | -0.484% | 0.408% | 0.1%
10 | 9 | 0.142% | -0.108% | 0.375% | 0.1%
10 | 10 | 0.142% | 0% | 0.178% | 0.1%
10 | 11 | 0.151% | 0.009% | 0.145% | 0.1%
10 | 12 | 0.134% | -0.017% | 0.142% | 0.1%
Deviations of %error on test id #11. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | -56.19% | 0% | 0% | 0.1%
11 | 1 | -3.242% | 52.95% | 0% | 0.1%
11 | 2 | -0.671% | 2.571% | -20.03% | 0.1%
11 | 3 | -0.583% | 0.088% | -1.499% | 0.1%
11 | 4 | 0.107% | 0.69% | -0.382% | 0.1%
11 | 5 | 0.125% | 0.018% | -0.117% | 0.1%
11 | 6 | 0.151% | 0.026% | 0.128% | 0.1%
11 | 7 | 0.116% | -0.035% | 0.131% | 0.1%
Deviations of %error on test id #12. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | -68.1% | 0% | 0% | 0.1%
12 | 1 | -11.02% | 57.08% | 0% | 0.1%
12 | 2 | -0.213% | 10.8% | -26.44% | 0.1%
12 | 3 | -1.368% | -1.155% | -4.199% | 0.1%
12 | 4 | -1.333% | 0.035% | -0.971% | 0.1%
12 | 5 | 0.089% | 1.422% | -0.871% | 0.1%
12 | 6 | 0.285% | 0.196% | -0.32% | 0.1%
12 | 7 | 0.987% | 0.702% | 0.454% | 0.1%
12 | 8 | 0.223% | -0.764% | 0.498% | 0.1%
12 | 9 | 0.169% | -0.054% | 0.46% | 0.1%
12 | 10 | 0.151% | -0.018% | 0.181% | 0.1%
12 | 11 | 0.169% | 0.018% | 0.163% | 0.1%
Deviations of %error on test id #13. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | -64.95% | 0% | 0% | 0.1%
13 | 1 | -6.523% | 58.42% | 0% | 0.1%
13 | 2 | -0.133% | 6.39% | -23.87% | 0.1%
13 | 3 | -1.549% | -1.416% | -2.735% | 0.1%
13 | 4 | -1.09% | 0.459% | -0.924% | 0.1%
13 | 5 | 0.098% | 1.188% | -0.847% | 0.1%
13 | 6 | 0.42% | 0.322% | -0.191% | 0.1%
13 | 7 | 0.843% | 0.423% | 0.454% | 0.1%
13 | 8 | 0.205% | -0.638% | 0.489% | 0.1%
13 | 9 | 0.116% | -0.089% | 0.388% | 0.1%
13 | 10 | 0.116% | 0% | 0.146% | 0.1%
13 | 11 | 0.125% | 0.009% | 0.119% | 0.1%
Deviations of %error on test id #14. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | 0.339% | 0% | 0% | 0.1%
14 | 1 | -0.609% | -0.948% | 0% | 0.1%
14 | 2 | -2.063% | -1.454% | -0.778% | 0.1%
14 | 3 | -0.186% | 1.877% | -0.953% | 0.1%
14 | 4 | 0.178% | 0.364% | -0.69% | 0.1%
14 | 5 | 0.906% | 0.728% | 0.299% | 0.1%
14 | 6 | 0.428% | -0.478% | 0.504% | 0.1%
14 | 7 | 0.196% | -0.232% | 0.51% | 0.1%
14 | 8 | 0.196% | 0% | 0.273% | 0.1%
14 | 9 | 0.178% | -0.018% | 0.19% | 0.1%
14 | 10 | 0.196% | 0.018% | 0.19% | 0.1%
Deviations of %error on test id #15. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | -4.304% | 0% | 0% | 0.1%
15 | 1 | -0.169% | 4.135% | 0% | 0.1%
15 | 2 | -0.627% | -0.458% | -1.7% | 0.1%
15 | 3 | -1.944% | -1.317% | -0.913% | 0.1%
15 | 4 | -0.636% | 1.308% | -1.069% | 0.1%
15 | 5 | 0.178% | 0.814% | -0.801% | 0.1%
15 | 6 | 1.452% | 1.274% | 0.331% | 0.1%
15 | 7 | 0.806% | -0.646% | 0.812% | 0.1%
15 | 8 | 0.142% | -0.664% | 0.8% | 0.1%
15 | 9 | 0.134% | -0.008% | 0.361% | 0.1%
15 | 10 | 0.169% | 0.035% | 0.148% | 0.1%
15 | 11 | 0.142% | -0.027% | 0.148% | 0.1%
Deviations of %error on test id #16. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | 0.258% | 0% | 0% | 0.1%
16 | 1 | -0.618% | -0.876% | 0% | 0.1%
16 | 2 | -0.381% | 0.237% | -0.247% | 0.1%
16 | 3 | -1.203% | -0.822% | -0.734% | 0.1%
16 | 4 | -0.899% | 0.304% | -0.828% | 0.1%
16 | 5 | 0.196% | 1.095% | -0.635% | 0.1%
16 | 6 | 0.951% | 0.755% | 0.083% | 0.1%
16 | 7 | 1.096% | 0.145% | 0.748% | 0.1%
16 | 8 | 0.214% | -0.882% | 0.754% | 0.1%
16 | 9 | 0.196% | -0.018% | 0.502% | 0.1%
16 | 10 | 0.196% | 0% | 0.202% | 0.1%
16 | 11 | 0.187% | -0.009% | 0.193% | 0.1%
Deviations of %error on test id #17. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | -12.84% | 0% | 0% | 0.1%
17 | 1 | 0.098% | 12.93% | 0% | 0.1%
17 | 2 | -0.636% | -0.734% | -4.458% | 0.1%
17 | 3 | -0.776% | -0.14% | -0.438% | 0.1%
17 | 4 | -0.009% | 0.767% | -0.474% | 0.1%
17 | 5 | 0.42% | 0.429% | -0.122% | 0.1%
17 | 6 | 0.942% | 0.522% | 0.451% | 0.1%
17 | 7 | 0.536% | -0.406% | 0.633% | 0.1%
17 | 8 | 0.25% | -0.286% | 0.576% | 0.1%
17 | 9 | 0.25% | 0% | 0.345% | 0.1%
17 | 10 | 0.232% | -0.018% | 0.244% | 0.1%
17 | 11 | 0.258% | 0.026% | 0.247% | 0.1%
Deviations of %error on test id #18. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | -20.09% | 0% | 0% | 0.1%
18 | 1 | 0.321% | 20.41% | 0% | 0.1%
18 | 2 | -0.177% | -0.498% | -6.65% | 0.1%
18 | 3 | -1.047% | -0.87% | -0.301% | 0.1%
18 | 4 | 0.205% | 1.252% | -0.34% | 0.1%
18 | 5 | 0.491% | 0.286% | -0.117% | 0.1%
18 | 6 | 0.978% | 0.487% | 0.558% | 0.1%
18 | 7 | 0.446% | -0.532% | 0.638% | 0.1%
18 | 8 | 0.267% | -0.179% | 0.564% | 0.1%
18 | 9 | 0.267% | 0% | 0.327% | 0.1%
18 | 10 | 0.258% | -0.009% | 0.264% | 0.1%
18 | 11 | 0.258% | 0% | 0.261% | 0.1%
Deviations of %error on test id #19. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | -53.01% | 0% | 0% | 0.1%
19 | 1 | 0.151% | 53.16% | 0% | 0.1%
19 | 2 | -1.09% | -1.241% | -17.98% | 0.1%
19 | 3 | -1.566% | -0.476% | -0.835% | 0.1%
19 | 4 | 0.125% | 1.691% | -0.844% | 0.1%
19 | 5 | 0.348% | 0.223% | -0.364% | 0.1%
19 | 6 | 1.042% | 0.694% | 0.505% | 0.1%
19 | 7 | 0.285% | -0.757% | 0.558% | 0.1%
19 | 8 | 0.223% | -0.062% | 0.517% | 0.1%
19 | 9 | 0.232% | 0.009% | 0.247% | 0.1%
19 | 10 | 0.241% | 0.009% | 0.232% | 0.1%
Deviations of %error on test id #20. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | -20.18% | 0% | 0% | 0.1%
20 | 1 | -0.968% | 19.22% | 0% | 0.1%
20 | 2 | -0.653% | 0.315% | -7.268% | 0.1%
20 | 3 | 0.062% | 0.715% | -0.52% | 0.1%
20 | 4 | 0.241% | 0.179% | -0.117% | 0.1%
20 | 5 | 0.214% | -0.027% | 0.172% | 0.1%
20 | 6 | 0.223% | 0.009% | 0.226% | 0.1%
20 | 7 | 0.241% | 0.018% | 0.226% | 0.1%
Deviations of %error on test id #21. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | -18.56% | 0% | 0% | 0.1%
21 | 1 | -0.089% | 18.47% | 0% | 0.1%
21 | 2 | -0.292% | -0.203% | -6.314% | 0.1%
21 | 3 | -1.532% | -1.24% | -0.638% | 0.1%
21 | 4 | -0.907% | 0.625% | -0.91% | 0.1%
21 | 5 | 0.312% | 1.219% | -0.709% | 0.1%
21 | 6 | 1.534% | 1.222% | 0.313% | 0.1%
21 | 7 | 1.033% | -0.501% | 0.96% | 0.1%
21 | 8 | 0.303% | -0.73% | 0.957% | 0.1%
21 | 9 | 0.276% | -0.027% | 0.537% | 0.1%
21 | 10 | 0.285% | 0.009% | 0.288% | 0.1%
21 | 11 | 0.294% | 0.009% | 0.285% | 0.1%
Deviations of %error on test id #22. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | -22.58% | 0% | 0% | 0.1%
22 | 1 | -0.142% | 22.44% | 0% | 0.1%
22 | 2 | -0.425% | -0.283% | -7.715% | 0.1%
22 | 3 | -0.486% | -0.061% | -0.351% | 0.1%
22 | 4 | 0.027% | 0.513% | -0.295% | 0.1%
22 | 5 | 0.446% | 0.419% | -0.004% | 0.1%
22 | 6 | 0.888% | 0.442% | 0.454% | 0.1%
22 | 7 | 0.473% | -0.415% | 0.602% | 0.1%
22 | 8 | 0.267% | -0.206% | 0.543% | 0.1%
22 | 9 | 0.223% | -0.044% | 0.321% | 0.1%
22 | 10 | 0.223% | 0% | 0.238% | 0.1%
22 | 11 | 0.205% | -0.018% | 0.217% | 0.1%
Deviations of %error on test id #23. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | -40.06% | 0% | 0% | 0.1%
23 | 1 | -0.018% | 40.04% | 0% | 0.1%
23 | 2 | -0.036% | -0.018% | -13.37% | 0.1%
23 | 3 | -0.794% | -0.758% | -0.283% | 0.1%
23 | 4 | 0.241% | 1.035% | -0.196% | 0.1%
23 | 5 | 0.563% | 0.322% | 0.003% | 0.1%
23 | 6 | 0.996% | 0.433% | 0.6% | 0.1%
23 | 7 | 0.455% | -0.541% | 0.671% | 0.1%
23 | 8 | 0.312% | -0.143% | 0.588% | 0.1%
23 | 9 | 0.339% | 0.027% | 0.369% | 0.1%
23 | 10 | 0.294% | -0.045% | 0.315% | 0.1%
23 | 11 | 0.339% | 0.045% | 0.324% | 0.1%
Deviations of %error on test id #24. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | -37.15% | 0% | 0% | 0.1%
24 | 1 | 0.151% | 37.3% | 0% | 0.1%
24 | 2 | -0.124% | -0.275% | -12.38% | 0.1%
24 | 3 | -0.75% | -0.626% | -0.241% | 0.1%
24 | 4 | -1.841% | -1.091% | -0.905% | 0.1%
24 | 5 | 0.25% | 2.091% | -0.78% | 0.1%
24 | 6 | 0.933% | 0.683% | -0.219% | 0.1%
24 | 7 | 1.562% | 0.629% | 0.915% | 0.1%
24 | 8 | 0.348% | -1.214% | 0.948% | 0.1%
24 | 9 | 0.303% | -0.045% | 0.738% | 0.1%
24 | 10 | 0.294% | -0.009% | 0.315% | 0.1%
24 | 11 | 0.303% | 0.009% | 0.3% | 0.1%
Deviations of %error on test id #25. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | -16.39% | 0% | 0% | 0.1%
25 | 1 | -0.009% | 16.38% | 0% | 0.1%
25 | 2 | -0.275% | -0.266% | -5.559% | 0.1%
25 | 3 | -0.741% | -0.466% | -0.342% | 0.1%
25 | 4 | 0.491% | 1.232% | -0.175% | 0.1%
25 | 5 | 0.572% | 0.081% | 0.107% | 0.1%
25 | 6 | 1.033% | 0.461% | 0.699% | 0.1%
25 | 7 | 0.348% | -0.685% | 0.651% | 0.1%
25 | 8 | 0.321% | -0.027% | 0.567% | 0.1%
25 | 9 | 0.33% | 0.009% | 0.333% | 0.1%
25 | 10 | 0.33% | 0% | 0.327% | 0.1%
Deviations of %error on test id #26. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | -35.06% | 0% | 0% | 0.1%
26 | 1 | -0.009% | 35.05% | 0% | 0.1%
26 | 2 | -0.863% | -0.854% | -11.98% | 0.1%
26 | 3 | 0.018% | 0.881% | -0.285% | 0.1%
26 | 4 | 0.339% | 0.321% | -0.169% | 0.1%
26 | 5 | 0.428% | 0.089% | 0.262% | 0.1%
26 | 6 | 0.554% | 0.126% | 0.44% | 0.1%
26 | 7 | 0.518% | -0.036% | 0.5% | 0.1%
26 | 8 | 0.437% | -0.081% | 0.503% | 0.1%
26 | 9 | 0.5% | 0.063% | 0.485% | 0.1%
Deviations of %error on test id #27. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | -6.956% | 0% | 0% | 0.1%
27 | 1 | -0.177% | 6.779% | 0% | 0.1%
27 | 2 | -1.781% | -1.604% | -2.971% | 0.1%
27 | 3 | 0.178% | 1.959% | -0.593% | 0.1%
27 | 4 | 0.464% | 0.286% | -0.38% | 0.1%
27 | 5 | 1.388% | 0.924% | 0.677% | 0.1%
27 | 6 | 0.464% | -0.924% | 0.772% | 0.1%
27 | 7 | 0.339% | -0.125% | 0.73% | 0.1%
27 | 8 | 0.42% | 0.081% | 0.408% | 0.1%
27 | 9 | 0.446% | 0.026% | 0.402% | 0.1%
27 | 10 | 0.581% | 0.135% | 0.482% | 0.1%
27 | 11 | 0.464% | -0.117% | 0.497% | 0.1%
27 | 12 | 0.375% | -0.089% | 0.473% | 0.1%
27 | 13 | 0.545% | 0.17% | 0.461% | 0.1%
27 | 14 | 0.312% | -0.233% | 0.411% | 0.1%
27 | 15 | 0.348% | 0.036% | 0.402% | 0.1%
27 | 16 | 0.563% | 0.215% | 0.408% | 0.1%
27 | 17 | 0.33% | -0.233% | 0.414% | 0.1%
27 | 18 | 0.428% | 0.098% | 0.44% | 0.1%
27 | 19 | 0.348% | -0.08% | 0.369% | 0.1%
Deviations of %error on test id #28. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | -43.76% | 0% | 0% | 0.1%
28 | 1 | 0.134% | 43.89% | 0% | 0.1%
28 | 2 | 0.25% | 0.116% | -14.46% | 0.1%
28 | 3 | -1.082% | -1.332% | -0.233% | 0.1%
28 | 4 | 0.232% | 1.314% | -0.2% | 0.1%
28 | 5 | 0.671% | 0.439% | -0.06% | 0.1%
28 | 6 | 1.224% | 0.553% | 0.709% | 0.1%
28 | 7 | 0.464% | -0.76% | 0.786% | 0.1%
28 | 8 | 0.375% | -0.089% | 0.688% | 0.1%
28 | 9 | 0.357% | -0.018% | 0.399% | 0.1%
28 | 10 | 0.348% | -0.009% | 0.36% | 0.1%
Deviations of %error on test id #29. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | -19.27% | 0% | 0% | 0.1%
29 | 1 | 0.232% | 19.5% | 0% | 0.1%
29 | 2 | -0.292% | -0.524% | -6.444% | 0.1%
29 | 3 | -0.776% | -0.484% | -0.279% | 0.1%
29 | 4 | 0.635% | 1.411% | -0.144% | 0.1%
29 | 5 | 0.879% | 0.244% | 0.246% | 0.1%
29 | 6 | 0.897% | 0.018% | 0.804% | 0.1%
29 | 7 | 0.339% | -0.558% | 0.705% | 0.1%
29 | 8 | 0.348% | 0.009% | 0.528% | 0.1%
29 | 9 | 0.339% | -0.009% | 0.342% | 0.1%
29 | 10 | 0.321% | -0.018% | 0.336% | 0.1%
Deviations of %error on test id #30. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | -55.53% | 0% | 0% | 0.1%
30 | 1 | 0.68% | 56.21% | 0% | 0.1%
30 | 2 | -0.213% | -0.893% | -18.35% | 0.1%
30 | 3 | -0.31% | -0.097% | 0.052% | 0.1%
30 | 4 | -0.257% | 0.053% | -0.26% | 0.1%
30 | 5 | 0.689% | 0.946% | 0.041% | 0.1%
30 | 6 | 1.306% | 0.617% | 0.579% | 0.1%
30 | 7 | 0.68% | -0.626% | 0.892% | 0.1%
30 | 8 | 0.375% | -0.305% | 0.787% | 0.1%
30 | 9 | 0.402% | 0.027% | 0.486% | 0.1%
30 | 10 | 0.608% | 0.206% | 0.462% | 0.1%
30 | 11 | 0.689% | 0.081% | 0.566% | 0.1%
30 | 12 | 0.473% | -0.216% | 0.59% | 0.1%
30 | 13 | 0.518% | 0.045% | 0.56% | 0.1%
30 | 14 | 0.536% | 0.018% | 0.509% | 0.1%
30 | 15 | 0.5% | -0.036% | 0.518% | 0.1%
Deviations of %error on test id #31. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | -19.75% | 0% | 0% | 0.1%
31 | 1 | 0.707% | 20.46% | 0% | 0.1%
31 | 2 | -0.08% | -0.787% | -6.375% | 0.1%
31 | 3 | -0.829% | -0.749% | -0.067% | 0.1%
31 | 4 | 0.491% | 1.32% | -0.139% | 0.1%
31 | 5 | 0.689% | 0.198% | 0.117% | 0.1%
31 | 6 | 1.105% | 0.416% | 0.762% | 0.1%
31 | 7 | 0.411% | -0.694% | 0.735% | 0.1%
31 | 8 | 0.357% | -0.054% | 0.624% | 0.1%
31 | 9 | 0.357% | 0% | 0.375% | 0.1%
31 | 10 | 0.366% | 0.009% | 0.36% | 0.1%
Deviations of %error on test id #32. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | -39.68% | 0% | 0% | 0.1%
32 | 1 | 0.232% | 39.91% | 0% | 0.1%
32 | 2 | -0.916% | -1.148% | -13.46% | 0.1%
32 | 3 | 0.08% | 0.996% | -0.201% | 0.1%
32 | 4 | 0.33% | 0.25% | -0.169% | 0.1%
32 | 5 | 0.321% | -0.009% | 0.244% | 0.1%
32 | 6 | 0.33% | 0.009% | 0.327% | 0.1%
32 | 7 | 0.321% | -0.009% | 0.324% | 0.1%
Deviations of %error on test id #33. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | 0.16% | 0% | 0% | 0.1%
33 | 1 | -0.968% | -1.128% | 0% | 0.1%
33 | 2 | 0.178% | 1.146% | -0.21% | 0.1%
33 | 3 | 0.276% | 0.098% | -0.171% | 0.1%
33 | 4 | 0.276% | 0% | 0.243% | 0.1%
33 | 5 | 0.294% | 0.018% | 0.282% | 0.1%
Deviations of %error on test id #34. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | -4.83% | 0% | 0% | 0.1%
34 | 1 | -0.724% | 4.106% | 0% | 0.1%
34 | 2 | -0.292% | 0.432% | -1.949% | 0.1%
34 | 3 | 0.375% | 0.667% | -0.214% | 0.1%
34 | 4 | 0.303% | -0.072% | 0.129% | 0.1%
34 | 5 | 0.312% | 0.009% | 0.33% | 0.1%
34 | 6 | 0.303% | -0.009% | 0.306% | 0.1%
Deviations of %error on test id #35. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | -13.88% | 0% | 0% | 0.1%
35 | 1 | 0.608% | 14.49% | 0% | 0.1%
35 | 2 | 0.411% | -0.197% | -4.286% | 0.1%
35 | 3 | -0.266% | -0.677% | 0.251% | 0.1%
35 | 4 | -2.174% | -1.908% | -0.676% | 0.1%
35 | 5 | 1.087% | 3.261% | -0.451% | 0.1%
35 | 6 | 2.059% | 0.972% | 0.324% | 0.1%
35 | 7 | 2.58% | 0.521% | 1.909% | 0.1%
35 | 8 | 0.509% | -2.071% | 1.716% | 0.1%
35 | 9 | 0.321% | -0.188% | 1.137% | 0.1%
35 | 10 | 0.375% | 0.054% | 0.402% | 0.1%
35 | 11 | 0.635% | 0.26% | 0.444% | 0.1%
35 | 12 | 0.635% | 0% | 0.548% | 0.1%
35 | 13 | 0.42% | -0.215% | 0.563% | 0.1%
35 | 14 | 0.366% | -0.054% | 0.474% | 0.1%
35 | 15 | 0.527% | 0.161% | 0.438% | 0.1%
35 | 16 | 0.626% | 0.099% | 0.506% | 0.1%
35 | 17 | 0.635% | 0.009% | 0.596% | 0.1%
35 | 18 | 0.491% | -0.144% | 0.584% | 0.1%
35 | 19 | 0.644% | 0.153% | 0.59% | 0.1%
Deviations of %error on test id #36. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | -6.242% | 0% | 0% | 0.1%
36 | 1 | 0.294% | 6.536% | 0% | 0.1%
36 | 2 | -0.811% | -1.105% | -2.253% | 0.1%
36 | 3 | 0.276% | 1.087% | -0.08% | 0.1%
36 | 4 | 0.617% | 0.341% | 0.027% | 0.1%
36 | 5 | 1.124% | 0.507% | 0.672% | 0.1%
36 | 6 | 0.384% | -0.74% | 0.708% | 0.1%
36 | 7 | 0.33% | -0.054% | 0.613% | 0.1%
36 | 8 | 0.339% | 0.009% | 0.351% | 0.1%
36 | 9 | 0.366% | 0.027% | 0.345% | 0.1%
Deviations of %error on test id #37. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | -6.172% | 0% | 0% | 0.1%
37 | 1 | -0.009% | 6.163% | 0% | 0.1%
37 | 2 | -0.31% | -0.301% | -2.164% | 0.1%
37 | 3 | 0.671% | 0.981% | 0.117% | 0.1%
37 | 4 | 0.473% | -0.198% | 0.278% | 0.1%
37 | 5 | 0.464% | -0.009% | 0.536% | 0.1%
37 | 6 | 0.626% | 0.162% | 0.521% | 0.1%
37 | 7 | 0.716% | 0.09% | 0.602% | 0.1%
37 | 8 | 0.707% | -0.009% | 0.683% | 0.1%
37 | 9 | 0.545% | -0.162% | 0.656% | 0.1%
37 | 10 | 0.473% | -0.072% | 0.575% | 0.1%
37 | 11 | 0.464% | -0.009% | 0.494% | 0.1%
37 | 12 | 0.572% | 0.108% | 0.503% | 0.1%
37 | 13 | 0.689% | 0.117% | 0.575% | 0.1%
37 | 14 | 0.861% | 0.172% | 0.707% | 0.1%
37 | 15 | 0.698% | -0.163% | 0.749% | 0.1%
37 | 16 | 0.951% | 0.253% | 0.837% | 0.1%
37 | 17 | 0.482% | -0.469% | 0.71% | 0.1%
37 | 18 | 0.671% | 0.189% | 0.701% | 0.1%
37 | 19 | 0.428% | -0.243% | 0.527% | 0.1%
Deviations of %error on test id #38. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | -8.208% | 0% | 0% | 0.1%
38 | 1 | 0.608% | 8.816% | 0% | 0.1%
38 | 2 | -0.398% | -1.006% | -2.666% | 0.1%
38 | 3 | 0.018% | 0.416% | 0.076% | 0.1%
38 | 4 | 0.806% | 0.788% | 0.142% | 0.1%
38 | 5 | 1.36% | 0.554% | 0.728% | 0.1%
38 | 6 | 0.825% | -0.535% | 0.997% | 0.1%
38 | 7 | 0.518% | -0.307% | 0.901% | 0.1%
38 | 8 | 0.428% | -0.09% | 0.59% | 0.1%
38 | 9 | 0.393% | -0.035% | 0.446% | 0.1%
38 | 10 | 0.357% | -0.036% | 0.393% | 0.1%
Deviations of %error on test id #39. Total 18 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | -7.968% | 0% | 0% | 0.1%
39 | 1 | -0.724% | 7.244% | 0% | 0.1%
39 | 2 | 0% | 0.724% | -2.897% | 0.1%
39 | 3 | 0.626% | 0.626% | -0.033% | 0.1%
39 | 4 | 0.843% | 0.217% | 0.49% | 0.1%
39 | 5 | 0.698% | -0.145% | 0.722% | 0.1%
39 | 6 | 1.042% | 0.344% | 0.861% | 0.1%
39 | 7 | 0.978% | -0.064% | 0.906% | 0.1%
39 | 8 | 0.689% | -0.289% | 0.903% | 0.1%
39 | 9 | 1.142% | 0.453% | 0.936% | 0.1%
39 | 10 | 0.924% | -0.218% | 0.918% | 0.1%
39 | 11 | 1.105% | 0.181% | 1.057% | 0.1%
39 | 12 | 1.06% | -0.045% | 1.03% | 0.1%
39 | 13 | 0.96% | -0.1% | 1.042% | 0.1%
39 | 14 | 0.788% | -0.172% | 0.936% | 0.1%
39 | 15 | 0.734% | -0.054% | 0.827% | 0.1%
39 | 16 | 0.752% | 0.018% | 0.758% | 0.1%
39 | 17 | 0.752% | 0% | 0.746% | 0.1%
Deviations of %error on test id #40. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | 102.8% | 0% | 0% | 0.1%
40 | 1 | 0.348% | -102.4% | 0% | 0.1%
40 | 2 | 0.267% | -0.081% | 34.46% | 0.1%
40 | 3 | -0.186% | -0.453% | 0.143% | 0.1%
40 | 4 | 0.987% | 1.173% | 0.356% | 0.1%
40 | 5 | 1.16% | 0.173% | 0.654% | 0.1%
40 | 6 | 1.37% | 0.21% | 1.172% | 0.1%
40 | 7 | 0.951% | -0.419% | 1.16% | 0.1%
40 | 8 | 0.77% | -0.181% | 1.03% | 0.1%
40 | 9 | 0.761% | -0.009% | 0.827% | 0.1%
40 | 10 | 0.77% | 0.009% | 0.767% | 0.1%
40 | 11 | 0.698% | -0.072% | 0.743% | 0.1%