## Summary
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:12:33 PM by version Alpha5 - Build 2023 Jul. 27
1. Configuration for 3P4W Meter
Equipment Setup
Option Name |  |  |  |  | Value
Communication Port |  |  |  |  | COM15
Communication Port Settings |  |  |  |  | 9600, 8, None, One
Timeout, Delay and Retries |  |  |  |  | 0, 50, 5000, 0, 100, 5000, 1
Equipment Version |  |  |  |  | GF302D
Control phase angle between YR (U1U2) and YB (U1U3) |  |  |  |  | Yes
Indicates RYB channels by colors |  |  |  |  | Yes
Meter Platform Name |  |  |  |  | Global Reference Meter
Serial Number |  |  |  |  | SED-IV0004
Meter Type |  |  |  |  | TYPE_3P4W
Meter Class |  |  |  |  | Class 0.5S
Meter Rating |  |  |  |  | Non-specified
Meter Maximum Current (Imax) |  |  |  |  | 40
Meter Basic Voltage (Un or Ub) |  |  |  |  | 220
Meter Basic Current (Ib or In) |  |  |  |  | 5
Meter Frequency (Fb) |  |  |  |  | 50
Meter Constant Active Power (P) |  |  |  |  | 3200
Meter Constant Reactive Power (Q) |  |  |  |  | 3200
Meter % Error Average Length |  |  |  |  | 3
Meter % Error Rounding Decimals |  |  |  |  | 3
Discard first few %error responses |  |  |  |  | Yes
Discard first few %error, number of responses will be discarded |  |  |  |  | 1
Chain voltage or current among test cases |  |  |  |  | No
Meter Configuration
Option Name |  |  |  |  | Value
Optical Port
Optical Port Settings |  |  |  |  | 9600, 8, None, One
External COM Port
External COM Port Settings |  |  |  |  | 19200, 8, None, One
Private Protocol Settings |  |  |  |  | abcd.efgh.xx@renesas.com, 0, 50, 100, 20000, 1
2. Accuracy Test
Below table summarizes all test categories and test result
No | Path In Workspace | Test Category | Passed | Failed | Inconclusive | Total | Overall Result | Accuracy Chart | Details
1 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / PF = 1.0 | PF = 1.0 | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart1. PF = 1.0.png | [>>]
2 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / PF = 0.5L | PF = 0.5L | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart2. PF = 0.5L.png | [>>]
3 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / PF = 0.8C | PF = 0.8C | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart3. PF = 0.8C.png | [>>]
4 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / 5th Harmonics | 5th Harmonics | 1 | 0 | 0 | 1 | PASSED | Temp\AccuracyChart4. 5th Harmonics.png | [>>]
5 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 1.0 | Voltage Variation, PF = 1.0 | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart5. Voltage Variatio....png | [>>]
6 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 0.5L | Voltage Variation, PF = 0.5L | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart6. Voltage Variatio....png | [>>]
7 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 1.0 | Frequency Variation, PF = 1.0 | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart7. Frequency Varia....png | [>>]
8 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 0.5L | Frequency Variation, PF = 0.5L | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart8. Frequency Varia....png | [>>]
9 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / PF = 1.0 | PF = 1.0 | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart9. PF = 1.0.png | [>>]
10 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / PF = 0.5L | PF = 0.5L | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart10. PF = 0.5L.png | [>>]
11 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / PF = 0.8C | PF = 0.8C | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart11. PF = 0.8C.png | [>>]
12 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / 5th Harmonics | 5th Harmonics | 1 | 0 | 0 | 1 | PASSED | Temp\AccuracyChart12. 5th Harmonics.png | [>>]
13 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 1.0 | Voltage Variation, PF = 1.0 | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart13. Voltage Variatio....png | [>>]
14 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 0.5L | Voltage Variation, PF = 0.5L | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart14. Voltage Variatio....png | [>>]
15 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 1.0 | Frequency Variation, PF = 1.0 | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart15. Frequency Varia....png | [>>]
16 | Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 0.5L | Frequency Variation, PF = 0.5L | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart16. Frequency Varia....png | [>>]
17 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / PF = 1.0 | PF = 1.0 | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart17. PF = 1.0.png | [>>]
18 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / PF = 0.5L | PF = 0.5L | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart18. PF = 0.5L.png | [>>]
19 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / PF = 0.8C | PF = 0.8C | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart19. PF = 0.8C.png | [>>]
20 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / 5th Harmonics | 5th Harmonics | 1 | 0 | 0 | 1 | PASSED | Temp\AccuracyChart20. 5th Harmonics.png | [>>]
21 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 1.0 | Voltage Variation, PF = 1.0 | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart21. Voltage Variatio....png | [>>]
22 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 0.5L | Voltage Variation, PF = 0.5L | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart22. Voltage Variatio....png | [>>]
23 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 1.0 | Frequency Variation, PF = 1.0 | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart23. Frequency Varia....png | [>>]
24 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 0.5L | Frequency Variation, PF = 0.5L | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart24. Frequency Varia....png | [>>]
25 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / PF = 1.0 | PF = 1.0 | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart25. PF = 1.0.png | [>>]
26 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / PF = 0.5L | PF = 0.5L | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart26. PF = 0.5L.png | [>>]
27 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / PF = 0.8C | PF = 0.8C | 5 | 0 | 0 | 5 | PASSED | Temp\AccuracyChart27. PF = 0.8C.png | [>>]
28 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / 5th Harmonics | 5th Harmonics | 1 | 0 | 0 | 1 | PASSED | Temp\AccuracyChart28. 5th Harmonics.png | [>>]
29 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 1.0 | Voltage Variation, PF = 1.0 | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart29. Voltage Variatio....png | [>>]
30 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 0.5L | Voltage Variation, PF = 0.5L | 6 | 0 | 0 | 6 | PASSED | Temp\AccuracyChart30. Voltage Variatio....png | [>>]
31 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 1.0 | Frequency Variation, PF = 1.0 | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart31. Frequency Varia....png | [>>]
32 | Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 0.5L | Frequency Variation, PF = 0.5L | 41 | 0 | 0 | 41 | PASSED | Temp\AccuracyChart32. Frequency Varia....png | [>>]

## 1. PF = 1.0
[<<]
Accuracy Test Report of 'PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:07:52 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 20 | 20 |  | ±0.1% | ±0.5% | 0.012% | PASSED | 2023-07-27 10:44
Yes | 1 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 12 | 20 |  | ±0.1% | ±0.5% | -0.002% | PASSED | 2023-07-27 10:44
Yes | 3 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.5% | 0.016% | PASSED | 2023-07-27 10:45
Yes | 4 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0.027% | PASSED | 2023-07-27 10:45
Yes | 7 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0.036% | PASSED | 2023-07-27 10:46
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 0.018% | 0% | 0% | 0.1%
0 | 1 | 0.018% | 0% | 0% | 0.1%
0 | 2 | 0.018% | 0% | 0.018% | 0.1%
0 | 3 | 0% | -0.018% | 0.012% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -0.068% | 0% | 0% | 0.1%
1 | 1 | -0.006% | 0.062% | 0% | 0.1%
1 | 2 | -0.009% | -0.003% | -0.028% | 0.1%
1 | 3 | 0.009% | 0.018% | -0.002% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.004% | 0% | 0% | 0.1%
3 | 1 | 0.018% | 0.014% | 0% | 0.1%
3 | 2 | 0.022% | 0.004% | 0.015% | 0.1%
3 | 3 | 0.009% | -0.013% | 0.016% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.027% | 0% | 0% | 0.1%
4 | 1 | 0.009% | -0.018% | 0% | 0.1%
4 | 2 | 0.018% | 0.009% | 0.018% | 0.1%
4 | 3 | 0.053% | 0.035% | 0.027% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.028% | 0% | 0% | 0.1%
7 | 1 | 0.03% | 0.002% | 0% | 0.1%
7 | 2 | 0.041% | 0.011% | 0.033% | 0.1%
7 | 3 | 0.036% | -0.005% | 0.036% | 0.1%

## 2. PF = 0.5L
[<<]
Accuracy Test Report of 'PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:07:54 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | -0.112% | PASSED | 2023-07-27 10:47
Yes | 1 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | -0.175% | PASSED | 2023-07-27 10:47
Yes | 3 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.354% | PASSED | 2023-07-27 10:48
Yes | 4 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.047% | PASSED | 2023-07-27 10:49
Yes | 7 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.076% | PASSED | 2023-07-27 10:50
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 18 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -35.2% | 0% | 0% | 0.1%
0 | 1 | -29.96% | 5.237% | 0% | 0.1%
0 | 2 | -21.02% | 8.939% | -28.72% | 0.1%
0 | 3 | -4.758% | 16.26% | -18.58% | 0.1%
0 | 4 | -1.798% | 2.96% | -9.192% | 0.1%
0 | 5 | -0.124% | 1.674% | -2.227% | 0.1%
0 | 6 | 0.393% | 0.517% | -0.51% | 0.1%
0 | 7 | 0.267% | -0.126% | 0.179% | 0.1%
0 | 8 | 0.232% | -0.035% | 0.297% | 0.1%
0 | 9 | 0.626% | 0.394% | 0.375% | 0.1%
0 | 10 | 1.133% | 0.507% | 0.664% | 0.1%
0 | 11 | 1.096% | -0.037% | 0.952% | 0.1%
0 | 12 | 0.752% | -0.344% | 0.994% | 0.1%
0 | 13 | 0.339% | -0.413% | 0.729% | 0.1%
0 | 14 | -0.106% | -0.445% | 0.328% | 0.1%
0 | 15 | -0.106% | 0% | 0.042% | 0.1%
0 | 16 | -0.106% | 0% | -0.106% | 0.1%
0 | 17 | -0.124% | -0.018% | -0.112% | 0.1%
Deviations of %error on test id #1. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -13.62% | 0% | 0% | 0.1%
1 | 1 | -2.878% | 10.74% | 0% | 0.1%
1 | 2 | 0.142% | 3.02% | -5.452% | 0.1%
1 | 3 | 0.03% | -0.112% | -0.902% | 0.1%
1 | 4 | -0.089% | -0.119% | 0.028% | 0.1%
1 | 5 | 0.232% | 0.321% | 0.058% | 0.1%
1 | 6 | 0.154% | -0.078% | 0.099% | 0.1%
1 | 7 | 0.012% | -0.142% | 0.133% | 0.1%
1 | 8 | 0.405% | 0.393% | 0.19% | 0.1%
1 | 9 | 0.47% | 0.065% | 0.296% | 0.1%
1 | 10 | 0.267% | -0.203% | 0.381% | 0.1%
1 | 11 | 0.047% | -0.22% | 0.261% | 0.1%
1 | 12 | -0.077% | -0.124% | 0.079% | 0.1%
1 | 13 | -0.083% | -0.006% | -0.038% | 0.1%
1 | 14 | -0.118% | -0.035% | -0.093% | 0.1%
1 | 15 | -0.325% | -0.207% | -0.175% | 0.1%
1 | 16 | -0.389% | -0.064% | -0.277% | 0.1%
1 | 17 | -0.26% | 0.129% | -0.325% | 0.1%
1 | 18 | -0.154% | 0.106% | -0.268% | 0.1%
1 | 19 | -0.112% | 0.042% | -0.175% | 0.1%
Deviations of %error on test id #3. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -0.115% | 0% | 0% | 0.1%
3 | 1 | 0.464% | 0.579% | 0% | 0.1%
3 | 2 | 0.321% | -0.143% | 0.223% | 0.1%
3 | 3 | 0.285% | -0.036% | 0.357% | 0.1%
3 | 4 | 0.348% | 0.063% | 0.318% | 0.1%
3 | 5 | 0.428% | 0.08% | 0.354% | 0.1%
Deviations of %error on test id #4. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -22.03% | 0% | 0% | 0.1%
4 | 1 | -6.932% | 15.09% | 0% | 0.1%
4 | 2 | -0.434% | 6.498% | -9.797% | 0.1%
4 | 3 | 0.062% | 0.496% | -2.435% | 0.1%
4 | 4 | 0.294% | 0.232% | -0.026% | 0.1%
4 | 5 | 0.527% | 0.233% | 0.294% | 0.1%
4 | 6 | 0.303% | -0.224% | 0.375% | 0.1%
4 | 7 | 0.107% | -0.196% | 0.312% | 0.1%
4 | 8 | 0.044% | -0.063% | 0.151% | 0.1%
4 | 9 | 0.053% | 0.009% | 0.068% | 0.1%
4 | 10 | 0.044% | -0.009% | 0.047% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.004% | 0% | 0% | 0.1%
7 | 1 | 0.082% | 0.078% | 0% | 0.1%
7 | 2 | 0.071% | -0.011% | 0.052% | 0.1%
7 | 3 | 0.075% | 0.004% | 0.076% | 0.1%

## 3. PF = 0.8C
[<<]
Accuracy Test Report of 'PF = 0.8C'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / PF = 0.8C'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:07:56 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | -0.071% | PASSED | 2023-07-27 10:51
Yes | 1 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | -0.083% | PASSED | 2023-07-27 10:51
Yes | 3 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | -0.012% | PASSED | 2023-07-27 10:51
Yes | 4 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | -0.166% | PASSED | 2023-07-27 10:52
Yes | 7 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.017% | PASSED | 2023-07-27 10:53
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -8.103% | 0% | 0% | 0.1%
0 | 1 | -9.913% | -1.81% | 0% | 0.1%
0 | 2 | -11.04% | -1.126% | -9.685% | 0.1%
0 | 3 | -6.484% | 4.555% | -9.145% | 0.1%
0 | 4 | -1.781% | 4.703% | -6.435% | 0.1%
0 | 5 | -1.815% | -0.034% | -3.36% | 0.1%
0 | 6 | -0.548% | 1.267% | -1.381% | 0.1%
0 | 7 | -0.142% | 0.406% | -0.835% | 0.1%
0 | 8 | -0.124% | 0.018% | -0.271% | 0.1%
0 | 9 | -0.089% | 0.035% | -0.118% | 0.1%
0 | 10 | 0% | 0.089% | -0.071% | 0.1%
Deviations of %error on test id #1. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -11.8% | 0% | 0% | 0.1%
1 | 1 | -6.39% | 5.406% | 0% | 0.1%
1 | 2 | -5.642% | 0.748% | -7.943% | 0.1%
1 | 3 | -1.437% | 4.205% | -4.49% | 0.1%
1 | 4 | -0.98% | 0.457% | -2.686% | 0.1%
1 | 5 | 0% | 0.98% | -0.806% | 0.1%
1 | 6 | 0.006% | 0.006% | -0.325% | 0.1%
1 | 7 | -0.124% | -0.13% | -0.039% | 0.1%
1 | 8 | -0.089% | 0.035% | -0.069% | 0.1%
1 | 9 | -0.106% | -0.017% | -0.106% | 0.1%
1 | 10 | -0.053% | 0.053% | -0.083% | 0.1%
Deviations of %error on test id #3. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -7.514% | 0% | 0% | 0.1%
3 | 1 | -7.506% | 0.008% | 0% | 0.1%
3 | 2 | -3.747% | 3.759% | -6.256% | 0.1%
3 | 3 | -2.631% | 1.116% | -4.628% | 0.1%
3 | 4 | -2.639% | -0.008% | -3.006% | 0.1%
3 | 5 | -0.044% | 2.595% | -1.771% | 0.1%
3 | 6 | 0% | 0.044% | -0.894% | 0.1%
3 | 7 | 0.018% | 0.018% | -0.009% | 0.1%
3 | 8 | -0.053% | -0.071% | -0.012% | 0.1%
Deviations of %error on test id #4. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -6.46% | 0% | 0% | 0.1%
4 | 1 | -3.574% | 2.886% | 0% | 0.1%
4 | 2 | -0.539% | 3.035% | -3.524% | 0.1%
4 | 3 | -0.275% | 0.264% | -1.463% | 0.1%
4 | 4 | -0.195% | 0.08% | -0.336% | 0.1%
4 | 5 | -0.142% | 0.053% | -0.204% | 0.1%
4 | 6 | -0.16% | -0.018% | -0.166% | 0.1%
Deviations of %error on test id #7. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | -1.229% | 0% | 0% | 0.1%
7 | 1 | -0.114% | 1.115% | 0% | 0.1%
7 | 2 | 0.06% | 0.174% | -0.428% | 0.1%
7 | 3 | 0.016% | -0.044% | -0.013% | 0.1%
7 | 4 | 0.023% | 0.007% | 0.033% | 0.1%
7 | 5 | 0.011% | -0.012% | 0.017% | 0.1%

## 4. 5th Harmonics
[<<]
Accuracy Test Report of '5th Harmonics'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / 5th Harmonics'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:07:59 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Harmonics' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
1 | 0 | 0 | 1
3. Test Cases
Total 1 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 1 | DP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° | {UcUbUa, 5th-10%, IcIbIa, 5th-40%} | 12 | 20 |  | ±0.1% | ±1% | -0.027% | PASSED | 2023-07-27 10:54
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #1. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -1.119% | 0% | 0% | 0.1%
1 | 1 | -6.112% | -4.993% | 0% | 0.1%
1 | 2 | -4.771% | 1.341% | -4.001% | 0.1%
1 | 3 | -2.406% | 2.365% | -4.43% | 0.1%
1 | 4 | 0.006% | 2.412% | -2.39% | 0.1%
1 | 5 | 0.399% | 0.393% | -0.667% | 0.1%
1 | 6 | 0.461% | 0.062% | 0.289% | 0.1%
1 | 7 | 0.273% | -0.188% | 0.378% | 0.1%
1 | 8 | 0.047% | -0.226% | 0.26% | 0.1%
1 | 9 | -0.009% | -0.056% | 0.104% | 0.1%
1 | 10 | -0.036% | -0.027% | 0.001% | 0.1%
1 | 11 | -0.036% | 0% | -0.027% | 0.1%

## 5. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:08:00 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DP | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | -0.003% | PASSED | 2023-07-27 10:54
Yes | 1 | DP | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.03% | PASSED | 2023-07-27 10:55
Yes | 2 | DP | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.028% | PASSED | 2023-07-27 10:55
Yes | 3 | DP | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.041% | PASSED | 2023-07-27 10:56
Yes | 4 | DP | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.025% | PASSED | 2023-07-27 10:56
Yes | 5 | DP | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | 0.015% | PASSED | 2023-07-27 10:57
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 2.44% | 0% | 0% | 0.1%
0 | 1 | 0.956% | -1.484% | 0% | 0.1%
0 | 2 | 0.027% | -0.929% | 1.141% | 0.1%
0 | 3 | -0.009% | -0.036% | 0.325% | 0.1%
0 | 4 | 0% | 0.009% | 0.006% | 0.1%
0 | 5 | 0% | 0% | -0.003% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0.018% | 0% | 0% | 0.1%
1 | 1 | 0.049% | 0.031% | 0% | 0.1%
1 | 2 | 0.022% | -0.027% | 0.03% | 0.1%
1 | 3 | 0.018% | -0.004% | 0.03% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | 0.031% | 0% | 0% | 0.1%
2 | 1 | 0.022% | -0.009% | 0% | 0.1%
2 | 2 | 0.036% | 0.014% | 0.03% | 0.1%
2 | 3 | 0.027% | -0.009% | 0.028% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -0.009% | 0% | 0% | 0.1%
3 | 1 | 0.044% | 0.053% | 0% | 0.1%
3 | 2 | 0.036% | -0.008% | 0.024% | 0.1%
3 | 3 | 0.044% | 0.008% | 0.041% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.022% | 0% | 0% | 0.1%
4 | 1 | 0.027% | 0.005% | 0% | 0.1%
4 | 2 | 0.062% | 0.035% | 0.037% | 0.1%
4 | 3 | -0.013% | -0.075% | 0.025% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -0.004% | 0% | 0% | 0.1%
5 | 1 | 0.049% | 0.053% | 0% | 0.1%
5 | 2 | -0.009% | -0.058% | 0.012% | 0.1%
5 | 3 | 0.004% | 0.013% | 0.015% | 0.1%

## 6. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:08:03 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DP | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | 0.021% | PASSED | 2023-07-27 10:57
Yes | 1 | DP | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.021% | PASSED | 2023-07-27 10:58
Yes | 2 | DP | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.116% | PASSED | 2023-07-27 10:59
Yes | 3 | DP | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.027% | PASSED | 2023-07-27 11:00
Yes | 4 | DP | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.834% | PASSED | 2023-07-27 11:00
Yes | 5 | DP | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | 0.038% | PASSED | 2023-07-27 11:01
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -16.55% | 0% | 0% | 0.1%
0 | 1 | -9.957% | 6.592% | 0% | 0.1%
0 | 2 | -0.933% | 9.024% | -9.146% | 0.1%
0 | 3 | 0.241% | 1.174% | -3.55% | 0.1%
0 | 4 | 0.16% | -0.081% | -0.177% | 0.1%
0 | 5 | 0.33% | 0.17% | 0.244% | 0.1%
0 | 6 | 0.446% | 0.116% | 0.312% | 0.1%
0 | 7 | 0.581% | 0.135% | 0.452% | 0.1%
0 | 8 | 0.348% | -0.233% | 0.458% | 0.1%
0 | 9 | 0.151% | -0.197% | 0.36% | 0.1%
0 | 10 | 0.089% | -0.062% | 0.196% | 0.1%
0 | 11 | 0.027% | -0.062% | 0.089% | 0.1%
0 | 12 | 0.009% | -0.018% | 0.042% | 0.1%
0 | 13 | 0.027% | 0.018% | 0.021% | 0.1%
Deviations of %error on test id #1. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0.134% | 0% | 0% | 0.1%
1 | 1 | 0.348% | 0.214% | 0% | 0.1%
1 | 2 | 0.232% | -0.116% | 0.238% | 0.1%
1 | 3 | 0.196% | -0.036% | 0.259% | 0.1%
1 | 4 | 0.196% | 0% | 0.208% | 0.1%
1 | 5 | 0.044% | -0.152% | 0.145% | 0.1%
1 | 6 | 0.036% | -0.008% | 0.092% | 0.1%
1 | 7 | 0.027% | -0.009% | 0.036% | 0.1%
1 | 8 | 0% | -0.027% | 0.021% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | 0.25% | 0% | 0% | 0.1%
2 | 1 | 0.196% | -0.054% | 0% | 0.1%
2 | 2 | 0.125% | -0.071% | 0.19% | 0.1%
2 | 3 | 0.027% | -0.098% | 0.116% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.071% | 0% | 0% | 0.1%
3 | 1 | 0.027% | -0.044% | 0% | 0.1%
3 | 2 | 0.027% | 0% | 0.042% | 0.1%
3 | 3 | 0.027% | 0% | 0.027% | 0.1%
Deviations of %error on test id #4. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -22.52% | 0% | 0% | 0.1%
4 | 1 | -18.25% | 4.268% | 0% | 0.1%
4 | 2 | -8.863% | 9.384% | -16.54% | 0.1%
4 | 3 | -6.289% | 2.574% | -11.13% | 0.1%
4 | 4 | -0.407% | 5.882% | -5.186% | 0.1%
4 | 5 | 0.375% | 0.782% | -2.107% | 0.1%
4 | 6 | 0.241% | -0.134% | 0.07% | 0.1%
4 | 7 | 0.276% | 0.035% | 0.297% | 0.1%
4 | 8 | 0.125% | -0.151% | 0.214% | 0.1%
4 | 9 | 0.151% | 0.026% | 0.184% | 0.1%
4 | 10 | 0.205% | 0.054% | 0.16% | 0.1%
4 | 11 | 0.393% | 0.188% | 0.25% | 0.1%
4 | 12 | 0.223% | -0.17% | 0.274% | 0.1%
4 | 13 | 0.232% | 0.009% | 0.283% | 0.1%
4 | 14 | 0.107% | -0.125% | 0.187% | 0.1%
4 | 15 | 0.393% | 0.286% | 0.244% | 0.1%
4 | 16 | 1.151% | 0.758% | 0.55% | 0.1%
4 | 17 | 1.042% | -0.109% | 0.862% | 0.1%
4 | 18 | 0.888% | -0.154% | 1.027% | 0.1%
4 | 19 | 0.572% | -0.316% | 0.834% | 0.1%
Deviations of %error on test id #5. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -9.172% | 0% | 0% | 0.1%
5 | 1 | 0.25% | 9.422% | 0% | 0.1%
5 | 2 | 0.42% | 0.17% | -2.834% | 0.1%
5 | 3 | 0.258% | -0.162% | 0.309% | 0.1%
5 | 4 | 0.25% | -0.008% | 0.309% | 0.1%
5 | 5 | 0.125% | -0.125% | 0.211% | 0.1%
5 | 6 | 0.125% | 0% | 0.167% | 0.1%
5 | 7 | -0.036% | -0.161% | 0.071% | 0.1%
5 | 8 | 0.321% | 0.357% | 0.137% | 0.1%
5 | 9 | 0.258% | -0.063% | 0.181% | 0.1%
5 | 10 | 0.258% | 0% | 0.279% | 0.1%
5 | 11 | 0.107% | -0.151% | 0.208% | 0.1%
5 | 12 | 0.142% | 0.035% | 0.169% | 0.1%
5 | 13 | 0.044% | -0.098% | 0.098% | 0.1%
5 | 14 | 0.044% | 0% | 0.077% | 0.1%
5 | 15 | 0.027% | -0.017% | 0.038% | 0.1%

## 7. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:08:07 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.042% | PASSED | 2023-07-27 11:01
Yes | 1 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.027% | PASSED | 2023-07-27 11:02
Yes | 2 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.021% | PASSED | 2023-07-27 11:02
Yes | 3 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.018% | PASSED | 2023-07-27 11:02
Yes | 4 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.009% | PASSED | 2023-07-27 11:03
Yes | 5 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.006% | PASSED | 2023-07-27 11:03
Yes | 6 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.003% | PASSED | 2023-07-27 11:04
Yes | 7 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.018% | PASSED | 2023-07-27 11:04
Yes | 8 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.016% | PASSED | 2023-07-27 11:05
Yes | 9 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.03% | PASSED | 2023-07-27 11:05
Yes | 10 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.033% | PASSED | 2023-07-27 11:05
Yes | 11 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.028% | PASSED | 2023-07-27 11:06
Yes | 12 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.04% | PASSED | 2023-07-27 11:06
Yes | 13 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.049% | PASSED | 2023-07-27 11:07
Yes | 14 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.052% | PASSED | 2023-07-27 11:07
Yes | 15 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.055% | PASSED | 2023-07-27 11:08
Yes | 16 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.045% | PASSED | 2023-07-27 11:08
Yes | 17 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.05% | PASSED | 2023-07-27 11:09
Yes | 18 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.059% | PASSED | 2023-07-27 11:10
Yes | 19 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.053% | PASSED | 2023-07-27 11:10
Yes | 20 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.067% | PASSED | 2023-07-27 11:11
Yes | 21 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.045% | PASSED | 2023-07-27 11:11
Yes | 22 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.064% | PASSED | 2023-07-27 11:12
Yes | 23 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.07% | PASSED | 2023-07-27 11:13
Yes | 24 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.089% | PASSED | 2023-07-27 11:14
Yes | 25 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.059% | PASSED | 2023-07-27 11:15
Yes | 26 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.067% | PASSED | 2023-07-27 11:16
Yes | 27 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.082% | PASSED | 2023-07-27 11:17
Yes | 28 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.076% | PASSED | 2023-07-27 11:17
Yes | 29 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.07% | PASSED | 2023-07-27 11:19
Yes | 30 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.065% | PASSED | 2023-07-27 11:20
Yes | 31 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.082% | PASSED | 2023-07-27 11:49
Yes | 32 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.079% | PASSED | 2023-07-27 11:49
Yes | 33 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.074% | PASSED | 2023-07-27 11:22
Yes | 34 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.089% | PASSED | 2023-07-27 11:50
Yes | 35 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.081% | PASSED | 2023-07-27 11:24
Yes | 36 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.095% | PASSED | 2023-07-27 11:50
Yes | 37 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.092% | PASSED | 2023-07-27 11:51
Yes | 38 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.081% | PASSED | 2023-07-27 11:51
Yes | 39 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.081% | PASSED | 2023-07-27 11:51
Yes | 40 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.078% | PASSED | 2023-07-27 11:52
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -0.071% | 0% | 0% | 0.1%
0 | 1 | -0.018% | 0.053% | 0% | 0.1%
0 | 2 | -0.049% | -0.031% | -0.046% | 0.1%
0 | 3 | -0.058% | -0.009% | -0.042% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -0.009% | 0% | 0% | 0.1%
1 | 1 | -0.018% | -0.009% | 0% | 0.1%
1 | 2 | -0.004% | 0.014% | -0.01% | 0.1%
1 | 3 | -0.058% | -0.054% | -0.027% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | 0.009% | 0% | 0% | 0.1%
2 | 1 | -0.053% | -0.062% | 0% | 0.1%
2 | 2 | 0.004% | 0.057% | -0.013% | 0.1%
2 | 3 | -0.013% | -0.017% | -0.021% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -0.013% | 0% | 0% | 0.1%
3 | 1 | -0.009% | 0.004% | 0% | 0.1%
3 | 2 | -0.031% | -0.022% | -0.018% | 0.1%
3 | 3 | -0.013% | 0.018% | -0.018% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.009% | 0% | 0% | 0.1%
4 | 1 | -0.004% | -0.013% | 0% | 0.1%
4 | 2 | -0.013% | -0.009% | -0.003% | 0.1%
4 | 3 | -0.009% | 0.004% | -0.009% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -0.013% | 0% | 0% | 0.1%
5 | 1 | 0.013% | 0.026% | 0% | 0.1%
5 | 2 | -0.009% | -0.022% | -0.003% | 0.1%
5 | 3 | 0.013% | 0.022% | 0.006% | 0.1%
Deviations of %error on test id #6. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | 0% | 0% | 0% | 0.1%
6 | 1 | 0.027% | 0.027% | 0% | 0.1%
6 | 2 | 0.031% | 0.004% | 0.019% | 0.1%
6 | 3 | -0.027% | -0.058% | 0.01% | 0.1%
6 | 4 | -0.013% | 0.014% | -0.003% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | -0.018% | 0% | 0% | 0.1%
7 | 1 | 0.027% | 0.045% | 0% | 0.1%
7 | 2 | 0.044% | 0.017% | 0.018% | 0.1%
7 | 3 | -0.018% | -0.062% | 0.018% | 0.1%
Deviations of %error on test id #8. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | 0.031% | 0% | 0% | 0.1%
8 | 1 | 0.031% | 0% | 0% | 0.1%
8 | 2 | 0.022% | -0.009% | 0.028% | 0.1%
8 | 3 | -0.004% | -0.026% | 0.016% | 0.1%
Deviations of %error on test id #9. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | 0.036% | 0% | 0% | 0.1%
9 | 1 | -0.013% | -0.049% | 0% | 0.1%
9 | 2 | 0.076% | 0.089% | 0.033% | 0.1%
9 | 3 | 0.027% | -0.049% | 0.03% | 0.1%
Deviations of %error on test id #10. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | 0.018% | 0% | 0% | 0.1%
10 | 1 | 0.027% | 0.009% | 0% | 0.1%
10 | 2 | 0.04% | 0.013% | 0.028% | 0.1%
10 | 3 | 0.031% | -0.009% | 0.033% | 0.1%
Deviations of %error on test id #11. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | 0.031% | 0% | 0% | 0.1%
11 | 1 | 0.062% | 0.031% | 0% | 0.1%
11 | 2 | 0.004% | -0.058% | 0.032% | 0.1%
11 | 3 | 0.018% | 0.014% | 0.028% | 0.1%
Deviations of %error on test id #12. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | 0.067% | 0% | 0% | 0.1%
12 | 1 | 0.031% | -0.036% | 0% | 0.1%
12 | 2 | 0.022% | -0.009% | 0.04% | 0.1%
12 | 3 | 0.067% | 0.045% | 0.04% | 0.1%
Deviations of %error on test id #13. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | 0.027% | 0% | 0% | 0.1%
13 | 1 | 0.049% | 0.022% | 0% | 0.1%
13 | 2 | 0.067% | 0.018% | 0.048% | 0.1%
13 | 3 | 0.031% | -0.036% | 0.049% | 0.1%
Deviations of %error on test id #14. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | 0.036% | 0% | 0% | 0.1%
14 | 1 | 0.076% | 0.04% | 0% | 0.1%
14 | 2 | 0.044% | -0.032% | 0.052% | 0.1%
14 | 3 | 0.036% | -0.008% | 0.052% | 0.1%
Deviations of %error on test id #15. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | 0.076% | 0% | 0% | 0.1%
15 | 1 | 0.062% | -0.014% | 0% | 0.1%
15 | 2 | 0.036% | -0.026% | 0.058% | 0.1%
15 | 3 | 0.067% | 0.031% | 0.055% | 0.1%
Deviations of %error on test id #16. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | 0.036% | 0% | 0% | 0.1%
16 | 1 | 0.067% | 0.031% | 0% | 0.1%
16 | 2 | 0.027% | -0.04% | 0.043% | 0.1%
16 | 3 | 0.04% | 0.013% | 0.045% | 0.1%
Deviations of %error on test id #17. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | 0.071% | 0% | 0% | 0.1%
17 | 1 | 0.067% | -0.004% | 0% | 0.1%
17 | 2 | 0.053% | -0.014% | 0.064% | 0.1%
17 | 3 | 0.031% | -0.022% | 0.05% | 0.1%
Deviations of %error on test id #18. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | 0.031% | 0% | 0% | 0.1%
18 | 1 | 0.076% | 0.045% | 0% | 0.1%
18 | 2 | 0.08% | 0.004% | 0.062% | 0.1%
18 | 3 | 0.022% | -0.058% | 0.059% | 0.1%
Deviations of %error on test id #19. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | 0.062% | 0% | 0% | 0.1%
19 | 1 | 0.058% | -0.004% | 0% | 0.1%
19 | 2 | 0.031% | -0.027% | 0.05% | 0.1%
19 | 3 | 0.071% | 0.04% | 0.053% | 0.1%
Deviations of %error on test id #20. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | 0.071% | 0% | 0% | 0.1%
20 | 1 | 0.067% | -0.004% | 0% | 0.1%
20 | 2 | 0.058% | -0.009% | 0.065% | 0.1%
20 | 3 | 0.076% | 0.018% | 0.067% | 0.1%
Deviations of %error on test id #21. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | 0.058% | 0% | 0% | 0.1%
21 | 1 | 0.058% | 0% | 0% | 0.1%
21 | 2 | 0.009% | -0.049% | 0.042% | 0.1%
21 | 3 | 0.067% | 0.058% | 0.045% | 0.1%
Deviations of %error on test id #22. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | 0.067% | 0% | 0% | 0.1%
22 | 1 | 0.071% | 0.004% | 0% | 0.1%
22 | 2 | 0.053% | -0.018% | 0.064% | 0.1%
22 | 3 | 0.067% | 0.014% | 0.064% | 0.1%
Deviations of %error on test id #23. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | 0.053% | 0% | 0% | 0.1%
23 | 1 | 0.098% | 0.045% | 0% | 0.1%
23 | 2 | 0.044% | -0.054% | 0.065% | 0.1%
23 | 3 | 0.067% | 0.023% | 0.07% | 0.1%
Deviations of %error on test id #24. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | 0.076% | 0% | 0% | 0.1%
24 | 1 | 0.062% | -0.014% | 0% | 0.1%
24 | 2 | 0.125% | 0.063% | 0.088% | 0.1%
24 | 3 | 0.08% | -0.045% | 0.089% | 0.1%
Deviations of %error on test id #25. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | 0.071% | 0% | 0% | 0.1%
25 | 1 | 0.058% | -0.013% | 0% | 0.1%
25 | 2 | 0.053% | -0.005% | 0.061% | 0.1%
25 | 3 | 0.067% | 0.014% | 0.059% | 0.1%
Deviations of %error on test id #26. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | 0.067% | 0% | 0% | 0.1%
26 | 1 | 0.027% | -0.04% | 0% | 0.1%
26 | 2 | 0.098% | 0.071% | 0.064% | 0.1%
26 | 3 | 0.076% | -0.022% | 0.067% | 0.1%
Deviations of %error on test id #27. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | 0.058% | 0% | 0% | 0.1%
27 | 1 | 0.107% | 0.049% | 0% | 0.1%
27 | 2 | 0.076% | -0.031% | 0.08% | 0.1%
27 | 3 | 0.062% | -0.014% | 0.082% | 0.1%
Deviations of %error on test id #28. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | 0.071% | 0% | 0% | 0.1%
28 | 1 | 0.053% | -0.018% | 0% | 0.1%
28 | 2 | 0.067% | 0.014% | 0.064% | 0.1%
28 | 3 | 0.107% | 0.04% | 0.076% | 0.1%
Deviations of %error on test id #29. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | 0.093% | 0% | 0% | 0.1%
29 | 1 | 0.062% | -0.031% | 0% | 0.1%
29 | 2 | 0.062% | 0% | 0.072% | 0.1%
29 | 3 | 0.085% | 0.023% | 0.07% | 0.1%
29 | 4 | 0.062% | -0.023% | 0.07% | 0.1%
Deviations of %error on test id #30. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | 0.102% | 0% | 0% | 0.1%
30 | 1 | 0.071% | -0.031% | 0% | 0.1%
30 | 2 | 0.067% | -0.004% | 0.08% | 0.1%
30 | 3 | 0.058% | -0.009% | 0.065% | 0.1%
30 | 4 | 0.071% | 0.013% | 0.065% | 0.1%
Deviations of %error on test id #31. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | 0.102% | 0% | 0% | 0.1%
31 | 1 | 0.085% | -0.017% | 0% | 0.1%
31 | 2 | 0.102% | 0.017% | 0.096% | 0.1%
31 | 3 | 0.058% | -0.044% | 0.082% | 0.1%
Deviations of %error on test id #32. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | 0.04% | 0% | 0% | 0.1%
32 | 1 | 0.098% | 0.058% | 0% | 0.1%
32 | 2 | 0.058% | -0.04% | 0.065% | 0.1%
32 | 3 | 0.08% | 0.022% | 0.079% | 0.1%
Deviations of %error on test id #33. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | 0.062% | 0% | 0% | 0.1%
33 | 1 | 0.058% | -0.004% | 0% | 0.1%
33 | 2 | 0.058% | 0% | 0.059% | 0.1%
33 | 3 | 0.107% | 0.049% | 0.074% | 0.1%
Deviations of %error on test id #34. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | 0.098% | 0% | 0% | 0.1%
34 | 1 | 0.067% | -0.031% | 0% | 0.1%
34 | 2 | 0.107% | 0.04% | 0.091% | 0.1%
34 | 3 | 0.093% | -0.014% | 0.089% | 0.1%
Deviations of %error on test id #35. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | 0.067% | 0% | 0% | 0.1%
35 | 1 | 0.053% | -0.014% | 0% | 0.1%
35 | 2 | 0.102% | 0.049% | 0.074% | 0.1%
35 | 3 | 0.089% | -0.013% | 0.081% | 0.1%
Deviations of %error on test id #36. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | 0.076% | 0% | 0% | 0.1%
36 | 1 | 0.116% | 0.04% | 0% | 0.1%
36 | 2 | 0.098% | -0.018% | 0.097% | 0.1%
36 | 3 | 0.071% | -0.027% | 0.095% | 0.1%
Deviations of %error on test id #37. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | 0.102% | 0% | 0% | 0.1%
37 | 1 | 0.107% | 0.005% | 0% | 0.1%
37 | 2 | 0.098% | -0.009% | 0.102% | 0.1%
37 | 3 | 0.071% | -0.027% | 0.092% | 0.1%
Deviations of %error on test id #38. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | 0.102% | 0% | 0% | 0.1%
38 | 1 | 0.102% | 0% | 0% | 0.1%
38 | 2 | 0.089% | -0.013% | 0.098% | 0.1%
38 | 3 | 0.053% | -0.036% | 0.081% | 0.1%
Deviations of %error on test id #39. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | 0.093% | 0% | 0% | 0.1%
39 | 1 | 0.062% | -0.031% | 0% | 0.1%
39 | 2 | 0.093% | 0.031% | 0.083% | 0.1%
39 | 3 | 0.089% | -0.004% | 0.081% | 0.1%
Deviations of %error on test id #40. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | 0.049% | 0% | 0% | 0.1%
40 | 1 | 0.098% | 0.049% | 0% | 0.1%
40 | 2 | 0.093% | -0.005% | 0.08% | 0.1%
40 | 3 | 0.044% | -0.049% | 0.078% | 0.1%

## 8. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:08:24 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.193% | PASSED | 2023-07-27 11:52
Yes | 1 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.041% | PASSED | 2023-07-27 11:53
Yes | 2 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.089% | PASSED | 2023-07-27 11:54
Yes | 3 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.047% | PASSED | 2023-07-27 11:54
Yes | 4 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.375% | PASSED | 2023-07-27 11:55
Yes | 5 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.632% | PASSED | 2023-07-27 11:55
Yes | 6 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.241% | PASSED | 2023-07-27 11:56
Yes | 7 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.426% | PASSED | 2023-07-27 11:56
Yes | 8 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.255% | PASSED | 2023-07-27 11:57
Yes | 9 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.131% | PASSED | 2023-07-27 11:57
Yes | 10 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.273% | PASSED | 2023-07-27 11:58
Yes | 11 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.149% | PASSED | 2023-07-27 11:58
Yes | 12 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.027% | PASSED | 2023-07-27 11:59
Yes | 13 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.285% | PASSED | 2023-07-27 12:00
Yes | 14 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.193% | PASSED | 2023-07-27 12:00
Yes | 15 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.009% | PASSED | 2023-07-27 12:01
Yes | 16 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.047% | PASSED | 2023-07-27 15:35
Yes | 17 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.033% | PASSED | 2023-07-27 12:02
Yes | 18 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.006% | PASSED | 2023-07-27 12:03
Yes | 19 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.027% | PASSED | 2023-07-27 12:03
Yes | 20 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.012% | PASSED | 2023-07-27 15:35
Yes | 21 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.003% | PASSED | 2023-07-27 12:05
Yes | 22 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.003% | PASSED | 2023-07-27 15:36
Yes | 23 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.015% | PASSED | 2023-07-27 12:06
Yes | 24 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.05% | PASSED | 2023-07-27 15:36
Yes | 25 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.062% | PASSED | 2023-07-27 12:08
Yes | 26 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.038% | PASSED | 2023-07-27 12:09
Yes | 27 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.033% | PASSED | 2023-07-27 12:36
Yes | 28 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.053% | PASSED | 2023-07-27 12:11
Yes | 29 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.059% | PASSED | 2023-07-27 12:13
Yes | 30 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.045% | PASSED | 2023-07-27 12:37
Yes | 31 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.136% | PASSED | 2023-07-27 12:37
Yes | 32 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.062% | PASSED | 2023-07-27 12:38
Yes | 33 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.255% | PASSED | 2023-07-27 12:38
Yes | 34 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.083% | PASSED | 2023-07-27 12:39
Yes | 35 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.083% | PASSED | 2023-07-27 12:40
Yes | 36 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.097% | PASSED | 2023-07-27 12:40
Yes | 37 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.116% | PASSED | 2023-07-27 12:41
Yes | 38 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.083% | PASSED | 2023-07-27 12:41
Yes | 39 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.051% | PASSED | 2023-07-27 12:42
Yes | 40 | DP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.097% | PASSED | 2023-07-27 12:43
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -25.44% | 0% | 0% | 0.1%
0 | 1 | -22.61% | 2.831% | 0% | 0.1%
0 | 2 | -9.95% | 12.66% | -19.33% | 0.1%
0 | 3 | -6.6% | 3.35% | -13.05% | 0.1%
0 | 4 | 0.276% | 6.876% | -5.425% | 0.1%
0 | 5 | 0.276% | 0% | -2.016% | 0.1%
0 | 6 | 0.178% | -0.098% | 0.243% | 0.1%
0 | 7 | 0.125% | -0.053% | 0.193% | 0.1%
Deviations of %error on test id #1. Total 19 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -29.59% | 0% | 0% | 0.1%
1 | 1 | -24.24% | 5.348% | 0% | 0.1%
1 | 2 | -12.72% | 11.53% | -22.18% | 0.1%
1 | 3 | -8.141% | 4.575% | -15.03% | 0.1%
1 | 4 | -0.548% | 7.593% | -7.135% | 0.1%
1 | 5 | 0.339% | 0.887% | -2.783% | 0.1%
1 | 6 | 0.178% | -0.161% | -0.01% | 0.1%
1 | 7 | 0.178% | 0% | 0.232% | 0.1%
1 | 8 | 0.196% | 0.018% | 0.184% | 0.1%
1 | 9 | 0.357% | 0.161% | 0.244% | 0.1%
1 | 10 | 0.437% | 0.08% | 0.33% | 0.1%
1 | 11 | 0.554% | 0.117% | 0.449% | 0.1%
1 | 12 | 0.393% | -0.161% | 0.461% | 0.1%
1 | 13 | 0.25% | -0.143% | 0.399% | 0.1%
1 | 14 | 0.169% | -0.081% | 0.271% | 0.1%
1 | 15 | 0.062% | -0.107% | 0.16% | 0.1%
1 | 16 | 0.053% | -0.009% | 0.095% | 0.1%
1 | 17 | 0.027% | -0.026% | 0.047% | 0.1%
1 | 18 | 0.044% | 0.017% | 0.041% | 0.1%
Deviations of %error on test id #2. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -25.32% | 0% | 0% | 0.1%
2 | 1 | -13.84% | 11.48% | 0% | 0.1%
2 | 2 | -10.02% | 3.824% | -16.4% | 0.1%
2 | 3 | -0.609% | 9.412% | -8.158% | 0.1%
2 | 4 | 0.062% | 0.671% | -3.523% | 0.1%
2 | 5 | 0.294% | 0.232% | -0.084% | 0.1%
2 | 6 | 0.42% | 0.126% | 0.259% | 0.1%
2 | 7 | 0.491% | 0.071% | 0.402% | 0.1%
2 | 8 | 0.554% | 0.063% | 0.488% | 0.1%
2 | 9 | 0.671% | 0.117% | 0.572% | 0.1%
2 | 10 | 0.608% | -0.063% | 0.611% | 0.1%
2 | 11 | 0.375% | -0.233% | 0.551% | 0.1%
2 | 12 | 0.196% | -0.179% | 0.393% | 0.1%
2 | 13 | 0.151% | -0.045% | 0.241% | 0.1%
2 | 14 | 0.071% | -0.08% | 0.139% | 0.1%
2 | 15 | 0.044% | -0.027% | 0.089% | 0.1%
Deviations of %error on test id #3. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -10.8% | 0% | 0% | 0.1%
3 | 1 | -6.932% | 3.867% | 0% | 0.1%
3 | 2 | -0.451% | 6.481% | -6.061% | 0.1%
3 | 3 | 0.223% | 0.674% | -2.387% | 0.1%
3 | 4 | 0.276% | 0.053% | 0.016% | 0.1%
3 | 5 | 0.402% | 0.126% | 0.3% | 0.1%
3 | 6 | 0.258% | -0.144% | 0.312% | 0.1%
3 | 7 | 0.285% | 0.027% | 0.315% | 0.1%
3 | 8 | 0.437% | 0.152% | 0.327% | 0.1%
3 | 9 | 0.402% | -0.035% | 0.375% | 0.1%
3 | 10 | 0.232% | -0.17% | 0.357% | 0.1%
3 | 11 | 0.178% | -0.054% | 0.271% | 0.1%
3 | 12 | 0.178% | 0% | 0.196% | 0.1%
3 | 13 | 0.044% | -0.134% | 0.133% | 0.1%
3 | 14 | 0.018% | -0.026% | 0.08% | 0.1%
3 | 15 | 0.062% | 0.044% | 0.041% | 0.1%
3 | 16 | 0.062% | 0% | 0.047% | 0.1%
Deviations of %error on test id #4. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -19.04% | 0% | 0% | 0.1%
4 | 1 | -12.75% | 6.291% | 0% | 0.1%
4 | 2 | -6.101% | 6.649% | -12.63% | 0.1%
4 | 3 | -1.471% | 4.63% | -6.774% | 0.1%
4 | 4 | 0.42% | 1.891% | -2.384% | 0.1%
4 | 5 | 0.375% | -0.045% | -0.225% | 0.1%
4 | 6 | 0.644% | 0.269% | 0.48% | 0.1%
4 | 7 | 0.402% | -0.242% | 0.474% | 0.1%
4 | 8 | 0.312% | -0.09% | 0.453% | 0.1%
4 | 9 | 0.411% | 0.099% | 0.375% | 0.1%
4 | 10 | 0.402% | -0.009% | 0.375% | 0.1%
Deviations of %error on test id #5. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -20.5% | 0% | 0% | 0.1%
5 | 1 | -16.26% | 4.243% | 0% | 0.1%
5 | 2 | -5.96% | 10.3% | -14.24% | 0.1%
5 | 3 | -4.328% | 1.632% | -8.848% | 0.1%
5 | 4 | 0.258% | 4.586% | -3.343% | 0.1%
5 | 5 | 0.411% | 0.153% | -1.22% | 0.1%
5 | 6 | 0.303% | -0.108% | 0.324% | 0.1%
5 | 7 | 0.258% | -0.045% | 0.324% | 0.1%
5 | 8 | 0.134% | -0.124% | 0.232% | 0.1%
5 | 9 | 0.089% | -0.045% | 0.16% | 0.1%
5 | 10 | 0.25% | 0.161% | 0.158% | 0.1%
5 | 11 | 0.285% | 0.035% | 0.208% | 0.1%
5 | 12 | 0.267% | -0.018% | 0.267% | 0.1%
5 | 13 | 0.16% | -0.107% | 0.237% | 0.1%
5 | 14 | 0.285% | 0.125% | 0.237% | 0.1%
5 | 15 | 0.978% | 0.693% | 0.474% | 0.1%
5 | 16 | 1.105% | 0.127% | 0.789% | 0.1%
5 | 17 | 0.879% | -0.226% | 0.987% | 0.1%
5 | 18 | 0.554% | -0.325% | 0.846% | 0.1%
5 | 19 | 0.464% | -0.09% | 0.632% | 0.1%
Deviations of %error on test id #6. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | -20.34% | 0% | 0% | 0.1%
6 | 1 | -8.417% | 11.92% | 0% | 0.1%
6 | 2 | -5.111% | 3.306% | -11.29% | 0.1%
6 | 3 | 0.348% | 5.459% | -4.393% | 0.1%
6 | 4 | 0.312% | -0.036% | -1.484% | 0.1%
6 | 5 | 0.241% | -0.071% | 0.3% | 0.1%
6 | 6 | 0.169% | -0.072% | 0.241% | 0.1%
Deviations of %error on test id #7. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | -21.14% | 0% | 0% | 0.1%
7 | 1 | -8.529% | 12.61% | 0% | 0.1%
7 | 2 | -5.295% | 3.234% | -11.65% | 0.1%
7 | 3 | 0.384% | 5.679% | -4.48% | 0.1%
7 | 4 | 0.455% | 0.071% | -1.485% | 0.1%
7 | 5 | 0.626% | 0.171% | 0.488% | 0.1%
7 | 6 | 0.482% | -0.144% | 0.521% | 0.1%
7 | 7 | 0.42% | -0.062% | 0.509% | 0.1%
7 | 8 | 0.384% | -0.036% | 0.429% | 0.1%
7 | 9 | 0.473% | 0.089% | 0.426% | 0.1%
Deviations of %error on test id #8. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | -7.491% | 0% | 0% | 0.1%
8 | 1 | 0.258% | 7.749% | 0% | 0.1%
8 | 2 | 0.303% | 0.045% | -2.31% | 0.1%
8 | 3 | 0.276% | -0.027% | 0.279% | 0.1%
8 | 4 | 0.187% | -0.089% | 0.255% | 0.1%
Deviations of %error on test id #9. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | -4.491% | 0% | 0% | 0.1%
9 | 1 | 0.411% | 4.902% | 0% | 0.1%
9 | 2 | 0.276% | -0.135% | -1.268% | 0.1%
9 | 3 | 0.187% | -0.089% | 0.291% | 0.1%
9 | 4 | 0.071% | -0.116% | 0.178% | 0.1%
9 | 5 | 0.027% | -0.044% | 0.095% | 0.1%
9 | 6 | 0.098% | 0.071% | 0.065% | 0.1%
9 | 7 | 0.33% | 0.232% | 0.152% | 0.1%
9 | 8 | 0.258% | -0.072% | 0.229% | 0.1%
9 | 9 | 0.223% | -0.035% | 0.27% | 0.1%
9 | 10 | 0.134% | -0.089% | 0.205% | 0.1%
9 | 11 | 0.036% | -0.098% | 0.131% | 0.1%
Deviations of %error on test id #10. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | -2.631% | 0% | 0% | 0.1%
10 | 1 | 0.027% | 2.658% | 0% | 0.1%
10 | 2 | 0.018% | -0.009% | -0.862% | 0.1%
10 | 3 | 0.303% | 0.285% | 0.116% | 0.1%
10 | 4 | 0.303% | 0% | 0.208% | 0.1%
10 | 5 | 0.214% | -0.089% | 0.273% | 0.1%
10 | 6 | 0.285% | 0.071% | 0.267% | 0.1%
10 | 7 | 0.321% | 0.036% | 0.273% | 0.1%
Deviations of %error on test id #11. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | -2.496% | 0% | 0% | 0.1%
11 | 1 | 0.116% | 2.612% | 0% | 0.1%
11 | 2 | 0.357% | 0.241% | -0.674% | 0.1%
11 | 3 | 0.214% | -0.143% | 0.229% | 0.1%
11 | 4 | 0.205% | -0.009% | 0.259% | 0.1%
11 | 5 | 0.107% | -0.098% | 0.175% | 0.1%
11 | 6 | 0.134% | 0.027% | 0.149% | 0.1%
Deviations of %error on test id #12. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | 0.25% | 0% | 0% | 0.1%
12 | 1 | 0.375% | 0.125% | 0% | 0.1%
12 | 2 | 0.294% | -0.081% | 0.306% | 0.1%
12 | 3 | 0.267% | -0.027% | 0.312% | 0.1%
12 | 4 | 0.482% | 0.215% | 0.348% | 0.1%
12 | 5 | 0.25% | -0.232% | 0.333% | 0.1%
12 | 6 | 0.232% | -0.018% | 0.321% | 0.1%
12 | 7 | 0.125% | -0.107% | 0.202% | 0.1%
12 | 8 | 0.044% | -0.081% | 0.134% | 0.1%
12 | 9 | 0.027% | -0.017% | 0.065% | 0.1%
12 | 10 | 0.009% | -0.018% | 0.027% | 0.1%
Deviations of %error on test id #13. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | 0.276% | 0% | 0% | 0.1%
13 | 1 | 0.312% | 0.036% | 0% | 0.1%
13 | 2 | 0.241% | -0.071% | 0.276% | 0.1%
13 | 3 | 0.303% | 0.062% | 0.285% | 0.1%
Deviations of %error on test id #14. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | 0.294% | 0% | 0% | 0.1%
14 | 1 | 0.276% | -0.018% | 0% | 0.1%
14 | 2 | 0.187% | -0.089% | 0.252% | 0.1%
14 | 3 | 0.116% | -0.071% | 0.193% | 0.1%
Deviations of %error on test id #15. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | 0.089% | 0% | 0% | 0.1%
15 | 1 | 0.018% | -0.071% | 0% | 0.1%
15 | 2 | -0.018% | -0.036% | 0.03% | 0.1%
15 | 3 | 0.027% | 0.045% | 0.009% | 0.1%
Deviations of %error on test id #16. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | -31.92% | 0% | 0% | 0.1%
16 | 1 | -15.52% | 16.4% | 0% | 0.1%
16 | 2 | -5.732% | 9.789% | -17.72% | 0.1%
16 | 3 | -0.133% | 5.599% | -7.129% | 0.1%
16 | 4 | 0.018% | 0.151% | -1.949% | 0.1%
16 | 5 | 0.009% | -0.009% | -0.035% | 0.1%
16 | 6 | 0.044% | 0.035% | 0.024% | 0.1%
16 | 7 | 0.089% | 0.045% | 0.047% | 0.1%
Deviations of %error on test id #17. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | -0.009% | 0% | 0% | 0.1%
17 | 1 | -0.018% | -0.009% | 0% | 0.1%
17 | 2 | 0.018% | 0.036% | -0.003% | 0.1%
17 | 3 | 0.098% | 0.08% | 0.033% | 0.1%
Deviations of %error on test id #18. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | 0% | 0% | 0% | 0.1%
18 | 1 | 0.036% | 0.036% | 0% | 0.1%
18 | 2 | 0.071% | 0.035% | 0.036% | 0.1%
18 | 3 | 0.509% | 0.438% | 0.205% | 0.1%
18 | 4 | 0.518% | 0.009% | 0.366% | 0.1%
18 | 5 | 0.464% | -0.054% | 0.497% | 0.1%
18 | 6 | 0.214% | -0.25% | 0.399% | 0.1%
18 | 7 | 0.134% | -0.08% | 0.271% | 0.1%
18 | 8 | 0.009% | -0.125% | 0.119% | 0.1%
18 | 9 | -0.027% | -0.036% | 0.039% | 0.1%
18 | 10 | 0.009% | 0.036% | -0.003% | 0.1%
18 | 11 | 0% | -0.009% | -0.006% | 0.1%
18 | 12 | -0.027% | -0.027% | -0.006% | 0.1%
Deviations of %error on test id #19. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | 0.036% | 0% | 0% | 0.1%
19 | 1 | 0% | -0.036% | 0% | 0.1%
19 | 2 | -0.027% | -0.027% | 0.003% | 0.1%
19 | 3 | -0.053% | -0.026% | -0.027% | 0.1%
Deviations of %error on test id #20. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | -32.38% | 0% | 0% | 0.1%
20 | 1 | -18.19% | 14.19% | 0% | 0.1%
20 | 2 | -5.295% | 12.9% | -18.62% | 0.1%
20 | 3 | -1.255% | 4.04% | -8.248% | 0.1%
20 | 4 | 0.044% | 1.299% | -2.169% | 0.1%
20 | 5 | 0.223% | 0.179% | -0.329% | 0.1%
20 | 6 | 0.33% | 0.107% | 0.199% | 0.1%
20 | 7 | 0.518% | 0.188% | 0.357% | 0.1%
20 | 8 | 0.617% | 0.099% | 0.488% | 0.1%
20 | 9 | 0.437% | -0.18% | 0.524% | 0.1%
20 | 10 | 0.178% | -0.259% | 0.411% | 0.1%
20 | 11 | 0.044% | -0.134% | 0.22% | 0.1%
20 | 12 | 0.009% | -0.035% | 0.077% | 0.1%
20 | 13 | 0.009% | 0% | 0.021% | 0.1%
20 | 14 | 0.018% | 0.009% | 0.012% | 0.1%
Deviations of %error on test id #21. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | 0.044% | 0% | 0% | 0.1%
21 | 1 | -0.062% | -0.106% | 0% | 0.1%
21 | 2 | 0.009% | 0.071% | -0.003% | 0.1%
21 | 3 | -0.018% | -0.027% | -0.024% | 0.1%
21 | 4 | 0% | 0.018% | -0.003% | 0.1%
Deviations of %error on test id #22. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | -29.15% | 0% | 0% | 0.1%
22 | 1 | -17.19% | 11.95% | 0% | 0.1%
22 | 2 | -6.328% | 10.87% | -17.56% | 0.1%
22 | 3 | 0% | 6.328% | -7.841% | 0.1%
22 | 4 | -0.018% | -0.018% | -2.115% | 0.1%
22 | 5 | 0.027% | 0.045% | 0.003% | 0.1%
22 | 6 | 0% | -0.027% | 0.003% | 0.1%
Deviations of %error on test id #23. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | 0.018% | 0% | 0% | 0.1%
23 | 1 | -0.036% | -0.054% | 0% | 0.1%
23 | 2 | -0.027% | 0.009% | -0.015% | 0.1%
23 | 3 | 0.009% | 0.036% | -0.018% | 0.1%
23 | 4 | -0.027% | -0.036% | -0.015% | 0.1%
Deviations of %error on test id #24. Total 21 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | -32.83% | 0% | 0% | 0.1%
24 | 1 | -21.37% | 11.46% | 0% | 0.1%
24 | 2 | -9.311% | 12.06% | -21.17% | 0.1%
24 | 3 | -0.266% | 9.045% | -10.32% | 0.1%
24 | 4 | -0.036% | 0.23% | -3.204% | 0.1%
24 | 5 | 0.285% | 0.321% | -0.006% | 0.1%
24 | 6 | 0.187% | -0.098% | 0.145% | 0.1%
24 | 7 | 0.125% | -0.062% | 0.199% | 0.1%
24 | 8 | 0.276% | 0.151% | 0.196% | 0.1%
24 | 9 | 0.267% | -0.009% | 0.223% | 0.1%
24 | 10 | 0.151% | -0.116% | 0.231% | 0.1%
24 | 11 | 0.089% | -0.062% | 0.169% | 0.1%
24 | 12 | 1.105% | 1.016% | 0.448% | 0.1%
24 | 13 | 0.924% | -0.181% | 0.706% | 0.1%
24 | 14 | 0.527% | -0.397% | 0.852% | 0.1%
24 | 15 | 0.142% | -0.385% | 0.531% | 0.1%
24 | 16 | 0.018% | -0.124% | 0.229% | 0.1%
24 | 17 | -0.027% | -0.045% | 0.044% | 0.1%
24 | 18 | 0.027% | 0.054% | 0.006% | 0.1%
24 | 19 | -0.036% | -0.063% | -0.012% | 0.1%
24 | 20 | -0.142% | -0.106% | -0.05% | 0.1%
Deviations of %error on test id #25. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | -0.018% | 0% | 0% | 0.1%
25 | 1 | -0.036% | -0.018% | 0% | 0.1%
25 | 2 | -0.062% | -0.026% | -0.039% | 0.1%
25 | 3 | -0.044% | 0.018% | -0.047% | 0.1%
25 | 4 | -0.08% | -0.036% | -0.062% | 0.1%
Deviations of %error on test id #26. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | -0.036% | 0% | 0% | 0.1%
26 | 1 | -0.071% | -0.035% | 0% | 0.1%
26 | 2 | 0.018% | 0.089% | -0.03% | 0.1%
26 | 3 | -0.062% | -0.08% | -0.038% | 0.1%
Deviations of %error on test id #27. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | -13.2% | 0% | 0% | 0.1%
27 | 1 | -7.056% | 6.145% | 0% | 0.1%
27 | 2 | 0.258% | 7.314% | -6.666% | 0.1%
27 | 3 | 0.16% | -0.098% | -2.213% | 0.1%
27 | 4 | 0.116% | -0.044% | 0.178% | 0.1%
27 | 5 | 0.303% | 0.187% | 0.193% | 0.1%
27 | 6 | 0.393% | 0.09% | 0.271% | 0.1%
27 | 7 | 0.536% | 0.143% | 0.411% | 0.1%
27 | 8 | 0.357% | -0.179% | 0.429% | 0.1%
27 | 9 | 0.16% | -0.197% | 0.351% | 0.1%
27 | 10 | 0.098% | -0.062% | 0.205% | 0.1%
27 | 11 | -0.062% | -0.16% | 0.065% | 0.1%
27 | 12 | -0.018% | 0.044% | 0.006% | 0.1%
27 | 13 | -0.018% | 0% | -0.033% | 0.1%
27 | 14 | -0.062% | -0.044% | -0.033% | 0.1%
Deviations of %error on test id #28. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | -0.044% | 0% | 0% | 0.1%
28 | 1 | -0.08% | -0.036% | 0% | 0.1%
28 | 2 | -0.009% | 0.071% | -0.044% | 0.1%
28 | 3 | -0.071% | -0.062% | -0.053% | 0.1%
Deviations of %error on test id #29. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | -0.036% | 0% | 0% | 0.1%
29 | 1 | -0.071% | -0.035% | 0% | 0.1%
29 | 2 | -0.036% | 0.035% | -0.048% | 0.1%
29 | 3 | -0.071% | -0.035% | -0.059% | 0.1%
Deviations of %error on test id #30. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | -12.59% | 0% | 0% | 0.1%
30 | 1 | -7.453% | 5.141% | 0% | 0.1%
30 | 2 | -2.783% | 4.67% | -7.61% | 0.1%
30 | 3 | -0.062% | 2.721% | -3.433% | 0.1%
30 | 4 | 0.018% | 0.08% | -0.942% | 0.1%
30 | 5 | 0.267% | 0.249% | 0.074% | 0.1%
30 | 6 | 0.205% | -0.062% | 0.163% | 0.1%
30 | 7 | 0.178% | -0.027% | 0.217% | 0.1%
30 | 8 | 0.357% | 0.179% | 0.247% | 0.1%
30 | 9 | 0.16% | -0.197% | 0.232% | 0.1%
30 | 10 | 0.134% | -0.026% | 0.217% | 0.1%
30 | 11 | 0.071% | -0.063% | 0.122% | 0.1%
30 | 12 | -0.08% | -0.151% | 0.042% | 0.1%
30 | 13 | -0.027% | 0.053% | -0.012% | 0.1%
30 | 14 | -0.036% | -0.009% | -0.048% | 0.1%
30 | 15 | -0.071% | -0.035% | -0.045% | 0.1%
Deviations of %error on test id #31. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | -0.204% | 0% | 0% | 0.1%
31 | 1 | 0.107% | 0.311% | 0% | 0.1%
31 | 2 | 0.366% | 0.259% | 0.09% | 0.1%
31 | 3 | 0.232% | -0.134% | 0.235% | 0.1%
31 | 4 | 0.125% | -0.107% | 0.241% | 0.1%
31 | 5 | 0.009% | -0.116% | 0.122% | 0.1%
31 | 6 | 0.178% | 0.169% | 0.104% | 0.1%
31 | 7 | 0.169% | -0.009% | 0.119% | 0.1%
31 | 8 | 0.142% | -0.027% | 0.163% | 0.1%
31 | 9 | 0.098% | -0.044% | 0.136% | 0.1%
Deviations of %error on test id #32. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | 0.009% | 0% | 0% | 0.1%
32 | 1 | 0.08% | 0.071% | 0% | 0.1%
32 | 2 | 0.08% | 0% | 0.056% | 0.1%
32 | 3 | 0.232% | 0.152% | 0.131% | 0.1%
32 | 4 | 0.178% | -0.054% | 0.163% | 0.1%
32 | 5 | 0.178% | 0% | 0.196% | 0.1%
32 | 6 | 0.321% | 0.143% | 0.226% | 0.1%
32 | 7 | 0.196% | -0.125% | 0.232% | 0.1%
32 | 8 | 0.134% | -0.062% | 0.217% | 0.1%
32 | 9 | 0.053% | -0.081% | 0.128% | 0.1%
32 | 10 | -0.062% | -0.115% | 0.042% | 0.1%
32 | 11 | -0.071% | -0.009% | -0.027% | 0.1%
32 | 12 | -0.044% | 0.027% | -0.059% | 0.1%
32 | 13 | -0.071% | -0.027% | -0.062% | 0.1%
Deviations of %error on test id #33. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | -12.75% | 0% | 0% | 0.1%
33 | 1 | -6.453% | 6.297% | 0% | 0.1%
33 | 2 | -2.386% | 4.067% | -7.196% | 0.1%
33 | 3 | 0.107% | 2.493% | -2.911% | 0.1%
33 | 4 | 0.098% | -0.009% | -0.727% | 0.1%
33 | 5 | 0.241% | 0.143% | 0.149% | 0.1%
33 | 6 | 0.178% | -0.063% | 0.172% | 0.1%
33 | 7 | 0.267% | 0.089% | 0.229% | 0.1%
33 | 8 | 0.321% | 0.054% | 0.255% | 0.1%
Deviations of %error on test id #34. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | -8.641% | 0% | 0% | 0.1%
34 | 1 | -3.184% | 5.457% | 0% | 0.1%
34 | 2 | -0.115% | 3.069% | -3.98% | 0.1%
34 | 3 | -0.027% | 0.088% | -1.109% | 0.1%
34 | 4 | 0.339% | 0.366% | 0.066% | 0.1%
34 | 5 | 0.187% | -0.152% | 0.166% | 0.1%
34 | 6 | 0.169% | -0.018% | 0.232% | 0.1%
34 | 7 | 0.446% | 0.277% | 0.267% | 0.1%
34 | 8 | 0.169% | -0.277% | 0.261% | 0.1%
34 | 9 | 0.107% | -0.062% | 0.241% | 0.1%
34 | 10 | 0.08% | -0.027% | 0.119% | 0.1%
34 | 11 | -0.089% | -0.169% | 0.033% | 0.1%
34 | 12 | -0.071% | 0.018% | -0.027% | 0.1%
34 | 13 | -0.08% | -0.009% | -0.08% | 0.1%
34 | 14 | -0.098% | -0.018% | -0.083% | 0.1%
Deviations of %error on test id #35. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | -3.4% | 0% | 0% | 0.1%
35 | 1 | -0.68% | 2.72% | 0% | 0.1%
35 | 2 | -0.036% | 0.644% | -1.372% | 0.1%
35 | 3 | 0.267% | 0.303% | -0.15% | 0.1%
35 | 4 | 0.25% | -0.017% | 0.16% | 0.1%
35 | 5 | 0.258% | 0.008% | 0.258% | 0.1%
35 | 6 | 0.464% | 0.206% | 0.324% | 0.1%
35 | 7 | 0.214% | -0.25% | 0.312% | 0.1%
35 | 8 | 0.125% | -0.089% | 0.268% | 0.1%
35 | 9 | 0.053% | -0.072% | 0.131% | 0.1%
35 | 10 | -0.062% | -0.115% | 0.039% | 0.1%
35 | 11 | -0.071% | -0.009% | -0.027% | 0.1%
35 | 12 | -0.098% | -0.027% | -0.077% | 0.1%
35 | 13 | -0.08% | 0.018% | -0.083% | 0.1%
Deviations of %error on test id #36. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | 0.285% | 0% | 0% | 0.1%
36 | 1 | 0.258% | -0.027% | 0% | 0.1%
36 | 2 | 0.214% | -0.044% | 0.252% | 0.1%
36 | 3 | 0.393% | 0.179% | 0.288% | 0.1%
36 | 4 | 0.205% | -0.188% | 0.271% | 0.1%
36 | 5 | 0.098% | -0.107% | 0.232% | 0.1%
36 | 6 | 0.044% | -0.054% | 0.116% | 0.1%
36 | 7 | -0.062% | -0.106% | 0.027% | 0.1%
36 | 8 | -0.106% | -0.044% | -0.041% | 0.1%
36 | 9 | -0.071% | 0.035% | -0.08% | 0.1%
36 | 10 | -0.115% | -0.044% | -0.097% | 0.1%
Deviations of %error on test id #37. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | 0.276% | 0% | 0% | 0.1%
37 | 1 | 0.16% | -0.116% | 0% | 0.1%
37 | 2 | 0.125% | -0.035% | 0.187% | 0.1%
37 | 3 | 0.009% | -0.116% | 0.098% | 0.1%
37 | 4 | 0.134% | 0.125% | 0.089% | 0.1%
37 | 5 | 0.187% | 0.053% | 0.11% | 0.1%
37 | 6 | 0.125% | -0.062% | 0.149% | 0.1%
37 | 7 | 0.036% | -0.089% | 0.116% | 0.1%
Deviations of %error on test id #38. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | 0.16% | 0% | 0% | 0.1%
38 | 1 | 0.08% | -0.08% | 0% | 0.1%
38 | 2 | 0.116% | 0.036% | 0.119% | 0.1%
38 | 3 | 0.285% | 0.169% | 0.16% | 0.1%
38 | 4 | 0.196% | -0.089% | 0.199% | 0.1%
38 | 5 | 0.089% | -0.107% | 0.19% | 0.1%
38 | 6 | 0.053% | -0.036% | 0.113% | 0.1%
38 | 7 | -0.133% | -0.186% | 0.003% | 0.1%
38 | 8 | -0.071% | 0.062% | -0.05% | 0.1%
38 | 9 | -0.106% | -0.035% | -0.103% | 0.1%
38 | 10 | -0.071% | 0.035% | -0.083% | 0.1%
38 | 11 | -0.071% | 0% | -0.083% | 0.1%
Deviations of %error on test id #39. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | 0.321% | 0% | 0% | 0.1%
39 | 1 | 0.142% | -0.179% | 0% | 0.1%
39 | 2 | 0.018% | -0.124% | 0.16% | 0.1%
39 | 3 | -0.018% | -0.036% | 0.047% | 0.1%
39 | 4 | -0.098% | -0.08% | -0.033% | 0.1%
39 | 5 | -0.036% | 0.062% | -0.051% | 0.1%
Deviations of %error on test id #40. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | -0.018% | 0% | 0% | 0.1%
40 | 1 | -0.115% | -0.097% | 0% | 0.1%
40 | 2 | -0.071% | 0.044% | -0.068% | 0.1%
40 | 3 | -0.106% | -0.035% | -0.097% | 0.1%

## 9. PF = 1.0
[<<]
Accuracy Test Report of 'PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:08:50 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 20 | 20 |  | ±0.1% | ±0.5% | -0.009% | PASSED | 2023-07-27 12:43
Yes | 1 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.5% | -0.039% | PASSED | 2023-07-27 12:44
Yes | 3 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.5% | -0.018% | PASSED | 2023-07-27 12:44
Yes | 4 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0% | PASSED | 2023-07-27 12:45
Yes | 7 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0.009% | PASSED | 2023-07-27 12:45
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 0.009% | 0% | 0% | 0.1%
0 | 1 | -0.009% | -0.018% | 0% | 0.1%
0 | 2 | -0.018% | -0.009% | -0.006% | 0.1%
0 | 3 | 0% | 0.018% | -0.009% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -0.092% | 0% | 0% | 0.1%
1 | 1 | -0.071% | 0.021% | 0% | 0.1%
1 | 2 | -0.036% | 0.035% | -0.066% | 0.1%
1 | 3 | -0.011% | 0.025% | -0.039% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.009% | 0% | 0% | 0.1%
3 | 1 | -0.036% | -0.045% | 0% | 0.1%
3 | 2 | 0.004% | 0.04% | -0.008% | 0.1%
3 | 3 | -0.022% | -0.026% | -0.018% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0% | 0% | 0% | 0.1%
4 | 1 | 0% | 0% | 0% | 0.1%
4 | 2 | -0.009% | -0.009% | -0.003% | 0.1%
4 | 3 | 0.009% | 0.018% | 0% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | -0.011% | 0% | 0% | 0.1%
7 | 1 | 0.012% | 0.023% | 0% | 0.1%
7 | 2 | 0.018% | 0.006% | 0.006% | 0.1%
7 | 3 | -0.002% | -0.02% | 0.009% | 0.1%

## 10. PF = 0.5L
[<<]
Accuracy Test Report of 'PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:08:52 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | -0.059% | PASSED | 2023-07-27 12:46
Yes | 1 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | 0.166% | PASSED | 2023-07-27 12:47
Yes | 3 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0% | PASSED | 2023-07-27 12:47
Yes | 4 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.021% | PASSED | 2023-07-27 12:48
Yes | 7 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.055% | PASSED | 2023-07-27 12:49
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -31.65% | 0% | 0% | 0.1%
0 | 1 | -28.7% | 2.954% | 0% | 0.1%
0 | 2 | -16.42% | 12.28% | -25.59% | 0.1%
0 | 3 | -8.67% | 7.749% | -17.93% | 0.1%
0 | 4 | -2.276% | 6.394% | -9.122% | 0.1%
0 | 5 | 0.303% | 2.579% | -3.548% | 0.1%
0 | 6 | 0.303% | 0% | -0.557% | 0.1%
0 | 7 | 0.142% | -0.161% | 0.249% | 0.1%
0 | 8 | 0% | -0.142% | 0.148% | 0.1%
0 | 9 | 0.196% | 0.196% | 0.113% | 0.1%
0 | 10 | 0.321% | 0.125% | 0.172% | 0.1%
0 | 11 | 0.178% | -0.143% | 0.232% | 0.1%
0 | 12 | 0.053% | -0.125% | 0.184% | 0.1%
0 | 13 | 1.242% | 1.189% | 0.491% | 0.1%
0 | 14 | 1.663% | 0.421% | 0.986% | 0.1%
0 | 15 | 1.242% | -0.421% | 1.382% | 0.1%
0 | 16 | 0.572% | -0.67% | 1.159% | 0.1%
0 | 17 | 0.053% | -0.519% | 0.622% | 0.1%
0 | 18 | -0.089% | -0.142% | 0.179% | 0.1%
0 | 19 | -0.142% | -0.053% | -0.059% | 0.1%
Deviations of %error on test id #1. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -21.72% | 0% | 0% | 0.1%
1 | 1 | -10.07% | 11.65% | 0% | 0.1%
1 | 2 | -1.884% | 8.183% | -11.22% | 0.1%
1 | 3 | -0.012% | 1.872% | -3.988% | 0.1%
1 | 4 | 0.196% | 0.208% | -0.567% | 0.1%
1 | 5 | 0.113% | -0.083% | 0.099% | 0.1%
1 | 6 | 0.184% | 0.071% | 0.164% | 0.1%
1 | 7 | 0.202% | 0.018% | 0.166% | 0.1%
Deviations of %error on test id #3. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -19.2% | 0% | 0% | 0.1%
3 | 1 | -7.232% | 11.97% | 0% | 0.1%
3 | 2 | -3.879% | 3.353% | -10.11% | 0.1%
3 | 3 | -0.036% | 3.843% | -3.716% | 0.1%
3 | 4 | 0.375% | 0.411% | -1.18% | 0.1%
3 | 5 | 0.906% | 0.531% | 0.415% | 0.1%
3 | 6 | 0.797% | -0.109% | 0.693% | 0.1%
3 | 7 | 0.852% | 0.055% | 0.852% | 0.1%
3 | 8 | 0.825% | -0.027% | 0.825% | 0.1%
3 | 9 | 0.599% | -0.226% | 0.759% | 0.1%
3 | 10 | 0.312% | -0.287% | 0.579% | 0.1%
3 | 11 | 0.125% | -0.187% | 0.345% | 0.1%
3 | 12 | 0.071% | -0.054% | 0.169% | 0.1%
3 | 13 | -0.009% | -0.08% | 0.062% | 0.1%
3 | 14 | -0.062% | -0.053% | 0% | 0.1%
Deviations of %error on test id #4. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -26.71% | 0% | 0% | 0.1%
4 | 1 | -11.22% | 15.49% | 0% | 0.1%
4 | 2 | -1.073% | 10.15% | -13% | 0.1%
4 | 3 | 0.142% | 1.215% | -4.051% | 0.1%
4 | 4 | 0.303% | 0.161% | -0.209% | 0.1%
4 | 5 | 0.42% | 0.117% | 0.288% | 0.1%
4 | 6 | 0.258% | -0.162% | 0.327% | 0.1%
4 | 7 | 0.116% | -0.142% | 0.265% | 0.1%
4 | 8 | 0.009% | -0.107% | 0.128% | 0.1%
4 | 9 | 0.044% | 0.035% | 0.056% | 0.1%
4 | 10 | 0.009% | -0.035% | 0.021% | 0.1%
4 | 11 | 0.009% | 0% | 0.021% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.098% | 0% | 0% | 0.1%
7 | 1 | 0.053% | -0.045% | 0% | 0.1%
7 | 2 | 0.053% | 0% | 0.068% | 0.1%
7 | 3 | 0.06% | 0.007% | 0.055% | 0.1%

## 11. PF = 0.8C
[<<]
Accuracy Test Report of 'PF = 0.8C'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / PF = 0.8C'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:08:56 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | -0.089% | PASSED | 2023-07-27 12:50
Yes | 1 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | -0.128% | PASSED | 2023-07-27 12:50
Yes | 3 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | -0.086% | PASSED | 2023-07-27 12:51
Yes | 4 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | -0.101% | PASSED | 2023-07-27 12:51
Yes | 7 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 0.8C, cos(ϕ), ϕ=323.13° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | -0.014% | PASSED | 2023-07-27 12:52
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -4.369% | 0% | 0% | 0.1%
0 | 1 | -3.417% | 0.952% | 0% | 0.1%
0 | 2 | -1.764% | 1.653% | -3.183% | 0.1%
0 | 3 | -0.195% | 1.569% | -1.792% | 0.1%
0 | 4 | -0.142% | 0.053% | -0.7% | 0.1%
0 | 5 | -0.071% | 0.071% | -0.136% | 0.1%
0 | 6 | -0.053% | 0.018% | -0.089% | 0.1%
Deviations of %error on test id #1. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -3.753% | 0% | 0% | 0.1%
1 | 1 | -2.822% | 0.931% | 0% | 0.1%
1 | 2 | -0.595% | 2.227% | -2.39% | 0.1%
1 | 3 | -0.319% | 0.276% | -1.245% | 0.1%
1 | 4 | -0.106% | 0.213% | -0.34% | 0.1%
1 | 5 | -0.136% | -0.03% | -0.187% | 0.1%
1 | 6 | -0.154% | -0.018% | -0.132% | 0.1%
1 | 7 | -0.095% | 0.059% | -0.128% | 0.1%
Deviations of %error on test id #3. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -8.156% | 0% | 0% | 0.1%
3 | 1 | -4.304% | 3.852% | 0% | 0.1%
3 | 2 | -4.369% | -0.065% | -5.61% | 0.1%
3 | 3 | -2.85% | 1.519% | -3.841% | 0.1%
3 | 4 | -0.907% | 1.943% | -2.709% | 0.1%
3 | 5 | -0.977% | -0.07% | -1.578% | 0.1%
3 | 6 | -0.372% | 0.605% | -0.752% | 0.1%
3 | 7 | -0.053% | 0.319% | -0.467% | 0.1%
3 | 8 | -0.16% | -0.107% | -0.195% | 0.1%
3 | 9 | -0.106% | 0.054% | -0.106% | 0.1%
3 | 10 | -0.115% | -0.009% | -0.127% | 0.1%
3 | 11 | -0.098% | 0.017% | -0.106% | 0.1%
3 | 12 | -0.044% | 0.054% | -0.086% | 0.1%
Deviations of %error on test id #4. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -7.923% | 0% | 0% | 0.1%
4 | 1 | -4.871% | 3.052% | 0% | 0.1%
4 | 2 | -1.116% | 3.755% | -4.637% | 0.1%
4 | 3 | -0.186% | 0.93% | -2.058% | 0.1%
4 | 4 | -0.036% | 0.15% | -0.446% | 0.1%
4 | 5 | -0.098% | -0.062% | -0.107% | 0.1%
4 | 6 | -0.098% | 0% | -0.077% | 0.1%
4 | 7 | -0.106% | -0.008% | -0.101% | 0.1%
Deviations of %error on test id #7. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | -0.13% | 0% | 0% | 0.1%
7 | 1 | 0.027% | 0.157% | 0% | 0.1%
7 | 2 | -0.014% | -0.041% | -0.039% | 0.1%
7 | 3 | -0.012% | 0.002% | 0% | 0.1%
7 | 4 | -0.016% | -0.004% | -0.014% | 0.1%

## 12. 5th Harmonics
[<<]
Accuracy Test Report of '5th Harmonics'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / 5th Harmonics'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:09:00 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Harmonics' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
1 | 0 | 0 | 1
3. Test Cases
Total 1 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 1 | RP | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° | {UcUbUa, 5th-10%, IcIbIa, 5th-40%} | 12 | 20 |  | ±0.1% | ±1% | -0.051% | PASSED | 2023-07-27 12:53
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #1. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -1.356% | 0% | 0% | 0.1%
1 | 1 | -6.644% | -5.288% | 0% | 0.1%
1 | 2 | -4.769% | 1.875% | -4.256% | 0.1%
1 | 3 | -2.411% | 2.358% | -4.608% | 0.1%
1 | 4 | -0.003% | 2.408% | -2.394% | 0.1%
1 | 5 | 0.443% | 0.446% | -0.657% | 0.1%
1 | 6 | 0.42% | -0.023% | 0.287% | 0.1%
1 | 7 | 0.247% | -0.173% | 0.37% | 0.1%
1 | 8 | 0.033% | -0.214% | 0.233% | 0.1%
1 | 9 | -0.038% | -0.071% | 0.081% | 0.1%
1 | 10 | -0.056% | -0.018% | -0.02% | 0.1%
1 | 11 | -0.059% | -0.003% | -0.051% | 0.1%

## 13. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:09:01 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RP | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | -0.022% | PASSED | 2023-07-27 12:53
Yes | 1 | RP | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.002% | PASSED | 2023-07-27 12:54
Yes | 2 | RP | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.021% | PASSED | 2023-07-27 12:54
Yes | 3 | RP | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.001% | PASSED | 2023-07-27 12:55
Yes | 4 | RP | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.003% | PASSED | 2023-07-27 12:55
Yes | 5 | RP | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | -0.007% | PASSED | 2023-07-27 12:55
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 0.581% | 0% | 0% | 0.1%
0 | 1 | 0.013% | -0.568% | 0% | 0.1%
0 | 2 | -0.018% | -0.031% | 0.192% | 0.1%
0 | 3 | -0.04% | -0.022% | -0.015% | 0.1%
0 | 4 | -0.009% | 0.031% | -0.022% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0.004% | 0% | 0% | 0.1%
1 | 1 | 0.004% | 0% | 0% | 0.1%
1 | 2 | -0.009% | -0.013% | 0% | 0.1%
1 | 3 | 0% | 0.009% | -0.002% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -0.031% | 0% | 0% | 0.1%
2 | 1 | 0% | 0.031% | 0% | 0.1%
2 | 2 | -0.022% | -0.022% | -0.018% | 0.1%
2 | 3 | -0.04% | -0.018% | -0.021% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.018% | 0% | 0% | 0.1%
3 | 1 | 0.009% | -0.009% | 0% | 0.1%
3 | 2 | -0.022% | -0.031% | 0.002% | 0.1%
3 | 3 | 0.009% | 0.031% | -0.001% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.004% | 0% | 0% | 0.1%
4 | 1 | 0% | -0.004% | 0% | 0.1%
4 | 2 | 0.013% | 0.013% | 0.006% | 0.1%
4 | 3 | -0.004% | -0.017% | 0.003% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -0.018% | 0% | 0% | 0.1%
5 | 1 | 0.009% | 0.027% | 0% | 0.1%
5 | 2 | -0.013% | -0.022% | -0.007% | 0.1%
5 | 3 | -0.018% | -0.005% | -0.007% | 0.1%

## 14. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:09:05 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RP | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | -0.139% | PASSED | 2023-07-27 12:56
Yes | 1 | RP | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | -0.021% | PASSED | 2023-07-27 12:57
Yes | 2 | RP | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.009% | PASSED | 2023-07-27 12:57
Yes | 3 | RP | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | -0.024% | PASSED | 2023-07-27 12:58
Yes | 4 | RP | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.268% | PASSED | 2023-07-27 12:58
Yes | 5 | RP | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | 0.756% | PASSED | 2023-07-27 12:59
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -18.31% | 0% | 0% | 0.1%
0 | 1 | -7.825% | 10.48% | 0% | 0.1%
0 | 2 | -2.698% | 5.127% | -9.61% | 0.1%
0 | 3 | 0.125% | 2.823% | -3.466% | 0.1%
0 | 4 | -0.027% | -0.152% | -0.867% | 0.1%
0 | 5 | -0.053% | -0.026% | 0.015% | 0.1%
0 | 6 | -0.044% | 0.009% | -0.041% | 0.1%
0 | 7 | 0.169% | 0.213% | 0.024% | 0.1%
0 | 8 | 0.232% | 0.063% | 0.119% | 0.1%
0 | 9 | 0.339% | 0.107% | 0.247% | 0.1%
0 | 10 | 0.752% | 0.413% | 0.441% | 0.1%
0 | 11 | 0.455% | -0.297% | 0.515% | 0.1%
0 | 12 | 0.321% | -0.134% | 0.509% | 0.1%
0 | 13 | 0.151% | -0.17% | 0.309% | 0.1%
0 | 14 | 0% | -0.151% | 0.157% | 0.1%
0 | 15 | -0.266% | -0.266% | -0.038% | 0.1%
0 | 16 | -0.319% | -0.053% | -0.195% | 0.1%
0 | 17 | -0.248% | 0.071% | -0.278% | 0.1%
0 | 18 | -0.124% | 0.124% | -0.23% | 0.1%
0 | 19 | -0.044% | 0.08% | -0.139% | 0.1%
Deviations of %error on test id #1. Total 19 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -10.01% | 0% | 0% | 0.1%
1 | 1 | -0.009% | 9.998% | 0% | 0.1%
1 | 2 | 0.134% | 0.143% | -3.294% | 0.1%
1 | 3 | 0.411% | 0.277% | 0.179% | 0.1%
1 | 4 | 0.303% | -0.108% | 0.283% | 0.1%
1 | 5 | 0.187% | -0.116% | 0.3% | 0.1%
1 | 6 | 0.384% | 0.197% | 0.291% | 0.1%
1 | 7 | 0.223% | -0.161% | 0.265% | 0.1%
1 | 8 | 0.178% | -0.045% | 0.262% | 0.1%
1 | 9 | 0.062% | -0.116% | 0.154% | 0.1%
1 | 10 | 0.214% | 0.152% | 0.151% | 0.1%
1 | 11 | 0.626% | 0.412% | 0.301% | 0.1%
1 | 12 | 0.446% | -0.18% | 0.429% | 0.1%
1 | 13 | 0.312% | -0.134% | 0.461% | 0.1%
1 | 14 | 0.169% | -0.143% | 0.309% | 0.1%
1 | 15 | -0.044% | -0.213% | 0.146% | 0.1%
1 | 16 | 0% | 0.044% | 0.042% | 0.1%
1 | 17 | -0.027% | -0.027% | -0.024% | 0.1%
1 | 18 | -0.036% | -0.009% | -0.021% | 0.1%
Deviations of %error on test id #2. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -19.4% | 0% | 0% | 0.1%
2 | 1 | -10.62% | 8.779% | 0% | 0.1%
2 | 2 | -7.163% | 3.459% | -12.39% | 0.1%
2 | 3 | 0.116% | 7.279% | -5.89% | 0.1%
2 | 4 | 0.053% | -0.063% | -2.331% | 0.1%
2 | 5 | -0.018% | -0.071% | 0.05% | 0.1%
2 | 6 | 0.089% | 0.107% | 0.041% | 0.1%
2 | 7 | 0.339% | 0.25% | 0.137% | 0.1%
2 | 8 | 0.527% | 0.188% | 0.318% | 0.1%
2 | 9 | 0.437% | -0.09% | 0.434% | 0.1%
2 | 10 | 0.366% | -0.071% | 0.443% | 0.1%
2 | 11 | 0.098% | -0.268% | 0.3% | 0.1%
2 | 12 | 0.036% | -0.062% | 0.167% | 0.1%
2 | 13 | -0.009% | -0.045% | 0.042% | 0.1%
2 | 14 | 0% | 0.009% | 0.009% | 0.1%
Deviations of %error on test id #3. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -9.23% | 0% | 0% | 0.1%
3 | 1 | 0.134% | 9.364% | 0% | 0.1%
3 | 2 | 1.388% | 1.254% | -2.569% | 0.1%
3 | 3 | 0.644% | -0.744% | 0.722% | 0.1%
3 | 4 | -0.036% | -0.68% | 0.665% | 0.1%
3 | 5 | -0.027% | 0.009% | 0.194% | 0.1%
3 | 6 | -0.027% | 0% | -0.03% | 0.1%
3 | 7 | -0.018% | 0.009% | -0.024% | 0.1%
Deviations of %error on test id #4. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -7.651% | 0% | 0% | 0.1%
4 | 1 | -6.422% | 1.229% | 0% | 0.1%
4 | 2 | 0.258% | 6.68% | -4.605% | 0.1%
4 | 3 | 0.33% | 0.072% | -1.945% | 0.1%
4 | 4 | 0.232% | -0.098% | 0.273% | 0.1%
4 | 5 | 0.241% | 0.009% | 0.268% | 0.1%
Deviations of %error on test id #5. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -25.48% | 0% | 0% | 0.1%
5 | 1 | -11.12% | 14.36% | 0% | 0.1%
5 | 2 | -10.23% | 0.887% | -15.61% | 0.1%
5 | 3 | -0.151% | 10.08% | -7.165% | 0.1%
5 | 4 | 0.294% | 0.445% | -3.362% | 0.1%
5 | 5 | 0.25% | -0.044% | 0.131% | 0.1%
5 | 6 | 0.241% | -0.009% | 0.262% | 0.1%
5 | 7 | 0.116% | -0.125% | 0.202% | 0.1%
5 | 8 | 0.16% | 0.044% | 0.172% | 0.1%
5 | 9 | 0.205% | 0.045% | 0.16% | 0.1%
5 | 10 | 0.348% | 0.143% | 0.238% | 0.1%
5 | 11 | 0.223% | -0.125% | 0.259% | 0.1%
5 | 12 | 0.232% | 0.009% | 0.268% | 0.1%
5 | 13 | 0.107% | -0.125% | 0.187% | 0.1%
5 | 14 | 0.116% | 0.009% | 0.152% | 0.1%
5 | 15 | 0.68% | 0.564% | 0.301% | 0.1%
5 | 16 | 1.324% | 0.644% | 0.707% | 0.1%
5 | 17 | 0.915% | -0.409% | 0.973% | 0.1%
5 | 18 | 0.906% | -0.009% | 1.048% | 0.1%
5 | 19 | 0.446% | -0.46% | 0.756% | 0.1%

## 15. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:09:11 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.056% | PASSED | 2023-07-27 12:59
Yes | 1 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.034% | PASSED | 2023-07-27 13:00
Yes | 2 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.043% | PASSED | 2023-07-27 13:00
Yes | 3 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.044% | PASSED | 2023-07-27 13:01
Yes | 4 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.015% | PASSED | 2023-07-27 13:01
Yes | 5 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.018% | PASSED | 2023-07-27 13:01
Yes | 6 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.031% | PASSED | 2023-07-27 13:02
Yes | 7 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.003% | PASSED | 2023-07-27 13:02
Yes | 8 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.011% | PASSED | 2023-07-27 13:03
Yes | 9 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.007% | PASSED | 2023-07-27 13:03
Yes | 10 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.003% | PASSED | 2023-07-27 13:03
Yes | 11 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.033% | PASSED
Yes | 12 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.018% | PASSED | 2023-07-27 13:04
Yes | 13 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.028% | PASSED | 2023-07-27 13:05
Yes | 14 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.024% | PASSED | 2023-07-27 13:05
Yes | 15 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.022% | PASSED | 2023-07-27 13:06
Yes | 16 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.03% | PASSED | 2023-07-27 13:06
Yes | 17 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.034% | PASSED | 2023-07-27 13:07
Yes | 18 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.033% | PASSED | 2023-07-27 13:07
Yes | 19 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.022% | PASSED | 2023-07-27 13:08
Yes | 20 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.043% | PASSED | 2023-07-27 13:08
Yes | 21 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.047% | PASSED | 2023-07-27 13:09
Yes | 22 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.035% | PASSED | 2023-07-27 13:09
Yes | 23 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.038% | PASSED | 2023-07-27 13:10
Yes | 24 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.056% | PASSED | 2023-07-27 13:10
Yes | 25 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.021% | PASSED | 2023-07-27 13:11
Yes | 26 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.042% | PASSED | 2023-07-27 13:12
Yes | 27 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.034% | PASSED | 2023-07-27 13:12
Yes | 28 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.041% | PASSED | 2023-07-27 13:13
Yes | 29 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.079% | PASSED | 2023-07-27 15:51
Yes | 30 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.064% | PASSED | 2023-07-27 13:14
Yes | 31 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.077% | PASSED | 2023-07-27 13:15
Yes | 32 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.053% | PASSED | 2023-07-27 13:16
Yes | 33 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.073% | PASSED | 2023-07-27 13:17
Yes | 34 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.061% | PASSED | 2023-07-27 13:18
Yes | 35 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.095% | PASSED | 2023-07-27 15:51
Yes | 36 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.071% | PASSED | 2023-07-27 13:19
Yes | 37 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.056% | PASSED | 2023-07-27 13:20
Yes | 38 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.067% | PASSED | 2023-07-27 13:22
Yes | 39 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.117% | PASSED | 2023-07-27 15:51
Yes | 40 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 1.0, cos(ϕ), ϕ=0° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.102% | PASSED | 2023-07-27 15:52
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -0.031% | 0% | 0% | 0.1%
0 | 1 | -0.075% | -0.044% | 0% | 0.1%
0 | 2 | -0.022% | 0.053% | -0.043% | 0.1%
0 | 3 | -0.071% | -0.049% | -0.056% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -0.071% | 0% | 0% | 0.1%
1 | 1 | -0.018% | 0.053% | 0% | 0.1%
1 | 2 | -0.067% | -0.049% | -0.052% | 0.1%
1 | 3 | -0.018% | 0.049% | -0.034% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -0.053% | 0% | 0% | 0.1%
2 | 1 | -0.031% | 0.022% | 0% | 0.1%
2 | 2 | -0.058% | -0.027% | -0.047% | 0.1%
2 | 3 | -0.04% | 0.018% | -0.043% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -0.071% | 0% | 0% | 0.1%
3 | 1 | -0.04% | 0.031% | 0% | 0.1%
3 | 2 | -0.022% | 0.018% | -0.044% | 0.1%
3 | 3 | -0.071% | -0.049% | -0.044% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -0.071% | 0% | 0% | 0.1%
4 | 1 | -0.031% | 0.04% | 0% | 0.1%
4 | 2 | -0.031% | 0% | -0.044% | 0.1%
4 | 3 | 0.018% | 0.049% | -0.015% | 0.1%
Deviations of %error on test id #5. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -0.031% | 0% | 0% | 0.1%
5 | 1 | -0.044% | -0.013% | 0% | 0.1%
5 | 2 | -0.027% | 0.017% | -0.034% | 0.1%
5 | 3 | -0.036% | -0.009% | -0.036% | 0.1%
5 | 4 | 0.009% | 0.045% | -0.018% | 0.1%
Deviations of %error on test id #6. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | -0.036% | 0% | 0% | 0.1%
6 | 1 | -0.044% | -0.008% | 0% | 0.1%
6 | 2 | -0.031% | 0.013% | -0.037% | 0.1%
6 | 3 | -0.018% | 0.013% | -0.031% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | -0.036% | 0% | 0% | 0.1%
7 | 1 | 0.009% | 0.045% | 0% | 0.1%
7 | 2 | -0.031% | -0.04% | -0.019% | 0.1%
7 | 3 | 0.013% | 0.044% | -0.003% | 0.1%
Deviations of %error on test id #8. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | 0.018% | 0% | 0% | 0.1%
8 | 1 | -0.036% | -0.054% | 0% | 0.1%
8 | 2 | 0.013% | 0.049% | -0.002% | 0.1%
8 | 3 | -0.009% | -0.022% | -0.011% | 0.1%
Deviations of %error on test id #9. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | 0.009% | 0% | 0% | 0.1%
9 | 1 | -0.004% | -0.013% | 0% | 0.1%
9 | 2 | -0.018% | -0.014% | -0.004% | 0.1%
9 | 3 | 0% | 0.018% | -0.007% | 0.1%
Deviations of %error on test id #10. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | -0.036% | 0% | 0% | 0.1%
10 | 1 | 0.022% | 0.058% | 0% | 0.1%
10 | 2 | -0.031% | -0.053% | -0.015% | 0.1%
10 | 3 | 0.018% | 0.049% | 0.003% | 0.1%
Deviations of %error on test id #11. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | -0.004% | 0% | 0% | 0.1%
11 | 1 | 0.031% | 0.035% | 0% | 0.1%
11 | 2 | 0.044% | 0.013% | 0.024% | 0.1%
11 | 3 | 0.071% | 0.027% | 0.049% | 0.1%
11 | 4 | 0.027% | -0.044% | 0.047% | 0.1%
11 | 5 | 0.004% | -0.023% | 0.034% | 0.1%
11 | 6 | 0.076% | 0.072% | 0.036% | 0.1%
11 | 7 | -0.013% | -0.089% | 0.022% | 0.1%
11 | 8 | 0.071% | 0.084% | 0.045% | 0.1%
11 | 9 | 0.04% | -0.031% | 0.033% | 0.1%
Deviations of %error on test id #12. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | -0.022% | 0% | 0% | 0.1%
12 | 1 | 0.04% | 0.062% | 0% | 0.1%
12 | 2 | 0.013% | -0.027% | 0.01% | 0.1%
12 | 3 | 0% | -0.013% | 0.018% | 0.1%
Deviations of %error on test id #13. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | 0.004% | 0% | 0% | 0.1%
13 | 1 | 0.018% | 0.014% | 0% | 0.1%
13 | 2 | 0.004% | -0.014% | 0.009% | 0.1%
13 | 3 | 0.062% | 0.058% | 0.028% | 0.1%
Deviations of %error on test id #14. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | 0.022% | 0% | 0% | 0.1%
14 | 1 | 0.009% | -0.013% | 0% | 0.1%
14 | 2 | 0.022% | 0.013% | 0.018% | 0.1%
14 | 3 | 0.04% | 0.018% | 0.024% | 0.1%
Deviations of %error on test id #15. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | 0.013% | 0% | 0% | 0.1%
15 | 1 | 0.049% | 0.036% | 0% | 0.1%
15 | 2 | 0.009% | -0.04% | 0.024% | 0.1%
15 | 3 | 0.009% | 0% | 0.022% | 0.1%
Deviations of %error on test id #16. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | 0.013% | 0% | 0% | 0.1%
16 | 1 | 0.022% | 0.009% | 0% | 0.1%
16 | 2 | 0.058% | 0.036% | 0.031% | 0.1%
16 | 3 | 0.009% | -0.049% | 0.03% | 0.1%
Deviations of %error on test id #17. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | 0.018% | 0% | 0% | 0.1%
17 | 1 | 0.044% | 0.026% | 0% | 0.1%
17 | 2 | 0.013% | -0.031% | 0.025% | 0.1%
17 | 3 | 0.044% | 0.031% | 0.034% | 0.1%
Deviations of %error on test id #18. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | 0.004% | 0% | 0% | 0.1%
18 | 1 | 0.013% | 0.009% | 0% | 0.1%
18 | 2 | 0.036% | 0.023% | 0.018% | 0.1%
18 | 3 | 0.049% | 0.013% | 0.033% | 0.1%
Deviations of %error on test id #19. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | 0.04% | 0% | 0% | 0.1%
19 | 1 | 0.013% | -0.027% | 0% | 0.1%
19 | 2 | 0.049% | 0.036% | 0.034% | 0.1%
19 | 3 | 0.004% | -0.045% | 0.022% | 0.1%
Deviations of %error on test id #20. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | 0.049% | 0% | 0% | 0.1%
20 | 1 | 0.009% | -0.04% | 0% | 0.1%
20 | 2 | 0.044% | 0.035% | 0.034% | 0.1%
20 | 3 | 0.04% | -0.004% | 0.031% | 0.1%
20 | 4 | 0.044% | 0.004% | 0.043% | 0.1%
Deviations of %error on test id #21. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | 0.053% | 0% | 0% | 0.1%
21 | 1 | 0.009% | -0.044% | 0% | 0.1%
21 | 2 | 0.049% | 0.04% | 0.037% | 0.1%
21 | 3 | 0.044% | -0.005% | 0.034% | 0.1%
21 | 4 | 0.049% | 0.005% | 0.047% | 0.1%
Deviations of %error on test id #22. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | -0.004% | 0% | 0% | 0.1%
22 | 1 | 0.044% | 0.048% | 0% | 0.1%
22 | 2 | 0.022% | -0.022% | 0.021% | 0.1%
22 | 3 | 0.04% | 0.018% | 0.035% | 0.1%
Deviations of %error on test id #23. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | 0.004% | 0% | 0% | 0.1%
23 | 1 | 0.067% | 0.063% | 0% | 0.1%
23 | 2 | 0.004% | -0.063% | 0.025% | 0.1%
23 | 3 | 0.044% | 0.04% | 0.038% | 0.1%
Deviations of %error on test id #24. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | 0.049% | 0% | 0% | 0.1%
24 | 1 | 0.049% | 0% | 0% | 0.1%
24 | 2 | 0.067% | 0.018% | 0.055% | 0.1%
24 | 3 | 0.053% | -0.014% | 0.056% | 0.1%
Deviations of %error on test id #25. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | 0.04% | 0% | 0% | 0.1%
25 | 1 | 0.004% | -0.036% | 0% | 0.1%
25 | 2 | 0.027% | 0.023% | 0.024% | 0.1%
25 | 3 | 0.031% | 0.004% | 0.021% | 0.1%
Deviations of %error on test id #26. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | 0.058% | 0% | 0% | 0.1%
26 | 1 | 0.036% | -0.022% | 0% | 0.1%
26 | 2 | 0.04% | 0.004% | 0.045% | 0.1%
26 | 3 | 0.049% | 0.009% | 0.042% | 0.1%
Deviations of %error on test id #27. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | 0.053% | 0% | 0% | 0.1%
27 | 1 | 0.04% | -0.013% | 0% | 0.1%
27 | 2 | 0.044% | 0.004% | 0.046% | 0.1%
27 | 3 | 0.018% | -0.026% | 0.034% | 0.1%
Deviations of %error on test id #28. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | 0.089% | 0% | 0% | 0.1%
28 | 1 | 0.049% | -0.04% | 0% | 0.1%
28 | 2 | 0.031% | -0.018% | 0.056% | 0.1%
28 | 3 | 0.044% | 0.013% | 0.041% | 0.1%
Deviations of %error on test id #29. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | 0.044% | 0% | 0% | 0.1%
29 | 1 | 0.129% | 0.085% | 0% | 0.1%
29 | 2 | 0.076% | -0.053% | 0.083% | 0.1%
29 | 3 | 0.08% | 0.004% | 0.095% | 0.1%
29 | 4 | 0.08% | 0% | 0.079% | 0.1%
Deviations of %error on test id #30. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | 0.058% | 0% | 0% | 0.1%
30 | 1 | 0.098% | 0.04% | 0% | 0.1%
30 | 2 | 0.04% | -0.058% | 0.065% | 0.1%
30 | 3 | 0.053% | 0.013% | 0.064% | 0.1%
Deviations of %error on test id #31. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | 0.089% | 0% | 0% | 0.1%
31 | 1 | 0.058% | -0.031% | 0% | 0.1%
31 | 2 | 0.093% | 0.035% | 0.08% | 0.1%
31 | 3 | 0.08% | -0.013% | 0.077% | 0.1%
Deviations of %error on test id #32. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | 0.044% | 0% | 0% | 0.1%
32 | 1 | 0.049% | 0.005% | 0% | 0.1%
32 | 2 | 0.04% | -0.009% | 0.044% | 0.1%
32 | 3 | 0.071% | 0.031% | 0.053% | 0.1%
32 | 4 | 0.049% | -0.022% | 0.053% | 0.1%
Deviations of %error on test id #33. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | 0.053% | 0% | 0% | 0.1%
33 | 1 | 0.085% | 0.032% | 0% | 0.1%
33 | 2 | 0.049% | -0.036% | 0.062% | 0.1%
33 | 3 | 0.085% | 0.036% | 0.073% | 0.1%
Deviations of %error on test id #34. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | 0.062% | 0% | 0% | 0.1%
34 | 1 | 0.089% | 0.027% | 0% | 0.1%
34 | 2 | 0.04% | -0.049% | 0.064% | 0.1%
34 | 3 | 0.053% | 0.013% | 0.061% | 0.1%
Deviations of %error on test id #35. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | 0.093% | 0% | 0% | 0.1%
35 | 1 | 0.107% | 0.014% | 0% | 0.1%
35 | 2 | 0.067% | -0.04% | 0.089% | 0.1%
35 | 3 | 0.111% | 0.044% | 0.095% | 0.1%
Deviations of %error on test id #36. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | 0.031% | 0% | 0% | 0.1%
36 | 1 | 0.098% | 0.067% | 0% | 0.1%
36 | 2 | 0.036% | -0.062% | 0.055% | 0.1%
36 | 3 | 0.08% | 0.044% | 0.071% | 0.1%
Deviations of %error on test id #37. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | 0.076% | 0% | 0% | 0.1%
37 | 1 | 0.089% | 0.013% | 0% | 0.1%
37 | 2 | 0.044% | -0.045% | 0.07% | 0.1%
37 | 3 | 0.076% | 0.032% | 0.07% | 0.1%
37 | 4 | 0.049% | -0.027% | 0.056% | 0.1%
Deviations of %error on test id #38. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | 0.125% | 0% | 0% | 0.1%
38 | 1 | 0.036% | -0.089% | 0% | 0.1%
38 | 2 | 0.08% | 0.044% | 0.08% | 0.1%
38 | 3 | 0.085% | 0.005% | 0.067% | 0.1%
Deviations of %error on test id #39. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | 0.116% | 0% | 0% | 0.1%
39 | 1 | 0.107% | -0.009% | 0% | 0.1%
39 | 2 | 0.125% | 0.018% | 0.116% | 0.1%
39 | 3 | 0.12% | -0.005% | 0.117% | 0.1%
Deviations of %error on test id #40. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | 0.125% | 0% | 0% | 0.1%
40 | 1 | 0.116% | -0.009% | 0% | 0.1%
40 | 2 | 0.08% | -0.036% | 0.107% | 0.1%
40 | 3 | 0.111% | 0.031% | 0.102% | 0.1%

## 16. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-21,22 Active Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:09:30 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.083% | PASSED | 2023-07-28 11:36
Yes | 1 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.306% | PASSED | 2023-07-28 11:36
Yes | 2 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.187% | PASSED | 2023-07-28 11:37
Yes | 3 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.363% | PASSED | 2023-07-28 11:37
Yes | 4 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.212% | PASSED | 2023-07-28 11:38
Yes | 5 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.122% | PASSED | 2023-07-28 11:38
Yes | 6 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.089% | PASSED | 2023-07-28 11:39
Yes | 7 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.092% | PASSED | 2023-07-28 11:40
Yes | 8 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.166% | PASSED | 2023-07-28 11:40
Yes | 9 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.062% | PASSED | 2023-07-28 11:41
Yes | 10 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.264% | PASSED | 2023-07-28 11:41
Yes | 11 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.056% | PASSED | 2023-07-28 11:42
Yes | 12 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.036% | PASSED | 2023-07-28 11:43
Yes | 13 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.021% | PASSED | 2023-07-28 11:44
Yes | 14 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.181% | PASSED | 2023-07-28 11:44
Yes | 15 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.03% | PASSED | 2023-07-28 11:45
Yes | 16 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.214% | PASSED | 2023-07-28 11:45
Yes | 17 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.054% | PASSED | 2023-07-28 11:46
Yes | 18 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.059% | PASSED | 2023-07-28 11:47
Yes | 19 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.027% | PASSED | 2023-07-28 11:48
Yes | 20 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.047% | PASSED | 2023-07-28 11:48
Yes | 21 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.018% | PASSED | 2023-07-28 11:49
Yes | 22 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.018% | PASSED | 2023-07-28 11:50
Yes | 23 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.003% | PASSED | 2023-07-28 11:51
Yes | 24 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.009% | PASSED | 2023-07-28 11:52
Yes | 25 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.006% | PASSED | 2023-07-28 11:53
Yes | 26 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.018% | PASSED | 2023-07-28 11:54
Yes | 27 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.024% | PASSED | 2023-07-28 11:56
Yes | 28 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.015% | PASSED | 2023-07-28 11:57
Yes | 29 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.009% | PASSED | 2023-07-28 11:59
Yes | 30 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.021% | PASSED | 2023-07-28 12:00
Yes | 31 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.053% | PASSED | 2023-07-28 14:46
Yes | 32 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.071% | PASSED | 2023-07-28 14:47
Yes | 33 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.092% | PASSED | 2023-07-28 14:47
Yes | 34 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.044% | PASSED | 2023-07-28 14:48
Yes | 35 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.095% | PASSED | 2023-07-28 14:48
Yes | 36 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.083% | PASSED | 2023-07-28 14:49
Yes | 37 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.112% | PASSED | 2023-07-28 14:49
Yes | 38 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.103% | PASSED | 2023-07-28 14:50
Yes | 39 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.048% | PASSED | 2023-07-28 14:50
Yes | 40 | RP | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 0.5L, cos(ϕ), ϕ=60° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.118% | PASSED | 2023-07-28 14:51
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 18 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -21.83% | 0% | 0% | 0.1%
0 | 1 | -18.11% | 3.721% | 0% | 0.1%
0 | 2 | -6.07% | 12.04% | -15.34% | 0.1%
0 | 3 | -4.076% | 1.994% | -9.419% | 0.1%
0 | 4 | 0.151% | 4.227% | -3.332% | 0.1%
0 | 5 | 0.187% | 0.036% | -1.246% | 0.1%
0 | 6 | 0.053% | -0.134% | 0.13% | 0.1%
0 | 7 | 0.098% | 0.045% | 0.113% | 0.1%
0 | 8 | 0.205% | 0.107% | 0.119% | 0.1%
0 | 9 | 0.366% | 0.161% | 0.223% | 0.1%
0 | 10 | 0.384% | 0.018% | 0.318% | 0.1%
0 | 11 | 0.527% | 0.143% | 0.426% | 0.1%
0 | 12 | 0.393% | -0.134% | 0.435% | 0.1%
0 | 13 | 0.294% | -0.099% | 0.405% | 0.1%
0 | 14 | 0.16% | -0.134% | 0.282% | 0.1%
0 | 15 | 0.098% | -0.062% | 0.184% | 0.1%
0 | 16 | 0.089% | -0.009% | 0.116% | 0.1%
0 | 17 | 0.062% | -0.027% | 0.083% | 0.1%
Deviations of %error on test id #1. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -30.35% | 0% | 0% | 0.1%
1 | 1 | -19.5% | 10.85% | 0% | 0.1%
1 | 2 | -16.94% | 2.561% | -22.27% | 0.1%
1 | 3 | -3.467% | 13.48% | -13.3% | 0.1%
1 | 4 | -2.774% | 0.693% | -7.728% | 0.1%
1 | 5 | 0.178% | 2.952% | -2.021% | 0.1%
1 | 6 | 0.116% | -0.062% | -0.827% | 0.1%
1 | 7 | 0.393% | 0.277% | 0.229% | 0.1%
1 | 8 | 0.357% | -0.036% | 0.289% | 0.1%
1 | 9 | 0.285% | -0.072% | 0.345% | 0.1%
1 | 10 | 0.276% | -0.009% | 0.306% | 0.1%
Deviations of %error on test id #2. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -25.4% | 0% | 0% | 0.1%
2 | 1 | -18.34% | 7.061% | 0% | 0.1%
2 | 2 | -9.747% | 8.589% | -17.83% | 0.1%
2 | 3 | -3.151% | 6.596% | -10.41% | 0.1%
2 | 4 | 0.366% | 3.517% | -4.177% | 0.1%
2 | 5 | 0.258% | -0.108% | -0.842% | 0.1%
2 | 6 | 0.482% | 0.224% | 0.369% | 0.1%
2 | 7 | 0.321% | -0.161% | 0.354% | 0.1%
2 | 8 | 0.294% | -0.027% | 0.366% | 0.1%
2 | 9 | 0.303% | 0.009% | 0.306% | 0.1%
2 | 10 | 0.473% | 0.17% | 0.357% | 0.1%
2 | 11 | 0.285% | -0.188% | 0.354% | 0.1%
2 | 12 | 0.258% | -0.027% | 0.339% | 0.1%
2 | 13 | 0.178% | -0.08% | 0.24% | 0.1%
2 | 14 | 0.125% | -0.053% | 0.187% | 0.1%
Deviations of %error on test id #3. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -12.11% | 0% | 0% | 0.1%
3 | 1 | -10.12% | 1.994% | 0% | 0.1%
3 | 2 | 0.062% | 10.18% | -7.387% | 0.1%
3 | 3 | 0.116% | 0.054% | -3.312% | 0.1%
3 | 4 | 0.384% | 0.268% | 0.187% | 0.1%
3 | 5 | 0.357% | -0.027% | 0.286% | 0.1%
3 | 6 | 0.375% | 0.018% | 0.372% | 0.1%
3 | 7 | 0.357% | -0.018% | 0.363% | 0.1%
Deviations of %error on test id #4. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -25.52% | 0% | 0% | 0.1%
4 | 1 | -14.02% | 11.5% | 0% | 0.1%
4 | 2 | -9.689% | 4.328% | -16.41% | 0.1%
4 | 3 | -2.327% | 7.362% | -8.678% | 0.1%
4 | 4 | -1.012% | 1.315% | -4.343% | 0.1%
4 | 5 | 0.044% | 1.056% | -1.098% | 0.1%
4 | 6 | 0.285% | 0.241% | -0.228% | 0.1%
4 | 7 | 0.906% | 0.621% | 0.412% | 0.1%
4 | 8 | 0.752% | -0.154% | 0.648% | 0.1%
4 | 9 | 0.951% | 0.199% | 0.87% | 0.1%
4 | 10 | 0.815% | -0.136% | 0.839% | 0.1%
4 | 11 | 0.788% | -0.027% | 0.851% | 0.1%
4 | 12 | 0.411% | -0.377% | 0.671% | 0.1%
4 | 13 | 0.285% | -0.126% | 0.495% | 0.1%
4 | 14 | 0.178% | -0.107% | 0.291% | 0.1%
4 | 15 | 0.089% | -0.089% | 0.184% | 0.1%
4 | 16 | 0.062% | -0.027% | 0.11% | 0.1%
4 | 17 | -0.044% | -0.106% | 0.036% | 0.1%
4 | 18 | -0.389% | -0.345% | -0.124% | 0.1%
4 | 19 | -0.204% | 0.185% | -0.212% | 0.1%
Deviations of %error on test id #5. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -19.91% | 0% | 0% | 0.1%
5 | 1 | -6.273% | 13.64% | 0% | 0.1%
5 | 2 | -4.758% | 1.515% | -10.31% | 0.1%
5 | 3 | 0.214% | 4.972% | -3.606% | 0.1%
5 | 4 | 0.169% | -0.045% | -1.458% | 0.1%
5 | 5 | 0.071% | -0.098% | 0.151% | 0.1%
5 | 6 | 0.125% | 0.054% | 0.122% | 0.1%
Deviations of %error on test id #6. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | -13.22% | 0% | 0% | 0.1%
6 | 1 | 0.608% | 13.83% | 0% | 0.1%
6 | 2 | 0.518% | -0.09% | -4.032% | 0.1%
6 | 3 | 0.071% | -0.447% | 0.399% | 0.1%
6 | 4 | 1.534% | 1.463% | 0.708% | 0.1%
6 | 5 | 1.033% | -0.501% | 0.879% | 0.1%
6 | 6 | 0.071% | -0.962% | 0.879% | 0.1%
6 | 7 | -0.018% | -0.089% | 0.362% | 0.1%
6 | 8 | -0.741% | -0.723% | -0.229% | 0.1%
6 | 9 | -0.133% | 0.608% | -0.297% | 0.1%
6 | 10 | 0.089% | 0.222% | -0.262% | 0.1%
6 | 11 | 0.089% | 0% | 0.015% | 0.1%
6 | 12 | 0.08% | -0.009% | 0.086% | 0.1%
6 | 13 | 0.098% | 0.018% | 0.089% | 0.1%
Deviations of %error on test id #7. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | -25.05% | 0% | 0% | 0.1%
7 | 1 | -18.04% | 7.011% | 0% | 0.1%
7 | 2 | -8.929% | 9.109% | -17.34% | 0.1%
7 | 3 | -5.605% | 3.324% | -10.86% | 0.1%
7 | 4 | 0.125% | 5.73% | -4.803% | 0.1%
7 | 5 | 0.446% | 0.321% | -1.678% | 0.1%
7 | 6 | 0.509% | 0.063% | 0.36% | 0.1%
7 | 7 | 0.599% | 0.09% | 0.518% | 0.1%
7 | 8 | 0.464% | -0.135% | 0.524% | 0.1%
7 | 9 | 0.366% | -0.098% | 0.476% | 0.1%
7 | 10 | 0.527% | 0.161% | 0.452% | 0.1%
7 | 11 | 0.321% | -0.206% | 0.405% | 0.1%
7 | 12 | 0.285% | -0.036% | 0.378% | 0.1%
7 | 13 | 0.169% | -0.116% | 0.258% | 0.1%
7 | 14 | 0.151% | -0.018% | 0.202% | 0.1%
7 | 15 | 0.062% | -0.089% | 0.127% | 0.1%
7 | 16 | 0.062% | 0% | 0.092% | 0.1%
Deviations of %error on test id #8. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | -11.77% | 0% | 0% | 0.1%
8 | 1 | -8.722% | 3.049% | 0% | 0.1%
8 | 2 | 0.446% | 9.168% | -6.682% | 0.1%
8 | 3 | 0.473% | 0.027% | -2.601% | 0.1%
8 | 4 | 0.285% | -0.188% | 0.401% | 0.1%
8 | 5 | 0.232% | -0.053% | 0.33% | 0.1%
8 | 6 | 0.16% | -0.072% | 0.226% | 0.1%
8 | 7 | 0.107% | -0.053% | 0.166% | 0.1%
Deviations of %error on test id #9. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | -5.889% | 0% | 0% | 0.1%
9 | 1 | 0.707% | 6.596% | 0% | 0.1%
9 | 2 | 0.59% | -0.117% | -1.531% | 0.1%
9 | 3 | 0.437% | -0.153% | 0.578% | 0.1%
9 | 4 | 0.223% | -0.214% | 0.417% | 0.1%
9 | 5 | 0.125% | -0.098% | 0.262% | 0.1%
9 | 6 | 0.178% | 0.053% | 0.175% | 0.1%
9 | 7 | 0.348% | 0.17% | 0.217% | 0.1%
9 | 8 | 0.312% | -0.036% | 0.279% | 0.1%
9 | 9 | 0.196% | -0.116% | 0.285% | 0.1%
9 | 10 | 0.169% | -0.027% | 0.226% | 0.1%
9 | 11 | 0.044% | -0.125% | 0.136% | 0.1%
9 | 12 | 0.062% | 0.018% | 0.092% | 0.1%
9 | 13 | 0.053% | -0.009% | 0.053% | 0.1%
9 | 14 | 0.071% | 0.018% | 0.062% | 0.1%
Deviations of %error on test id #10. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | -1.229% | 0% | 0% | 0.1%
10 | 1 | 0.223% | 1.452% | 0% | 0.1%
10 | 2 | 0.196% | -0.027% | -0.27% | 0.1%
10 | 3 | 0.151% | -0.045% | 0.19% | 0.1%
10 | 4 | 0.027% | -0.124% | 0.125% | 0.1%
10 | 5 | 0.08% | 0.053% | 0.086% | 0.1%
10 | 6 | 0.285% | 0.205% | 0.131% | 0.1%
10 | 7 | 0.312% | 0.027% | 0.226% | 0.1%
10 | 8 | 0.276% | -0.036% | 0.291% | 0.1%
10 | 9 | 0.205% | -0.071% | 0.264% | 0.1%
Deviations of %error on test id #11. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | -0.071% | 0% | 0% | 0.1%
11 | 1 | 0.205% | 0.276% | 0% | 0.1%
11 | 2 | 0.258% | 0.053% | 0.131% | 0.1%
11 | 3 | 0.402% | 0.144% | 0.288% | 0.1%
11 | 4 | 0.285% | -0.117% | 0.315% | 0.1%
11 | 5 | 0.348% | 0.063% | 0.345% | 0.1%
11 | 6 | 0.527% | 0.179% | 0.387% | 0.1%
11 | 7 | 0.258% | -0.269% | 0.378% | 0.1%
11 | 8 | 0.232% | -0.026% | 0.339% | 0.1%
11 | 9 | 0.196% | -0.036% | 0.229% | 0.1%
11 | 10 | 0.08% | -0.116% | 0.169% | 0.1%
11 | 11 | 0.036% | -0.044% | 0.104% | 0.1%
11 | 12 | 0.044% | 0.008% | 0.053% | 0.1%
11 | 13 | 0.089% | 0.045% | 0.056% | 0.1%
Deviations of %error on test id #12. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | -1.454% | 0% | 0% | 0.1%
12 | 1 | 0.241% | 1.695% | 0% | 0.1%
12 | 2 | 0.214% | -0.027% | -0.333% | 0.1%
12 | 3 | 0.375% | 0.161% | 0.277% | 0.1%
12 | 4 | 0.267% | -0.108% | 0.285% | 0.1%
12 | 5 | 0.214% | -0.053% | 0.285% | 0.1%
12 | 6 | 0.276% | 0.062% | 0.252% | 0.1%
12 | 7 | 0.393% | 0.117% | 0.294% | 0.1%
12 | 8 | 0.285% | -0.108% | 0.318% | 0.1%
12 | 9 | 0.16% | -0.125% | 0.279% | 0.1%
12 | 10 | 0.169% | 0.009% | 0.205% | 0.1%
12 | 11 | 0.062% | -0.107% | 0.13% | 0.1%
12 | 12 | 0.027% | -0.035% | 0.086% | 0.1%
12 | 13 | 0.053% | 0.026% | 0.047% | 0.1%
12 | 14 | 0.027% | -0.026% | 0.036% | 0.1%
Deviations of %error on test id #13. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | -5.223% | 0% | 0% | 0.1%
13 | 1 | 0.545% | 5.768% | 0% | 0.1%
13 | 2 | 0.053% | -0.492% | -1.542% | 0.1%
13 | 3 | 0.116% | 0.063% | 0.238% | 0.1%
13 | 4 | 1.434% | 1.318% | 0.534% | 0.1%
13 | 5 | 0.375% | -1.059% | 0.642% | 0.1%
13 | 6 | 0.018% | -0.357% | 0.609% | 0.1%
13 | 7 | -0.407% | -0.425% | -0.005% | 0.1%
13 | 8 | -0.416% | -0.009% | -0.268% | 0.1%
13 | 9 | 0.036% | 0.452% | -0.262% | 0.1%
13 | 10 | 0.018% | -0.018% | -0.121% | 0.1%
13 | 11 | 0.027% | 0.009% | 0.027% | 0.1%
13 | 12 | 0.018% | -0.009% | 0.021% | 0.1%
Deviations of %error on test id #14. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | 0.294% | 0% | 0% | 0.1%
14 | 1 | 0.151% | -0.143% | 0% | 0.1%
14 | 2 | 0.169% | 0.018% | 0.205% | 0.1%
14 | 3 | 0.33% | 0.161% | 0.217% | 0.1%
14 | 4 | 0.276% | -0.054% | 0.258% | 0.1%
14 | 5 | 0.276% | 0% | 0.294% | 0.1%
14 | 6 | 0.178% | -0.098% | 0.243% | 0.1%
14 | 7 | 0.089% | -0.089% | 0.181% | 0.1%
Deviations of %error on test id #15. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | 0.312% | 0% | 0% | 0.1%
15 | 1 | 0.205% | -0.107% | 0% | 0.1%
15 | 2 | 0.169% | -0.036% | 0.229% | 0.1%
15 | 3 | 0.044% | -0.125% | 0.139% | 0.1%
15 | 4 | 0.053% | 0.009% | 0.089% | 0.1%
15 | 5 | 0.027% | -0.026% | 0.041% | 0.1%
15 | 6 | 0.009% | -0.018% | 0.03% | 0.1%
Deviations of %error on test id #16. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | 0.214% | 0% | 0% | 0.1%
16 | 1 | 0.339% | 0.125% | 0% | 0.1%
16 | 2 | 0.294% | -0.045% | 0.282% | 0.1%
16 | 3 | 0.205% | -0.089% | 0.279% | 0.1%
16 | 4 | 0.142% | -0.063% | 0.214% | 0.1%
Deviations of %error on test id #17. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | 0.142% | 0% | 0% | 0.1%
17 | 1 | 0.089% | -0.053% | 0% | 0.1%
17 | 2 | 0.036% | -0.053% | 0.089% | 0.1%
17 | 3 | 0.036% | 0% | 0.054% | 0.1%
Deviations of %error on test id #18. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | 0.16% | 0% | 0% | 0.1%
18 | 1 | 0.08% | -0.08% | 0% | 0.1%
18 | 2 | 0.071% | -0.009% | 0.104% | 0.1%
18 | 3 | 0.027% | -0.044% | 0.059% | 0.1%
Deviations of %error on test id #19. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | 0.053% | 0% | 0% | 0.1%
19 | 1 | 0.08% | 0.027% | 0% | 0.1%
19 | 2 | -0.009% | -0.089% | 0.041% | 0.1%
19 | 3 | 0.009% | 0.018% | 0.027% | 0.1%
Deviations of %error on test id #20. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | 0.062% | 0% | 0% | 0.1%
20 | 1 | 0.053% | -0.009% | 0% | 0.1%
20 | 2 | 0.036% | -0.017% | 0.05% | 0.1%
20 | 3 | 0.062% | 0.026% | 0.05% | 0.1%
20 | 4 | 0.044% | -0.018% | 0.047% | 0.1%
Deviations of %error on test id #21. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | -0.009% | 0% | 0% | 0.1%
21 | 1 | 0.027% | 0.036% | 0% | 0.1%
21 | 2 | 0.018% | -0.009% | 0.012% | 0.1%
21 | 3 | 0.009% | -0.009% | 0.018% | 0.1%
Deviations of %error on test id #22. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | 0.018% | 0% | 0% | 0.1%
22 | 1 | 0.009% | -0.009% | 0% | 0.1%
22 | 2 | 0.018% | 0.009% | 0.015% | 0.1%
22 | 3 | -0.027% | -0.045% | 0% | 0.1%
22 | 4 | 0.062% | 0.089% | 0.018% | 0.1%
Deviations of %error on test id #23. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | 0.009% | 0% | 0% | 0.1%
23 | 1 | 0.009% | 0% | 0% | 0.1%
23 | 2 | -0.027% | -0.036% | -0.003% | 0.1%
23 | 3 | 0.009% | 0.036% | -0.003% | 0.1%
Deviations of %error on test id #24. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | -0.018% | 0% | 0% | 0.1%
24 | 1 | 0.018% | 0.036% | 0% | 0.1%
24 | 2 | 0.018% | 0% | 0.006% | 0.1%
24 | 3 | -0.062% | -0.08% | -0.009% | 0.1%
Deviations of %error on test id #25. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | -0.036% | 0% | 0% | 0.1%
25 | 1 | 0.027% | 0.063% | 0% | 0.1%
25 | 2 | 0.009% | -0.018% | 0% | 0.1%
25 | 3 | -0.018% | -0.027% | 0.006% | 0.1%
Deviations of %error on test id #26. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | -0.053% | 0% | 0% | 0.1%
26 | 1 | 0.018% | 0.071% | 0% | 0.1%
26 | 2 | 0.036% | 0.018% | 0% | 0.1%
26 | 3 | 0% | -0.036% | 0.018% | 0.1%
Deviations of %error on test id #27. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | -0.009% | 0% | 0% | 0.1%
27 | 1 | 0.018% | 0.027% | 0% | 0.1%
27 | 2 | -0.062% | -0.08% | -0.018% | 0.1%
27 | 3 | -0.027% | 0.035% | -0.024% | 0.1%
Deviations of %error on test id #28. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | 0.036% | 0% | 0% | 0.1%
28 | 1 | -0.053% | -0.089% | 0% | 0.1%
28 | 2 | 0.036% | 0.089% | 0.006% | 0.1%
28 | 3 | -0.027% | -0.063% | -0.015% | 0.1%
Deviations of %error on test id #29. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | -0.018% | 0% | 0% | 0.1%
29 | 1 | -0.044% | -0.026% | 0% | 0.1%
29 | 2 | 0.027% | 0.071% | -0.012% | 0.1%
29 | 3 | -0.009% | -0.036% | -0.009% | 0.1%
Deviations of %error on test id #30. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | -0.053% | 0% | 0% | 0.1%
30 | 1 | 0% | 0.053% | 0% | 0.1%
30 | 2 | 0% | 0% | -0.018% | 0.1%
30 | 3 | -0.053% | -0.053% | -0.018% | 0.1%
30 | 4 | -0.009% | 0.044% | -0.021% | 0.1%
Deviations of %error on test id #31. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | -18.45% | 0% | 0% | 0.1%
31 | 1 | -7.522% | 10.93% | 0% | 0.1%
31 | 2 | -3.508% | 4.014% | -9.828% | 0.1%
31 | 3 | 0.348% | 3.856% | -3.561% | 0.1%
31 | 4 | 0.214% | -0.134% | -0.982% | 0.1%
31 | 5 | 0.151% | -0.063% | 0.238% | 0.1%
31 | 6 | -0.009% | -0.16% | 0.119% | 0.1%
31 | 7 | 0.178% | 0.187% | 0.107% | 0.1%
31 | 8 | 0.241% | 0.063% | 0.137% | 0.1%
31 | 9 | 0.134% | -0.107% | 0.184% | 0.1%
31 | 10 | 0.071% | -0.063% | 0.149% | 0.1%
31 | 11 | 0.393% | 0.322% | 0.199% | 0.1%
31 | 12 | 1.379% | 0.986% | 0.614% | 0.1%
31 | 13 | 1.006% | -0.373% | 0.926% | 0.1%
31 | 14 | 0.725% | -0.281% | 1.037% | 0.1%
31 | 15 | 0.42% | -0.305% | 0.717% | 0.1%
31 | 16 | -0.036% | -0.456% | 0.37% | 0.1%
31 | 17 | -0.044% | -0.008% | 0.113% | 0.1%
31 | 18 | -0.08% | -0.036% | -0.053% | 0.1%
31 | 19 | -0.036% | 0.044% | -0.053% | 0.1%
Deviations of %error on test id #32. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | -21.69% | 0% | 0% | 0.1%
32 | 1 | -13% | 8.691% | 0% | 0.1%
32 | 2 | -3.276% | 9.723% | -12.65% | 0.1%
32 | 3 | -1.419% | 1.857% | -5.898% | 0.1%
32 | 4 | 0.312% | 1.731% | -1.461% | 0.1%
32 | 5 | 0.384% | 0.072% | -0.241% | 0.1%
32 | 6 | 0.366% | -0.018% | 0.354% | 0.1%
32 | 7 | 0.196% | -0.17% | 0.315% | 0.1%
32 | 8 | 0.348% | 0.152% | 0.303% | 0.1%
32 | 9 | 0.294% | -0.054% | 0.279% | 0.1%
32 | 10 | 0.205% | -0.089% | 0.282% | 0.1%
32 | 11 | 0.071% | -0.134% | 0.19% | 0.1%
32 | 12 | -0.018% | -0.089% | 0.086% | 0.1%
32 | 13 | -0.089% | -0.071% | -0.012% | 0.1%
32 | 14 | -0.106% | -0.017% | -0.071% | 0.1%
Deviations of %error on test id #33. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | -8.522% | 0% | 0% | 0.1%
33 | 1 | 0.107% | 8.629% | 0% | 0.1%
33 | 2 | 0.635% | 0.528% | -2.593% | 0.1%
33 | 3 | 0% | -0.635% | 0.247% | 0.1%
33 | 4 | -0.08% | -0.08% | 0.185% | 0.1%
33 | 5 | -0.089% | -0.009% | -0.056% | 0.1%
33 | 6 | -0.106% | -0.017% | -0.092% | 0.1%
Deviations of %error on test id #34. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | -13.02% | 0% | 0% | 0.1%
34 | 1 | -3.217% | 9.802% | 0% | 0.1%
34 | 2 | -1.333% | 1.884% | -5.856% | 0.1%
34 | 3 | 0.232% | 1.565% | -1.439% | 0.1%
34 | 4 | 0.276% | 0.044% | -0.275% | 0.1%
34 | 5 | 0.33% | 0.054% | 0.279% | 0.1%
34 | 6 | 0.187% | -0.143% | 0.264% | 0.1%
34 | 7 | 0.321% | 0.134% | 0.279% | 0.1%
34 | 8 | 0.223% | -0.098% | 0.244% | 0.1%
34 | 9 | 0.151% | -0.072% | 0.232% | 0.1%
34 | 10 | 0.071% | -0.08% | 0.148% | 0.1%
34 | 11 | -0.089% | -0.16% | 0.044% | 0.1%
Deviations of %error on test id #35. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | -6.878% | 0% | 0% | 0.1%
35 | 1 | 0.223% | 7.101% | 0% | 0.1%
35 | 2 | 0.116% | -0.107% | -2.18% | 0.1%
35 | 3 | 0.036% | -0.08% | 0.125% | 0.1%
35 | 4 | 0.285% | 0.249% | 0.146% | 0.1%
35 | 5 | 0.428% | 0.143% | 0.25% | 0.1%
35 | 6 | 0.446% | 0.018% | 0.386% | 0.1%
35 | 7 | 0.285% | -0.161% | 0.386% | 0.1%
35 | 8 | 0.044% | -0.241% | 0.258% | 0.1%
35 | 9 | -0.044% | -0.088% | 0.095% | 0.1%
35 | 10 | -0.098% | -0.054% | -0.033% | 0.1%
35 | 11 | -0.142% | -0.044% | -0.095% | 0.1%
Deviations of %error on test id #36. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | -9.726% | 0% | 0% | 0.1%
36 | 1 | 0.08% | 9.806% | 0% | 0.1%
36 | 2 | -0.036% | -0.116% | -3.227% | 0.1%
36 | 3 | -0.098% | -0.062% | -0.018% | 0.1%
36 | 4 | 0.107% | 0.205% | -0.009% | 0.1%
36 | 5 | 0.25% | 0.143% | 0.086% | 0.1%
36 | 6 | 0.482% | 0.232% | 0.28% | 0.1%
36 | 7 | 0.312% | -0.17% | 0.348% | 0.1%
36 | 8 | 0.08% | -0.232% | 0.291% | 0.1%
36 | 9 | 0.018% | -0.062% | 0.137% | 0.1%
36 | 10 | -0.115% | -0.133% | -0.006% | 0.1%
36 | 11 | -0.071% | 0.044% | -0.056% | 0.1%
36 | 12 | -0.106% | -0.035% | -0.097% | 0.1%
36 | 13 | -0.071% | 0.035% | -0.083% | 0.1%
Deviations of %error on test id #37. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | -9.921% | 0% | 0% | 0.1%
37 | 1 | -2.799% | 7.122% | 0% | 0.1%
37 | 2 | 0.169% | 2.968% | -4.184% | 0.1%
37 | 3 | 0.089% | -0.08% | -0.847% | 0.1%
37 | 4 | 0.036% | -0.053% | 0.098% | 0.1%
37 | 5 | 0.214% | 0.178% | 0.113% | 0.1%
37 | 6 | 0.134% | -0.08% | 0.128% | 0.1%
37 | 7 | 0.027% | -0.107% | 0.125% | 0.1%
37 | 8 | 0.009% | -0.018% | 0.057% | 0.1%
37 | 9 | 0.644% | 0.635% | 0.227% | 0.1%
37 | 10 | 0.482% | -0.162% | 0.378% | 0.1%
37 | 11 | 0.267% | -0.215% | 0.464% | 0.1%
37 | 12 | 0.125% | -0.142% | 0.291% | 0.1%
37 | 13 | -0.098% | -0.223% | 0.098% | 0.1%
37 | 14 | -0.115% | -0.017% | -0.029% | 0.1%
37 | 15 | -0.124% | -0.009% | -0.112% | 0.1%
37 | 16 | -0.098% | 0.026% | -0.112% | 0.1%
Deviations of %error on test id #38. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | -8.566% | 0% | 0% | 0.1%
38 | 1 | -1.601% | 6.965% | 0% | 0.1%
38 | 2 | 0.027% | 1.628% | -3.38% | 0.1%
38 | 3 | 0.276% | 0.249% | -0.433% | 0.1%
38 | 4 | 0.196% | -0.08% | 0.166% | 0.1%
38 | 5 | 0.116% | -0.08% | 0.196% | 0.1%
38 | 6 | 0.339% | 0.223% | 0.217% | 0.1%
38 | 7 | 0.142% | -0.197% | 0.199% | 0.1%
38 | 8 | 0.071% | -0.071% | 0.184% | 0.1%
38 | 9 | 0.036% | -0.035% | 0.083% | 0.1%
38 | 10 | -0.133% | -0.169% | -0.009% | 0.1%
38 | 11 | -0.071% | 0.062% | -0.056% | 0.1%
38 | 12 | -0.115% | -0.044% | -0.106% | 0.1%
38 | 13 | -0.089% | 0.026% | -0.092% | 0.1%
38 | 14 | -0.106% | -0.017% | -0.103% | 0.1%
Deviations of %error on test id #39. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | 0.16% | 0% | 0% | 0.1%
39 | 1 | 0.098% | -0.062% | 0% | 0.1%
39 | 2 | 0.018% | -0.08% | 0.092% | 0.1%
39 | 3 | 0.027% | 0.009% | 0.048% | 0.1%
Deviations of %error on test id #40. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | 0.009% | 0% | 0% | 0.1%
40 | 1 | 0.303% | 0.294% | 0% | 0.1%
40 | 2 | 0.196% | -0.107% | 0.169% | 0.1%
40 | 3 | 0.089% | -0.107% | 0.196% | 0.1%
40 | 4 | 0.312% | 0.223% | 0.199% | 0.1%
40 | 5 | 0.169% | -0.143% | 0.19% | 0.1%
40 | 6 | 0.089% | -0.08% | 0.19% | 0.1%
40 | 7 | 0.027% | -0.062% | 0.095% | 0.1%
40 | 8 | -0.133% | -0.16% | -0.006% | 0.1%
40 | 9 | -0.098% | 0.035% | -0.068% | 0.1%
40 | 10 | -0.098% | 0% | -0.11% | 0.1%
40 | 11 | -0.151% | -0.053% | -0.116% | 0.1%
40 | 12 | -0.106% | 0.045% | -0.118% | 0.1%

## 17. PF = 1.0
[<<]
Accuracy Test Report of 'PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:10:01 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 20 | 20 |  | ±0.1% | ±0.5% | 0.05% | PASSED | 2023-07-27 16:07
Yes | 1 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 12 | 20 |  | ±0.1% | ±0.5% | 0.039% | PASSED | 2023-07-27 16:07
Yes | 3 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.5% | 0.05% | PASSED | 2023-07-27 16:08
Yes | 4 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0.089% | PASSED | 2023-07-27 16:08
Yes | 7 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0.091% | PASSED | 2023-07-27 16:09
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 0.036% | 0% | 0% | 0.1%
0 | 1 | 0.071% | 0.035% | 0% | 0.1%
0 | 2 | 0.044% | -0.027% | 0.05% | 0.1%
0 | 3 | 0.036% | -0.008% | 0.05% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -0.036% | 0% | 0% | 0.1%
1 | 1 | 0.03% | 0.066% | 0% | 0.1%
1 | 2 | 0.033% | 0.003% | 0.009% | 0.1%
1 | 3 | 0.053% | 0.02% | 0.039% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.022% | 0% | 0% | 0.1%
3 | 1 | 0.044% | 0.022% | 0% | 0.1%
3 | 2 | 0.053% | 0.009% | 0.04% | 0.1%
3 | 3 | 0.053% | 0% | 0.05% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.009% | 0% | 0% | 0.1%
4 | 1 | 0.08% | 0.071% | 0% | 0.1%
4 | 2 | 0.089% | 0.009% | 0.059% | 0.1%
4 | 3 | 0.098% | 0.009% | 0.089% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.087% | 0% | 0% | 0.1%
7 | 1 | 0.085% | -0.002% | 0% | 0.1%
7 | 2 | 0.098% | 0.013% | 0.09% | 0.1%
7 | 3 | 0.089% | -0.009% | 0.091% | 0.1%

## 18. PF = 0.5L
[<<]
Accuracy Test Report of 'PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:10:04 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | 0.315% | PASSED | 2023-07-27 16:09
Yes | 1 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | 0.309% | PASSED | 2023-07-27 16:10
Yes | 3 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.273% | PASSED | 2023-07-27 16:10
Yes | 4 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.232% | PASSED | 2023-07-27 16:11
Yes | 7 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.236% | PASSED | 2023-07-27 16:12
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -31.16% | 0% | 0% | 0.1%
0 | 1 | -15.96% | 15.2% | 0% | 0.1%
0 | 2 | -4.141% | 11.82% | -17.09% | 0.1%
0 | 3 | 0.339% | 4.48% | -6.586% | 0.1%
0 | 4 | 0.107% | -0.232% | -1.232% | 0.1%
0 | 5 | -0.301% | -0.408% | 0.048% | 0.1%
0 | 6 | -0.601% | -0.3% | -0.265% | 0.1%
0 | 7 | -0.337% | 0.264% | -0.413% | 0.1%
0 | 8 | 0.018% | 0.355% | -0.307% | 0.1%
0 | 9 | 0.232% | 0.214% | -0.029% | 0.1%
0 | 10 | 0.321% | 0.089% | 0.19% | 0.1%
0 | 11 | 0.303% | -0.018% | 0.285% | 0.1%
0 | 12 | 0.321% | 0.018% | 0.315% | 0.1%
Deviations of %error on test id #1. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -33.92% | 0% | 0% | 0.1%
1 | 1 | -15.52% | 18.4% | 0% | 0.1%
1 | 2 | -0.307% | 15.21% | -16.58% | 0.1%
1 | 3 | 0.333% | 0.64% | -5.164% | 0.1%
1 | 4 | 0.184% | -0.149% | 0.07% | 0.1%
1 | 5 | -0.124% | -0.308% | 0.131% | 0.1%
1 | 6 | -0.089% | 0.035% | -0.01% | 0.1%
1 | 7 | -0.03% | 0.059% | -0.081% | 0.1%
1 | 8 | -0.295% | -0.265% | -0.138% | 0.1%
1 | 9 | 0.047% | 0.342% | -0.093% | 0.1%
1 | 10 | 0.25% | 0.203% | 0.001% | 0.1%
1 | 11 | 0.297% | 0.047% | 0.198% | 0.1%
1 | 12 | 0.315% | 0.018% | 0.287% | 0.1%
1 | 13 | 0.315% | 0% | 0.309% | 0.1%
Deviations of %error on test id #3. Total 18 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -33.82% | 0% | 0% | 0.1%
3 | 1 | -18.54% | 15.28% | 0% | 0.1%
3 | 2 | -8.715% | 9.828% | -20.36% | 0.1%
3 | 3 | -2.055% | 6.66% | -9.771% | 0.1%
3 | 4 | 0.25% | 2.305% | -3.507% | 0.1%
3 | 5 | 0.25% | 0% | -0.518% | 0.1%
3 | 6 | 0.258% | 0.008% | 0.253% | 0.1%
3 | 7 | 0.134% | -0.124% | 0.214% | 0.1%
3 | 8 | -0.151% | -0.285% | 0.08% | 0.1%
3 | 9 | -0.204% | -0.053% | -0.074% | 0.1%
3 | 10 | -0.248% | -0.044% | -0.201% | 0.1%
3 | 11 | -0.469% | -0.221% | -0.307% | 0.1%
3 | 12 | -0.284% | 0.185% | -0.334% | 0.1%
3 | 13 | -0.053% | 0.231% | -0.269% | 0.1%
3 | 14 | 0.196% | 0.249% | -0.047% | 0.1%
3 | 15 | 0.223% | 0.027% | 0.122% | 0.1%
3 | 16 | 0.258% | 0.035% | 0.226% | 0.1%
3 | 17 | 0.339% | 0.081% | 0.273% | 0.1%
Deviations of %error on test id #4. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -32.59% | 0% | 0% | 0.1%
4 | 1 | -7.506% | 25.08% | 0% | 0.1%
4 | 2 | 0.214% | 7.72% | -13.29% | 0.1%
4 | 3 | -0.009% | -0.223% | -2.434% | 0.1%
4 | 4 | -0.106% | -0.097% | 0.033% | 0.1%
4 | 5 | 0.018% | 0.124% | -0.032% | 0.1%
4 | 6 | -0.089% | -0.107% | -0.059% | 0.1%
4 | 7 | -0.186% | -0.097% | -0.086% | 0.1%
4 | 8 | 0.134% | 0.32% | -0.047% | 0.1%
4 | 9 | 0.214% | 0.08% | 0.054% | 0.1%
4 | 10 | 0.214% | 0% | 0.187% | 0.1%
4 | 11 | 0.366% | 0.152% | 0.265% | 0.1%
4 | 12 | 0.437% | 0.071% | 0.339% | 0.1%
4 | 13 | 0.276% | -0.161% | 0.36% | 0.1%
4 | 14 | 0.223% | -0.053% | 0.312% | 0.1%
4 | 15 | 0.232% | 0.009% | 0.244% | 0.1%
4 | 16 | 0.241% | 0.009% | 0.232% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.244% | 0% | 0% | 0.1%
7 | 1 | 0.212% | -0.032% | 0% | 0.1%
7 | 2 | 0.253% | 0.041% | 0.236% | 0.1%
7 | 3 | 0.242% | -0.011% | 0.236% | 0.1%

## 19. PF = 0.8C
[<<]
Accuracy Test Report of 'PF = 0.8C'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / PF = 0.8C'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:10:09 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | -0.024% | PASSED | 2023-07-27 16:13
Yes | 1 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | 0.159% | PASSED | 2023-07-27 16:13
Yes | 3 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.062% | PASSED | 2023-07-27 16:13
Yes | 4 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.012% | PASSED | 2023-07-27 16:14
Yes | 7 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.011% | PASSED | 2023-07-27 16:15
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -21.33% | 0% | 0% | 0.1%
0 | 1 | 294.5% | 315.8% | 0% | 0.1%
0 | 2 | 1285% | 991% | 519.5% | 0.1%
0 | 3 | 53.6% | -1232% | 544.5% | 0.1%
0 | 4 | 37.13% | -16.48% | 458.7% | 0.1%
0 | 5 | -0.036% | -37.16% | 30.23% | 0.1%
0 | 6 | -0.053% | -0.017% | 12.35% | 0.1%
0 | 7 | 0.036% | 0.089% | -0.018% | 0.1%
0 | 8 | 0.285% | 0.249% | 0.089% | 0.1%
0 | 9 | 0.16% | -0.125% | 0.16% | 0.1%
0 | 10 | 0.196% | 0.036% | 0.214% | 0.1%
0 | 11 | 0.16% | -0.036% | 0.172% | 0.1%
0 | 12 | 0.59% | 0.43% | 0.315% | 0.1%
0 | 13 | 0.375% | -0.215% | 0.375% | 0.1%
0 | 14 | 0.303% | -0.072% | 0.423% | 0.1%
0 | 15 | 0.178% | -0.125% | 0.285% | 0.1%
0 | 16 | 0% | -0.178% | 0.16% | 0.1%
0 | 17 | 0% | 0% | 0.059% | 0.1%
0 | 18 | -0.036% | -0.036% | -0.012% | 0.1%
0 | 19 | -0.036% | 0% | -0.024% | 0.1%
Deviations of %error on test id #1. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 176.1% | 0% | 0% | 0.1%
1 | 1 | 160.7% | -15.31% | 0% | 0.1%
1 | 2 | 89.59% | -71.15% | 142.1% | 0.1%
1 | 3 | 6.926% | -82.66% | 85.75% | 0.1%
1 | 4 | 1.577% | -5.349% | 32.7% | 0.1%
1 | 5 | -0.106% | -1.683% | 2.799% | 0.1%
1 | 6 | -0.083% | 0.023% | 0.463% | 0.1%
1 | 7 | -0.065% | 0.018% | -0.085% | 0.1%
1 | 8 | 0.041% | 0.106% | -0.036% | 0.1%
1 | 9 | 0.125% | 0.084% | 0.034% | 0.1%
1 | 10 | 0.131% | 0.006% | 0.099% | 0.1%
1 | 11 | 0.22% | 0.089% | 0.159% | 0.1%
Deviations of %error on test id #3. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 80.61% | 0% | 0% | 0.1%
3 | 1 | 197.5% | 116.9% | 0% | 0.1%
3 | 2 | 339.1% | 141.6% | 205.7% | 0.1%
3 | 3 | 121.2% | -217.9% | 219.3% | 0.1%
3 | 4 | 26.12% | -95.07% | 162.1% | 0.1%
3 | 5 | 21.39% | -4.736% | 56.23% | 0.1%
3 | 6 | 0.018% | -21.37% | 15.84% | 0.1%
3 | 7 | 0.071% | 0.053% | 7.158% | 0.1%
3 | 8 | 0.044% | -0.027% | 0.044% | 0.1%
3 | 9 | 0.018% | -0.026% | 0.044% | 0.1%
3 | 10 | 0.125% | 0.107% | 0.062% | 0.1%
Deviations of %error on test id #4. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 437% | 0% | 0% | 0.1%
4 | 1 | 103.4% | -333.6% | 0% | 0.1%
4 | 2 | 16.3% | -87.06% | 185.6% | 0.1%
4 | 3 | 0.009% | -16.29% | 39.89% | 0.1%
4 | 4 | 0.018% | 0.009% | 5.443% | 0.1%
4 | 5 | 0.044% | 0.026% | 0.024% | 0.1%
4 | 6 | 0.16% | 0.116% | 0.074% | 0.1%
4 | 7 | 0.169% | 0.009% | 0.124% | 0.1%
4 | 8 | 0.196% | 0.027% | 0.175% | 0.1%
4 | 9 | 0.357% | 0.161% | 0.241% | 0.1%
4 | 10 | 0.223% | -0.134% | 0.259% | 0.1%
4 | 11 | 0.098% | -0.125% | 0.226% | 0.1%
4 | 12 | 0.009% | -0.089% | 0.11% | 0.1%
4 | 13 | 0.044% | 0.035% | 0.05% | 0.1%
4 | 14 | -0.018% | -0.062% | 0.012% | 0.1%
Deviations of %error on test id #7. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.16% | 0% | 0% | 0.1%
7 | 1 | -0.039% | -0.199% | 0% | 0.1%
7 | 2 | 0.02% | 0.059% | 0.047% | 0.1%
7 | 3 | 0.014% | -0.006% | -0.002% | 0.1%
7 | 4 | 0% | -0.014% | 0.011% | 0.1%

## 20. 5th Harmonics
[<<]
Accuracy Test Report of '5th Harmonics'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / 5th Harmonics'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:10:13 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Harmonics' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
1 | 0 | 0 | 1
3. Test Cases
Total 1 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 1 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° | {UcUbUa, 5th-10%, IcIbIa, 5th-40%} | 12 | 20 |  | ±0.1% | ±1% | 0.031% | PASSED | 2023-07-27 16:15
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #1. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -0.236% | 0% | 0% | 0.1%
1 | 1 | -0.151% | 0.085% | 0% | 0.1%
1 | 2 | -1.278% | -1.127% | -0.555% | 0.1%
1 | 3 | -4.426% | -3.148% | -1.952% | 0.1%
1 | 4 | -3.433% | 0.993% | -3.046% | 0.1%
1 | 5 | -0.942% | 2.491% | -2.934% | 0.1%
1 | 6 | 0.003% | 0.945% | -1.457% | 0.1%
1 | 7 | 0.288% | 0.285% | -0.217% | 0.1%
1 | 8 | 0.396% | 0.108% | 0.229% | 0.1%
1 | 9 | 0.184% | -0.212% | 0.289% | 0.1%
1 | 10 | 0.047% | -0.137% | 0.209% | 0.1%
1 | 11 | 0.039% | -0.008% | 0.09% | 0.1%
1 | 12 | 0.03% | -0.009% | 0.039% | 0.1%
1 | 13 | 0.024% | -0.006% | 0.031% | 0.1%

## 21. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:10:15 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DQ | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | 0.052% | PASSED | 2023-07-27 16:16
Yes | 1 | DQ | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.08% | PASSED | 2023-07-27 16:16
Yes | 2 | DQ | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.069% | PASSED | 2023-07-27 16:16
Yes | 3 | DQ | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.07% | PASSED | 2023-07-27 16:17
Yes | 4 | DQ | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.071% | PASSED | 2023-07-27 16:17
Yes | 5 | DQ | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | 0.073% | PASSED | 2023-07-27 16:17
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 3.907% | 0% | 0% | 0.1%
0 | 1 | 1.722% | -2.185% | 0% | 0.1%
0 | 2 | 0.397% | -1.325% | 2.009% | 0.1%
0 | 3 | 0.089% | -0.308% | 0.736% | 0.1%
0 | 4 | 0.027% | -0.062% | 0.171% | 0.1%
0 | 5 | 0.08% | 0.053% | 0.065% | 0.1%
0 | 6 | 0.049% | -0.031% | 0.052% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0.071% | 0% | 0% | 0.1%
1 | 1 | 0.076% | 0.005% | 0% | 0.1%
1 | 2 | 0.085% | 0.009% | 0.077% | 0.1%
1 | 3 | 0.08% | -0.005% | 0.08% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | 0.089% | 0% | 0% | 0.1%
2 | 1 | 0.071% | -0.018% | 0% | 0.1%
2 | 2 | 0.044% | -0.027% | 0.068% | 0.1%
2 | 3 | 0.093% | 0.049% | 0.069% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.053% | 0% | 0% | 0.1%
3 | 1 | 0.058% | 0.005% | 0% | 0.1%
3 | 2 | 0.102% | 0.044% | 0.071% | 0.1%
3 | 3 | 0.049% | -0.053% | 0.07% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.08% | 0% | 0% | 0.1%
4 | 1 | 0.044% | -0.036% | 0% | 0.1%
4 | 2 | 0.093% | 0.049% | 0.072% | 0.1%
4 | 3 | 0.076% | -0.017% | 0.071% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | 0.071% | 0% | 0% | 0.1%
5 | 1 | 0.076% | 0.005% | 0% | 0.1%
5 | 2 | 0.062% | -0.014% | 0.07% | 0.1%
5 | 3 | 0.08% | 0.018% | 0.073% | 0.1%

## 22. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Voltage Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:10:18 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DQ | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | 0.193% | PASSED | 2023-07-27 16:18
Yes | 1 | DQ | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.247% | PASSED | 2023-07-27 16:18
Yes | 2 | DQ | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.264% | PASSED | 2023-07-27 16:19
Yes | 3 | DQ | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.238% | PASSED | 2023-07-27 16:19
Yes | 4 | DQ | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | -0.036% | PASSED | 2023-07-27 16:20
Yes | 5 | DQ | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | -0.086% | PASSED | 2023-07-27 16:20
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -31.95% | 0% | 0% | 0.1%
0 | 1 | -18.19% | 13.76% | 0% | 0.1%
0 | 2 | -5.676% | 12.52% | -18.61% | 0.1%
0 | 3 | 0.241% | 5.917% | -7.876% | 0.1%
0 | 4 | 0.196% | -0.045% | -1.746% | 0.1%
0 | 5 | 0.223% | 0.027% | 0.22% | 0.1%
0 | 6 | 0.16% | -0.063% | 0.193% | 0.1%
Deviations of %error on test id #1. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -34.27% | 0% | 0% | 0.1%
1 | 1 | -19.57% | 14.7% | 0% | 0.1%
1 | 2 | -8.804% | 10.77% | -20.88% | 0.1%
1 | 3 | -1.549% | 7.255% | -9.975% | 0.1%
1 | 4 | 0.214% | 1.763% | -3.38% | 0.1%
1 | 5 | 0.232% | 0.018% | -0.368% | 0.1%
1 | 6 | 0.071% | -0.161% | 0.172% | 0.1%
1 | 7 | 0.018% | -0.053% | 0.107% | 0.1%
1 | 8 | -0.106% | -0.124% | -0.006% | 0.1%
1 | 9 | -0.416% | -0.31% | -0.168% | 0.1%
1 | 10 | -0.522% | -0.106% | -0.348% | 0.1%
1 | 11 | -0.266% | 0.256% | -0.401% | 0.1%
1 | 12 | -0.009% | 0.257% | -0.266% | 0.1%
1 | 13 | 0.196% | 0.205% | -0.026% | 0.1%
1 | 14 | 0.267% | 0.071% | 0.151% | 0.1%
1 | 15 | 0.232% | -0.035% | 0.232% | 0.1%
1 | 16 | 0.241% | 0.009% | 0.247% | 0.1%
Deviations of %error on test id #2. Total 18 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -34.63% | 0% | 0% | 0.1%
2 | 1 | -24.09% | 10.54% | 0% | 0.1%
2 | 2 | -13.34% | 10.75% | -24.02% | 0.1%
2 | 3 | -2.411% | 10.93% | -13.28% | 0.1%
2 | 4 | 0.098% | 2.509% | -5.218% | 0.1%
2 | 5 | 0.25% | 0.152% | -0.688% | 0.1%
2 | 6 | 0.16% | -0.09% | 0.169% | 0.1%
2 | 7 | -0.115% | -0.275% | 0.098% | 0.1%
2 | 8 | 0.027% | 0.142% | 0.024% | 0.1%
2 | 9 | -0.169% | -0.196% | -0.086% | 0.1%
2 | 10 | -0.495% | -0.326% | -0.212% | 0.1%
2 | 11 | -0.513% | -0.018% | -0.392% | 0.1%
2 | 12 | -0.195% | 0.318% | -0.401% | 0.1%
2 | 13 | 0.027% | 0.222% | -0.227% | 0.1%
2 | 14 | 0.196% | 0.169% | 0.009% | 0.1%
2 | 15 | 0.267% | 0.071% | 0.163% | 0.1%
2 | 16 | 0.276% | 0.009% | 0.246% | 0.1%
2 | 17 | 0.25% | -0.026% | 0.264% | 0.1%
Deviations of %error on test id #3. Total 18 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -28.09% | 0% | 0% | 0.1%
3 | 1 | -18.97% | 9.121% | 0% | 0.1%
3 | 2 | -2.942% | 16.03% | -16.67% | 0.1%
3 | 3 | -1.35% | 1.592% | -7.754% | 0.1%
3 | 4 | 0.258% | 1.608% | -1.345% | 0.1%
3 | 5 | 0.16% | -0.098% | -0.311% | 0.1%
3 | 6 | -0.009% | -0.169% | 0.136% | 0.1%
3 | 7 | -0.044% | -0.035% | 0.036% | 0.1%
3 | 8 | 0.044% | 0.088% | -0.003% | 0.1%
3 | 9 | -0.062% | -0.106% | -0.021% | 0.1%
3 | 10 | -0.539% | -0.477% | -0.186% | 0.1%
3 | 11 | -0.566% | -0.027% | -0.389% | 0.1%
3 | 12 | -0.328% | 0.238% | -0.478% | 0.1%
3 | 13 | -0.089% | 0.239% | -0.328% | 0.1%
3 | 14 | 0.151% | 0.24% | -0.089% | 0.1%
3 | 15 | 0.241% | 0.09% | 0.101% | 0.1%
3 | 16 | 0.223% | -0.018% | 0.205% | 0.1%
3 | 17 | 0.25% | 0.027% | 0.238% | 0.1%
Deviations of %error on test id #4. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -29.41% | 0% | 0% | 0.1%
4 | 1 | -27.59% | 1.826% | 0% | 0.1%
4 | 2 | -17.09% | 10.5% | -24.7% | 0.1%
4 | 3 | -5.637% | 11.45% | -16.77% | 0.1%
4 | 4 | -1.832% | 3.805% | -8.186% | 0.1%
4 | 5 | 0.214% | 2.046% | -2.418% | 0.1%
4 | 6 | 0.223% | 0.009% | -0.465% | 0.1%
4 | 7 | 0.018% | -0.205% | 0.152% | 0.1%
4 | 8 | -0.062% | -0.08% | 0.06% | 0.1%
4 | 9 | -0.027% | 0.035% | -0.024% | 0.1%
4 | 10 | -0.018% | 0.009% | -0.036% | 0.1%
Deviations of %error on test id #5. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -36.94% | 0% | 0% | 0.1%
5 | 1 | -25.06% | 11.88% | 0% | 0.1%
5 | 2 | -16.09% | 8.971% | -26.03% | 0.1%
5 | 3 | -8.178% | 7.91% | -16.44% | 0.1%
5 | 4 | -0.345% | 7.833% | -8.204% | 0.1%
5 | 5 | 0.258% | 0.603% | -2.755% | 0.1%
5 | 6 | 0.044% | -0.214% | -0.014% | 0.1%
5 | 7 | -0.151% | -0.195% | 0.05% | 0.1%
5 | 8 | -0.133% | 0.018% | -0.08% | 0.1%
5 | 9 | -0.036% | 0.097% | -0.107% | 0.1%
5 | 10 | -0.089% | -0.053% | -0.086% | 0.1%

## 23. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:10:25 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.011% | PASSED | 2023-07-27 16:20
Yes | 1 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.009% | PASSED | 2023-07-27 16:21
Yes | 2 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.038% | PASSED | 2023-07-27 16:21
Yes | 3 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.043% | PASSED | 2023-07-27 16:21
Yes | 4 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.01% | PASSED | 2023-07-27 16:21
Yes | 5 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.027% | PASSED | 2023-07-27 16:22
Yes | 6 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.061% | PASSED | 2023-07-27 16:22
Yes | 7 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.076% | PASSED | 2023-07-27 16:22
Yes | 8 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.046% | PASSED | 2023-07-27 16:23
Yes | 9 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.04% | PASSED | 2023-07-27 16:23
Yes | 10 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.086% | PASSED | 2023-07-27 16:23
Yes | 11 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.083% | PASSED | 2023-07-27 16:24
Yes | 12 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.08% | PASSED | 2023-07-27 16:24
Yes | 13 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.091% | PASSED | 2023-07-27 16:24
Yes | 14 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.064% | PASSED | 2023-07-27 16:25
Yes | 15 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.049% | PASSED | 2023-07-27 16:25
Yes | 16 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.089% | PASSED | 2023-07-27 16:25
Yes | 17 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.113% | PASSED | 2023-07-27 16:26
Yes | 18 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.119% | PASSED | 2023-07-27 16:26
Yes | 19 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.091% | PASSED | 2023-07-27 16:26
Yes | 20 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.098% | PASSED | 2023-07-27 16:27
Yes | 21 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.1% | PASSED | 2023-07-27 16:27
Yes | 22 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.11% | PASSED | 2023-07-27 16:27
Yes | 23 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.132% | PASSED | 2023-07-27 16:27
Yes | 24 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.126% | PASSED | 2023-07-27 16:28
Yes | 25 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.12% | PASSED | 2023-07-27 16:28
Yes | 26 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.123% | PASSED | 2023-07-27 16:28
Yes | 27 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.108% | PASSED | 2023-07-27 16:29
Yes | 28 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.137% | PASSED | 2023-07-27 16:29
Yes | 29 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.115% | PASSED | 2023-07-27 16:29
Yes | 30 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.146% | PASSED | 2023-07-27 16:30
Yes | 31 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.14% | PASSED | 2023-07-27 16:30
Yes | 32 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.132% | PASSED | 2023-07-27 16:30
Yes | 33 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.111% | PASSED | 2023-07-27 16:31
Yes | 34 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.111% | PASSED | 2023-07-27 16:31
Yes | 35 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.117% | PASSED | 2023-07-27 16:31
Yes | 36 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.126% | PASSED | 2023-07-27 16:32
Yes | 37 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.156% | PASSED | 2023-07-27 16:32
Yes | 38 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.154% | PASSED | 2023-07-27 16:32
Yes | 39 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.143% | PASSED | 2023-07-27 16:33
Yes | 40 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.135% | PASSED | 2023-07-27 16:33
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -0.062% | 0% | 0% | 0.1%
0 | 1 | 0% | 0.062% | 0% | 0.1%
0 | 2 | -0.036% | -0.036% | -0.033% | 0.1%
0 | 3 | 0.004% | 0.04% | -0.011% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0.018% | 0% | 0% | 0.1%
1 | 1 | 0.013% | -0.005% | 0% | 0.1%
1 | 2 | -0.004% | -0.017% | 0.009% | 0.1%
1 | 3 | 0.018% | 0.022% | 0.009% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | 0.053% | 0% | 0% | 0.1%
2 | 1 | 0.009% | -0.044% | 0% | 0.1%
2 | 2 | 0.044% | 0.035% | 0.035% | 0.1%
2 | 3 | 0.062% | 0.018% | 0.038% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.013% | 0% | 0% | 0.1%
3 | 1 | 0.022% | 0.009% | 0% | 0.1%
3 | 2 | 0.049% | 0.027% | 0.028% | 0.1%
3 | 3 | 0.058% | 0.009% | 0.043% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.062% | 0% | 0% | 0.1%
4 | 1 | 0.018% | -0.044% | 0% | 0.1%
4 | 2 | 0% | -0.018% | 0.027% | 0.1%
4 | 3 | 0.013% | 0.013% | 0.01% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | 0.013% | 0% | 0% | 0.1%
5 | 1 | 0.031% | 0.018% | 0% | 0.1%
5 | 2 | 0.036% | 0.005% | 0.027% | 0.1%
5 | 3 | 0.013% | -0.023% | 0.027% | 0.1%
Deviations of %error on test id #6. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | 0.062% | 0% | 0% | 0.1%
6 | 1 | 0.049% | -0.013% | 0% | 0.1%
6 | 2 | 0.067% | 0.018% | 0.059% | 0.1%
6 | 3 | 0.067% | 0% | 0.061% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.062% | 0% | 0% | 0.1%
7 | 1 | 0.071% | 0.009% | 0% | 0.1%
7 | 2 | 0.049% | -0.022% | 0.061% | 0.1%
7 | 3 | 0.107% | 0.058% | 0.076% | 0.1%
Deviations of %error on test id #8. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | 0.089% | 0% | 0% | 0.1%
8 | 1 | 0.004% | -0.085% | 0% | 0.1%
8 | 2 | 0.049% | 0.045% | 0.047% | 0.1%
8 | 3 | 0.085% | 0.036% | 0.046% | 0.1%
Deviations of %error on test id #9. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | 0.058% | 0% | 0% | 0.1%
9 | 1 | 0.027% | -0.031% | 0% | 0.1%
9 | 2 | 0.036% | 0.009% | 0.04% | 0.1%
9 | 3 | 0.058% | 0.022% | 0.04% | 0.1%
Deviations of %error on test id #10. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | 0.053% | 0% | 0% | 0.1%
10 | 1 | 0.102% | 0.049% | 0% | 0.1%
10 | 2 | 0.062% | -0.04% | 0.072% | 0.1%
10 | 3 | 0.093% | 0.031% | 0.086% | 0.1%
Deviations of %error on test id #11. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | 0.116% | 0% | 0% | 0.1%
11 | 1 | 0.076% | -0.04% | 0% | 0.1%
11 | 2 | 0.111% | 0.035% | 0.101% | 0.1%
11 | 3 | 0.062% | -0.049% | 0.083% | 0.1%
Deviations of %error on test id #12. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | 0.102% | 0% | 0% | 0.1%
12 | 1 | 0.062% | -0.04% | 0% | 0.1%
12 | 2 | 0.093% | 0.031% | 0.086% | 0.1%
12 | 3 | 0.085% | -0.008% | 0.08% | 0.1%
Deviations of %error on test id #13. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | 0.111% | 0% | 0% | 0.1%
13 | 1 | 0.067% | -0.044% | 0% | 0.1%
13 | 2 | 0.08% | 0.013% | 0.086% | 0.1%
13 | 3 | 0.125% | 0.045% | 0.091% | 0.1%
Deviations of %error on test id #14. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | 0.098% | 0% | 0% | 0.1%
14 | 1 | 0.058% | -0.04% | 0% | 0.1%
14 | 2 | 0.049% | -0.009% | 0.068% | 0.1%
14 | 3 | 0.085% | 0.036% | 0.064% | 0.1%
Deviations of %error on test id #15. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | 0.062% | 0% | 0% | 0.1%
15 | 1 | 0.067% | 0.005% | 0% | 0.1%
15 | 2 | 0.067% | 0% | 0.065% | 0.1%
15 | 3 | 0.013% | -0.054% | 0.049% | 0.1%
Deviations of %error on test id #16. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | 0.089% | 0% | 0% | 0.1%
16 | 1 | 0.067% | -0.022% | 0% | 0.1%
16 | 2 | 0.098% | 0.031% | 0.085% | 0.1%
16 | 3 | 0.102% | 0.004% | 0.089% | 0.1%
Deviations of %error on test id #17. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | 0.138% | 0% | 0% | 0.1%
17 | 1 | 0.093% | -0.045% | 0% | 0.1%
17 | 2 | 0.107% | 0.014% | 0.113% | 0.1%
17 | 3 | 0.138% | 0.031% | 0.113% | 0.1%
Deviations of %error on test id #18. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | 0.089% | 0% | 0% | 0.1%
18 | 1 | 0.129% | 0.04% | 0% | 0.1%
18 | 2 | 0.093% | -0.036% | 0.104% | 0.1%
18 | 3 | 0.134% | 0.041% | 0.119% | 0.1%
Deviations of %error on test id #19. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | 0.134% | 0% | 0% | 0.1%
19 | 1 | 0.085% | -0.049% | 0% | 0.1%
19 | 2 | 0.098% | 0.013% | 0.106% | 0.1%
19 | 3 | 0.089% | -0.009% | 0.091% | 0.1%
Deviations of %error on test id #20. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | 0.098% | 0% | 0% | 0.1%
20 | 1 | 0.102% | 0.004% | 0% | 0.1%
20 | 2 | 0.093% | -0.009% | 0.098% | 0.1%
20 | 3 | 0.098% | 0.005% | 0.098% | 0.1%
Deviations of %error on test id #21. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | 0.089% | 0% | 0% | 0.1%
21 | 1 | 0.125% | 0.036% | 0% | 0.1%
21 | 2 | 0.049% | -0.076% | 0.088% | 0.1%
21 | 3 | 0.125% | 0.076% | 0.1% | 0.1%
Deviations of %error on test id #22. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | 0.044% | 0% | 0% | 0.1%
22 | 1 | 0.138% | 0.094% | 0% | 0.1%
22 | 2 | 0.093% | -0.045% | 0.092% | 0.1%
22 | 3 | 0.098% | 0.005% | 0.11% | 0.1%
Deviations of %error on test id #23. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | 0.134% | 0% | 0% | 0.1%
23 | 1 | 0.111% | -0.023% | 0% | 0.1%
23 | 2 | 0.125% | 0.014% | 0.123% | 0.1%
23 | 3 | 0.16% | 0.035% | 0.132% | 0.1%
Deviations of %error on test id #24. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | 0.134% | 0% | 0% | 0.1%
24 | 1 | 0.107% | -0.027% | 0% | 0.1%
24 | 2 | 0.138% | 0.031% | 0.126% | 0.1%
24 | 3 | 0.134% | -0.004% | 0.126% | 0.1%
Deviations of %error on test id #25. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | 0.12% | 0% | 0% | 0.1%
25 | 1 | 0.125% | 0.005% | 0% | 0.1%
25 | 2 | 0.134% | 0.009% | 0.126% | 0.1%
25 | 3 | 0.102% | -0.032% | 0.12% | 0.1%
Deviations of %error on test id #26. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | 0.138% | 0% | 0% | 0.1%
26 | 1 | 0.098% | -0.04% | 0% | 0.1%
26 | 2 | 0.134% | 0.036% | 0.123% | 0.1%
26 | 3 | 0.138% | 0.004% | 0.123% | 0.1%
Deviations of %error on test id #27. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | 0.085% | 0% | 0% | 0.1%
27 | 1 | 0.111% | 0.026% | 0% | 0.1%
27 | 2 | 0.12% | 0.009% | 0.105% | 0.1%
27 | 3 | 0.093% | -0.027% | 0.108% | 0.1%
Deviations of %error on test id #28. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | 0.125% | 0% | 0% | 0.1%
28 | 1 | 0.129% | 0.004% | 0% | 0.1%
28 | 2 | 0.116% | -0.013% | 0.123% | 0.1%
28 | 3 | 0.165% | 0.049% | 0.137% | 0.1%
Deviations of %error on test id #29. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | 0.134% | 0% | 0% | 0.1%
29 | 1 | 0.125% | -0.009% | 0% | 0.1%
29 | 2 | 0.134% | 0.009% | 0.131% | 0.1%
29 | 3 | 0.085% | -0.049% | 0.115% | 0.1%
Deviations of %error on test id #30. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | 0.174% | 0% | 0% | 0.1%
30 | 1 | 0.134% | -0.04% | 0% | 0.1%
30 | 2 | 0.125% | -0.009% | 0.144% | 0.1%
30 | 3 | 0.178% | 0.053% | 0.146% | 0.1%
Deviations of %error on test id #31. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | 0.129% | 0% | 0% | 0.1%
31 | 1 | 0.165% | 0.036% | 0% | 0.1%
31 | 2 | 0.12% | -0.045% | 0.138% | 0.1%
31 | 3 | 0.134% | 0.014% | 0.14% | 0.1%
Deviations of %error on test id #32. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | 0.138% | 0% | 0% | 0.1%
32 | 1 | 0.138% | 0% | 0% | 0.1%
32 | 2 | 0.134% | -0.004% | 0.137% | 0.1%
32 | 3 | 0.125% | -0.009% | 0.132% | 0.1%
Deviations of %error on test id #33. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | 0.125% | 0% | 0% | 0.1%
33 | 1 | 0.08% | -0.045% | 0% | 0.1%
33 | 2 | 0.16% | 0.08% | 0.122% | 0.1%
33 | 3 | 0.093% | -0.067% | 0.111% | 0.1%
Deviations of %error on test id #34. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | 0.111% | 0% | 0% | 0.1%
34 | 1 | 0.125% | 0.014% | 0% | 0.1%
34 | 2 | 0.08% | -0.045% | 0.105% | 0.1%
34 | 3 | 0.129% | 0.049% | 0.111% | 0.1%
Deviations of %error on test id #35. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | 0.12% | 0% | 0% | 0.1%
35 | 1 | 0.125% | 0.005% | 0% | 0.1%
35 | 2 | 0.098% | -0.027% | 0.114% | 0.1%
35 | 3 | 0.129% | 0.031% | 0.117% | 0.1%
Deviations of %error on test id #36. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | 0.129% | 0% | 0% | 0.1%
36 | 1 | 0.134% | 0.005% | 0% | 0.1%
36 | 2 | 0.134% | 0% | 0.132% | 0.1%
36 | 3 | 0.116% | -0.018% | 0.128% | 0.1%
36 | 4 | 0.129% | 0.013% | 0.126% | 0.1%
Deviations of %error on test id #37. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | 0.125% | 0% | 0% | 0.1%
37 | 1 | 0.129% | 0.004% | 0% | 0.1%
37 | 2 | 0.205% | 0.076% | 0.153% | 0.1%
37 | 3 | 0.134% | -0.071% | 0.156% | 0.1%
Deviations of %error on test id #38. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | 0.169% | 0% | 0% | 0.1%
38 | 1 | 0.142% | -0.027% | 0% | 0.1%
38 | 2 | 0.183% | 0.041% | 0.165% | 0.1%
38 | 3 | 0.138% | -0.045% | 0.154% | 0.1%
Deviations of %error on test id #39. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | 0.165% | 0% | 0% | 0.1%
39 | 1 | 0.129% | -0.036% | 0% | 0.1%
39 | 2 | 0.174% | 0.045% | 0.156% | 0.1%
39 | 3 | 0.125% | -0.049% | 0.143% | 0.1%
Deviations of %error on test id #40. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | 0.125% | 0% | 0% | 0.1%
40 | 1 | 0.16% | 0.035% | 0% | 0.1%
40 | 2 | 0.125% | -0.035% | 0.137% | 0.1%
40 | 3 | 0.12% | -0.005% | 0.135% | 0.1%

## 24. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Forward / Frequency Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:10:44 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.08% | PASSED | 2023-07-27 16:33
Yes | 1 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.083% | PASSED | 2023-07-27 16:34
Yes | 2 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.131% | PASSED | 2023-07-27 16:34
Yes | 3 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.092% | PASSED | 2023-07-27 16:34
Yes | 4 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.139% | PASSED | 2023-07-27 16:35
Yes | 5 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.125% | PASSED | 2023-07-27 16:35
Yes | 6 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.148% | PASSED | 2023-07-27 16:36
Yes | 7 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.039% | PASSED | 2023-07-27 16:36
Yes | 8 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.071% | PASSED | 2023-07-27 16:36
Yes | 9 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.214% | PASSED | 2023-07-27 16:37
Yes | 10 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.113% | PASSED | 2023-07-27 16:37
Yes | 11 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.059% | PASSED | 2023-07-27 16:38
Yes | 12 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.056% | PASSED | 2023-07-27 16:38
Yes | 13 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.244% | PASSED | 2023-07-27 16:38
Yes | 14 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.255% | PASSED | 2023-07-27 16:39
Yes | 15 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.256% | PASSED | 2023-07-27 16:39
Yes | 16 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.041% | PASSED | 2023-07-27 16:40
Yes | 17 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.315% | PASSED | 2023-07-27 16:40
Yes | 18 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.012% | PASSED | 2023-07-27 16:40
Yes | 19 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.282% | PASSED | 2023-07-27 16:41
Yes | 20 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.288% | PASSED | 2023-07-27 16:41
Yes | 21 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.3% | PASSED | 2023-07-27 16:42
Yes | 22 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.291% | PASSED | 2023-07-27 16:42
Yes | 23 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.375% | PASSED | 2023-07-27 16:42
Yes | 24 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.009% | PASSED | 2023-07-27 16:43
Yes | 25 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.044% | PASSED | 2023-07-27 16:43
Yes | 26 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.33% | PASSED | 2023-07-27 16:44
Yes | 27 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.241% | PASSED | 2023-07-27 16:44
Yes | 28 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.363% | PASSED | 2023-07-27 16:44
Yes | 29 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.411% | PASSED | 2023-07-27 16:45
Yes | 30 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.303% | PASSED | 2023-07-27 16:45
Yes | 31 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.405% | PASSED | 2023-07-27 16:46
Yes | 32 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.375% | PASSED | 2023-07-27 16:46
Yes | 33 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.467% | PASSED | 2023-07-27 16:47
Yes | 34 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.452% | PASSED | 2023-07-27 16:47
Yes | 35 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.458% | PASSED | 2023-07-27 16:48
Yes | 36 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.443% | PASSED | 2023-07-27 16:48
Yes | 37 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.47% | PASSED | 2023-07-27 16:48
Yes | 38 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.458% | PASSED | 2023-07-27 16:49
Yes | 39 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.509% | PASSED | 2023-07-27 16:49
Yes | 40 | DQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.432% | PASSED | 2023-07-27 16:50
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -30.84% | 0% | 0% | 0.1%
0 | 1 | -24% | 6.849% | 0% | 0.1%
0 | 2 | -7.063% | 16.93% | -20.64% | 0.1%
0 | 3 | -3.78% | 3.283% | -11.61% | 0.1%
0 | 4 | 0.08% | 3.86% | -3.588% | 0.1%
0 | 5 | 0.134% | 0.054% | -1.189% | 0.1%
0 | 6 | 0.098% | -0.036% | 0.104% | 0.1%
0 | 7 | 0.009% | -0.089% | 0.08% | 0.1%
Deviations of %error on test id #1. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -38.05% | 0% | 0% | 0.1%
1 | 1 | -30.25% | 7.805% | 0% | 0.1%
1 | 2 | -15.28% | 14.97% | -27.86% | 0.1%
1 | 3 | -9.039% | 6.24% | -18.19% | 0.1%
1 | 4 | 0.125% | 9.164% | -8.064% | 0.1%
1 | 5 | -0.106% | -0.231% | -3.007% | 0.1%
1 | 6 | -0.177% | -0.071% | -0.053% | 0.1%
1 | 7 | -0.071% | 0.106% | -0.118% | 0.1%
1 | 8 | -0.044% | 0.027% | -0.097% | 0.1%
1 | 9 | -0.195% | -0.151% | -0.103% | 0.1%
1 | 10 | -0.133% | 0.062% | -0.124% | 0.1%
1 | 11 | -0.071% | 0.062% | -0.133% | 0.1%
1 | 12 | -0.044% | 0.027% | -0.083% | 0.1%
Deviations of %error on test id #2. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -28.77% | 0% | 0% | 0.1%
2 | 1 | -19.6% | 9.168% | 0% | 0.1%
2 | 2 | -11.58% | 8.018% | -19.98% | 0.1%
2 | 3 | 0.116% | 11.7% | -10.36% | 0.1%
2 | 4 | 0.134% | 0.018% | -3.778% | 0.1%
2 | 5 | 0.125% | -0.009% | 0.125% | 0.1%
2 | 6 | 0.134% | 0.009% | 0.131% | 0.1%
Deviations of %error on test id #3. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -34.11% | 0% | 0% | 0.1%
3 | 1 | -27.06% | 7.05% | 0% | 0.1%
3 | 2 | -10.81% | 16.24% | -23.99% | 0.1%
3 | 3 | -5.207% | 5.606% | -14.36% | 0.1%
3 | 4 | 0.062% | 5.269% | -5.319% | 0.1%
3 | 5 | 0.036% | -0.026% | -1.703% | 0.1%
3 | 6 | 0.125% | 0.089% | 0.074% | 0.1%
3 | 7 | 0.071% | -0.054% | 0.077% | 0.1%
3 | 8 | 0.08% | 0.009% | 0.092% | 0.1%
Deviations of %error on test id #4. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -33.78% | 0% | 0% | 0.1%
4 | 1 | -21.37% | 12.42% | 0% | 0.1%
4 | 2 | -13.69% | 7.68% | -22.95% | 0.1%
4 | 3 | -2.386% | 11.3% | -12.48% | 0.1%
4 | 4 | -0.724% | 1.662% | -5.599% | 0.1%
4 | 5 | -0.115% | 0.609% | -1.075% | 0.1%
4 | 6 | -0.142% | -0.027% | -0.327% | 0.1%
4 | 7 | -0.151% | -0.009% | -0.136% | 0.1%
4 | 8 | -0.124% | 0.027% | -0.139% | 0.1%
Deviations of %error on test id #5. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -29.66% | 0% | 0% | 0.1%
5 | 1 | -21.97% | 7.684% | 0% | 0.1%
5 | 2 | -9.091% | 12.88% | -20.24% | 0.1%
5 | 3 | -2.08% | 7.011% | -11.05% | 0.1%
5 | 4 | 0.125% | 2.205% | -3.682% | 0.1%
5 | 5 | 0.071% | -0.054% | -0.628% | 0.1%
5 | 6 | 0.125% | 0.054% | 0.107% | 0.1%
5 | 7 | -0.027% | -0.152% | 0.056% | 0.1%
5 | 8 | -0.337% | -0.31% | -0.08% | 0.1%
5 | 9 | -0.354% | -0.017% | -0.239% | 0.1%
5 | 10 | -0.222% | 0.132% | -0.304% | 0.1%
5 | 11 | -0.044% | 0.178% | -0.207% | 0.1%
5 | 12 | 0.071% | 0.115% | -0.065% | 0.1%
5 | 13 | -0.478% | -0.549% | -0.15% | 0.1%
5 | 14 | -0.407% | 0.071% | -0.271% | 0.1%
5 | 15 | -0.23% | 0.177% | -0.372% | 0.1%
5 | 16 | -0.098% | 0.132% | -0.245% | 0.1%
5 | 17 | 0.089% | 0.187% | -0.08% | 0.1%
5 | 18 | 0.178% | 0.089% | 0.056% | 0.1%
5 | 19 | 0.107% | -0.071% | 0.125% | 0.1%
Deviations of %error on test id #6. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | -17.34% | 0% | 0% | 0.1%
6 | 1 | -9.398% | 7.942% | 0% | 0.1%
6 | 2 | 0.125% | 9.523% | -8.871% | 0.1%
6 | 3 | 0.142% | 0.017% | -3.044% | 0.1%
6 | 4 | 0.187% | 0.045% | 0.151% | 0.1%
6 | 5 | 0.116% | -0.071% | 0.148% | 0.1%
Deviations of %error on test id #7. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | -32.68% | 0% | 0% | 0.1%
7 | 1 | -16.91% | 15.77% | 0% | 0.1%
7 | 2 | -9.296% | 7.616% | -19.63% | 0.1%
7 | 3 | -2.242% | 7.054% | -9.483% | 0.1%
7 | 4 | -0.062% | 2.18% | -3.867% | 0.1%
7 | 5 | 0% | 0.062% | -0.768% | 0.1%
7 | 6 | 0.08% | 0.08% | 0.006% | 0.1%
7 | 7 | 0.036% | -0.044% | 0.039% | 0.1%
Deviations of %error on test id #8. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | -33.25% | 0% | 0% | 0.1%
8 | 1 | -20.93% | 12.32% | 0% | 0.1%
8 | 2 | -12.64% | 8.284% | -22.27% | 0.1%
8 | 3 | -2.454% | 10.19% | -12.01% | 0.1%
8 | 4 | -0.328% | 2.126% | -5.141% | 0.1%
8 | 5 | -0.027% | 0.301% | -0.936% | 0.1%
8 | 6 | -0.115% | -0.088% | -0.157% | 0.1%
8 | 7 | -0.098% | 0.017% | -0.08% | 0.1%
8 | 8 | 0% | 0.098% | -0.071% | 0.1%
Deviations of %error on test id #9. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | -34.5% | 0% | 0% | 0.1%
9 | 1 | -23.19% | 11.31% | 0% | 0.1%
9 | 2 | -14.09% | 9.092% | -23.93% | 0.1%
9 | 3 | -2.589% | 11.51% | -13.29% | 0.1%
9 | 4 | -0.257% | 2.332% | -5.647% | 0.1%
9 | 5 | 0.018% | 0.275% | -0.943% | 0.1%
9 | 6 | -0.106% | -0.124% | -0.115% | 0.1%
9 | 7 | -0.062% | 0.044% | -0.05% | 0.1%
9 | 8 | -0.222% | -0.16% | -0.13% | 0.1%
9 | 9 | -0.222% | 0% | -0.169% | 0.1%
9 | 10 | -0.151% | 0.071% | -0.198% | 0.1%
9 | 11 | -0.036% | 0.115% | -0.136% | 0.1%
9 | 12 | 0.08% | 0.116% | -0.036% | 0.1%
9 | 13 | 0.214% | 0.134% | 0.086% | 0.1%
9 | 14 | 0.223% | 0.009% | 0.172% | 0.1%
9 | 15 | 0.196% | -0.027% | 0.211% | 0.1%
9 | 16 | 0.223% | 0.027% | 0.214% | 0.1%
Deviations of %error on test id #10. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | -28.21% | 0% | 0% | 0.1%
10 | 1 | -17.42% | 10.79% | 0% | 0.1%
10 | 2 | -10.09% | 7.333% | -18.57% | 0.1%
10 | 3 | -2.157% | 7.929% | -9.887% | 0.1%
10 | 4 | 0.151% | 2.308% | -4.031% | 0.1%
10 | 5 | -0.036% | -0.187% | -0.681% | 0.1%
10 | 6 | 0.027% | 0.063% | 0.047% | 0.1%
10 | 7 | 0.107% | 0.08% | 0.033% | 0.1%
10 | 8 | 0.205% | 0.098% | 0.113% | 0.1%
Deviations of %error on test id #11. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | -30.89% | 0% | 0% | 0.1%
11 | 1 | -16.55% | 14.33% | 0% | 0.1%
11 | 2 | -7.567% | 8.988% | -18.34% | 0.1%
11 | 3 | -1.858% | 5.709% | -8.66% | 0.1%
11 | 4 | 0.125% | 1.983% | -3.1% | 0.1%
11 | 5 | -0.027% | -0.152% | -0.587% | 0.1%
11 | 6 | 0.009% | 0.036% | 0.036% | 0.1%
11 | 7 | 0.08% | 0.071% | 0.021% | 0.1%
11 | 8 | 0.089% | 0.009% | 0.059% | 0.1%
Deviations of %error on test id #12. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | -29.82% | 0% | 0% | 0.1%
12 | 1 | -17.13% | 12.69% | 0% | 0.1%
12 | 2 | -5.7% | 11.43% | -17.55% | 0.1%
12 | 3 | -1.419% | 4.281% | -8.082% | 0.1%
12 | 4 | 0.232% | 1.651% | -2.296% | 0.1%
12 | 5 | 0.205% | -0.027% | -0.327% | 0.1%
12 | 6 | -0.036% | -0.241% | 0.134% | 0.1%
12 | 7 | -0.053% | -0.017% | 0.039% | 0.1%
12 | 8 | -0.036% | 0.017% | -0.042% | 0.1%
12 | 9 | -0.08% | -0.044% | -0.056% | 0.1%
Deviations of %error on test id #13. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | -23.68% | 0% | 0% | 0.1%
13 | 1 | -15.12% | 8.558% | 0% | 0.1%
13 | 2 | -6.693% | 8.426% | -15.16% | 0.1%
13 | 3 | 0.169% | 6.862% | -7.214% | 0.1%
13 | 4 | 0.232% | 0.063% | -2.097% | 0.1%
13 | 5 | 0.036% | -0.196% | 0.146% | 0.1%
13 | 6 | -0.044% | -0.08% | 0.075% | 0.1%
13 | 7 | -0.009% | 0.035% | -0.006% | 0.1%
13 | 8 | -0.16% | -0.151% | -0.071% | 0.1%
13 | 9 | -0.486% | -0.326% | -0.218% | 0.1%
13 | 10 | -0.557% | -0.071% | -0.401% | 0.1%
13 | 11 | -0.16% | 0.397% | -0.401% | 0.1%
13 | 12 | 0.018% | 0.178% | -0.233% | 0.1%
13 | 13 | 0.214% | 0.196% | 0.024% | 0.1%
13 | 14 | 0.223% | 0.009% | 0.152% | 0.1%
13 | 15 | 0.276% | 0.053% | 0.238% | 0.1%
13 | 16 | 0.232% | -0.044% | 0.244% | 0.1%
Deviations of %error on test id #14. Total 18 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | -31.71% | 0% | 0% | 0.1%
14 | 1 | -22.99% | 8.722% | 0% | 0.1%
14 | 2 | -12.94% | 10.05% | -22.55% | 0.1%
14 | 3 | 0.169% | 13.11% | -11.92% | 0.1%
14 | 4 | 0.232% | 0.063% | -4.179% | 0.1%
14 | 5 | 0.125% | -0.107% | 0.175% | 0.1%
14 | 6 | -0.044% | -0.169% | 0.104% | 0.1%
14 | 7 | 0.071% | 0.115% | 0.051% | 0.1%
14 | 8 | -0.142% | -0.213% | -0.038% | 0.1%
14 | 9 | -0.062% | 0.08% | -0.044% | 0.1%
14 | 10 | -0.266% | -0.204% | -0.157% | 0.1%
14 | 11 | -0.53% | -0.264% | -0.286% | 0.1%
14 | 12 | -0.31% | 0.22% | -0.369% | 0.1%
14 | 13 | -0.106% | 0.204% | -0.315% | 0.1%
14 | 14 | 0.16% | 0.266% | -0.085% | 0.1%
14 | 15 | 0.214% | 0.054% | 0.089% | 0.1%
14 | 16 | 0.285% | 0.071% | 0.22% | 0.1%
14 | 17 | 0.267% | -0.018% | 0.255% | 0.1%
Deviations of %error on test id #15. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | -28.03% | 0% | 0% | 0.1%
15 | 1 | -18.4% | 9.631% | 0% | 0.1%
15 | 2 | -9.076% | 9.325% | -18.5% | 0.1%
15 | 3 | 0.223% | 9.299% | -9.085% | 0.1%
15 | 4 | 0.232% | 0.009% | -2.874% | 0.1%
15 | 5 | 0.25% | 0.018% | 0.235% | 0.1%
15 | 6 | 0.285% | 0.035% | 0.256% | 0.1%
Deviations of %error on test id #16. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | -26.84% | 0% | 0% | 0.1%
16 | 1 | -18.44% | 8.396% | 0% | 0.1%
16 | 2 | -9.42% | 9.022% | -18.23% | 0.1%
16 | 3 | 0.196% | 9.616% | -9.222% | 0.1%
16 | 4 | 0.276% | 0.08% | -2.983% | 0.1%
16 | 5 | -0.106% | -0.382% | 0.122% | 0.1%
16 | 6 | -0.08% | 0.026% | 0.03% | 0.1%
16 | 7 | 0.018% | 0.098% | -0.056% | 0.1%
16 | 8 | -0.062% | -0.08% | -0.041% | 0.1%
Deviations of %error on test id #17. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | -25.13% | 0% | 0% | 0.1%
17 | 1 | -10.69% | 14.45% | 0% | 0.1%
17 | 2 | -2.496% | 8.189% | -12.77% | 0.1%
17 | 3 | 0.303% | 2.799% | -4.293% | 0.1%
17 | 4 | 0.196% | -0.107% | -0.666% | 0.1%
17 | 5 | 0.009% | -0.187% | 0.169% | 0.1%
17 | 6 | 0.089% | 0.08% | 0.098% | 0.1%
17 | 7 | -0.062% | -0.151% | 0.012% | 0.1%
17 | 8 | -0.434% | -0.372% | -0.136% | 0.1%
17 | 9 | -0.469% | -0.035% | -0.322% | 0.1%
17 | 10 | -0.186% | 0.283% | -0.363% | 0.1%
17 | 11 | 0.018% | 0.204% | -0.212% | 0.1%
17 | 12 | 0.241% | 0.223% | 0.024% | 0.1%
17 | 13 | 0.294% | 0.053% | 0.184% | 0.1%
17 | 14 | 0.303% | 0.009% | 0.279% | 0.1%
17 | 15 | 0.348% | 0.045% | 0.315% | 0.1%
Deviations of %error on test id #18. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | -28.42% | 0% | 0% | 0.1%
18 | 1 | -17.95% | 10.46% | 0% | 0.1%
18 | 2 | -8.477% | 9.478% | -18.28% | 0.1%
18 | 3 | 0.312% | 8.789% | -8.707% | 0.1%
18 | 4 | 0.223% | -0.089% | -2.647% | 0.1%
18 | 5 | -0.044% | -0.267% | 0.164% | 0.1%
18 | 6 | 0.053% | 0.097% | 0.077% | 0.1%
18 | 7 | -0.044% | -0.097% | -0.012% | 0.1%
18 | 8 | 0% | 0.044% | 0.003% | 0.1%
18 | 9 | 0.08% | 0.08% | 0.012% | 0.1%
Deviations of %error on test id #19. Total 19 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | -26.17% | 0% | 0% | 0.1%
19 | 1 | -16.34% | 9.828% | 0% | 0.1%
19 | 2 | -7.445% | 8.893% | -16.65% | 0.1%
19 | 3 | 0.312% | 7.757% | -7.824% | 0.1%
19 | 4 | 0.25% | -0.062% | -2.294% | 0.1%
19 | 5 | 0.303% | 0.053% | 0.288% | 0.1%
19 | 6 | 0.151% | -0.152% | 0.235% | 0.1%
19 | 7 | -0.133% | -0.284% | 0.107% | 0.1%
19 | 8 | -0.133% | 0% | -0.038% | 0.1%
19 | 9 | 0.036% | 0.169% | -0.077% | 0.1%
19 | 10 | 0.142% | 0.106% | 0.015% | 0.1%
19 | 11 | -0.177% | -0.319% | 0% | 0.1%
19 | 12 | -0.337% | -0.16% | -0.124% | 0.1%
19 | 13 | -0.089% | 0.248% | -0.201% | 0.1%
19 | 14 | 0.125% | 0.214% | -0.1% | 0.1%
19 | 15 | 0.267% | 0.142% | 0.101% | 0.1%
19 | 16 | 0.267% | 0% | 0.22% | 0.1%
19 | 17 | 0.312% | 0.045% | 0.282% | 0.1%
19 | 18 | 0.267% | -0.045% | 0.282% | 0.1%
Deviations of %error on test id #20. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | -29.67% | 0% | 0% | 0.1%
20 | 1 | -16.72% | 12.95% | 0% | 0.1%
20 | 2 | -6.038% | 10.68% | -17.48% | 0.1%
20 | 3 | -1.489% | 4.549% | -8.083% | 0.1%
20 | 4 | -0.071% | 1.418% | -2.533% | 0.1%
20 | 5 | 0.062% | 0.133% | -0.499% | 0.1%
20 | 6 | 0.089% | 0.027% | 0.027% | 0.1%
20 | 7 | -0.044% | -0.133% | 0.036% | 0.1%
20 | 8 | -0.133% | -0.089% | -0.029% | 0.1%
20 | 9 | 0.062% | 0.195% | -0.038% | 0.1%
20 | 10 | 0.169% | 0.107% | 0.033% | 0.1%
20 | 11 | 0.285% | 0.116% | 0.172% | 0.1%
20 | 12 | 0.294% | 0.009% | 0.249% | 0.1%
20 | 13 | 0.267% | -0.027% | 0.282% | 0.1%
20 | 14 | 0.303% | 0.036% | 0.288% | 0.1%
Deviations of %error on test id #21. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | -26.67% | 0% | 0% | 0.1%
21 | 1 | -13.45% | 13.22% | 0% | 0.1%
21 | 2 | -1.995% | 11.45% | -14.04% | 0.1%
21 | 3 | 0.196% | 2.191% | -5.082% | 0.1%
21 | 4 | 0.294% | 0.098% | -0.502% | 0.1%
21 | 5 | 0.285% | -0.009% | 0.258% | 0.1%
21 | 6 | -0.089% | -0.374% | 0.163% | 0.1%
21 | 7 | -0.23% | -0.141% | -0.011% | 0.1%
21 | 8 | -0.284% | -0.054% | -0.201% | 0.1%
21 | 9 | -0.16% | 0.124% | -0.225% | 0.1%
21 | 10 | 0.098% | 0.258% | -0.115% | 0.1%
21 | 11 | 0.232% | 0.134% | 0.057% | 0.1%
21 | 12 | 0.276% | 0.044% | 0.202% | 0.1%
21 | 13 | 0.294% | 0.018% | 0.267% | 0.1%
21 | 14 | 0.33% | 0.036% | 0.3% | 0.1%
Deviations of %error on test id #22. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | -28.48% | 0% | 0% | 0.1%
22 | 1 | -16.76% | 11.72% | 0% | 0.1%
22 | 2 | -6.986% | 9.773% | -17.41% | 0.1%
22 | 3 | 0.241% | 7.227% | -7.835% | 0.1%
22 | 4 | 0.285% | 0.044% | -2.153% | 0.1%
22 | 5 | 0.285% | 0% | 0.27% | 0.1%
22 | 6 | 0.303% | 0.018% | 0.291% | 0.1%
Deviations of %error on test id #23. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | -21.48% | 0% | 0% | 0.1%
23 | 1 | -11.13% | 10.35% | 0% | 0.1%
23 | 2 | -2.411% | 8.719% | -11.67% | 0.1%
23 | 3 | 0.366% | 2.777% | -4.392% | 0.1%
23 | 4 | 0.33% | -0.036% | -0.572% | 0.1%
23 | 5 | 0.134% | -0.196% | 0.277% | 0.1%
23 | 6 | -0.195% | -0.329% | 0.09% | 0.1%
23 | 7 | -0.266% | -0.071% | -0.109% | 0.1%
23 | 8 | -0.062% | 0.204% | -0.174% | 0.1%
23 | 9 | 0.134% | 0.196% | -0.065% | 0.1%
23 | 10 | 0.276% | 0.142% | 0.116% | 0.1%
23 | 11 | 0.366% | 0.09% | 0.259% | 0.1%
23 | 12 | 0.375% | 0.009% | 0.339% | 0.1%
23 | 13 | 0.384% | 0.009% | 0.375% | 0.1%
Deviations of %error on test id #24. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | -23.19% | 0% | 0% | 0.1%
24 | 1 | -9.223% | 13.96% | 0% | 0.1%
24 | 2 | -1.79% | 7.433% | -11.4% | 0.1%
24 | 3 | 0.384% | 2.174% | -3.543% | 0.1%
24 | 4 | 0.178% | -0.206% | -0.409% | 0.1%
24 | 5 | 0.089% | -0.089% | 0.217% | 0.1%
24 | 6 | 0.151% | 0.062% | 0.139% | 0.1%
24 | 7 | 0.053% | -0.098% | 0.098% | 0.1%
24 | 8 | -0.177% | -0.23% | 0.009% | 0.1%
Deviations of %error on test id #25. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | -19.39% | 0% | 0% | 0.1%
25 | 1 | -6.832% | 12.56% | 0% | 0.1%
25 | 2 | -1.635% | 5.197% | -9.285% | 0.1%
25 | 3 | 0.25% | 1.885% | -2.739% | 0.1%
25 | 4 | -0.018% | -0.268% | -0.468% | 0.1%
25 | 5 | 0.027% | 0.045% | 0.086% | 0.1%
25 | 6 | -0.036% | -0.063% | -0.009% | 0.1%
25 | 7 | -0.124% | -0.088% | -0.044% | 0.1%
Deviations of %error on test id #26. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | -19.29% | 0% | 0% | 0.1%
26 | 1 | -9.106% | 10.18% | 0% | 0.1%
26 | 2 | -2.055% | 7.051% | -10.15% | 0.1%
26 | 3 | 0.348% | 2.403% | -3.604% | 0.1%
26 | 4 | 0.214% | -0.134% | -0.498% | 0.1%
26 | 5 | -0.036% | -0.25% | 0.175% | 0.1%
26 | 6 | 0.089% | 0.125% | 0.089% | 0.1%
26 | 7 | -0.027% | -0.116% | 0.009% | 0.1%
26 | 8 | -0.124% | -0.097% | -0.021% | 0.1%
26 | 9 | -0.46% | -0.336% | -0.204% | 0.1%
26 | 10 | -0.177% | 0.283% | -0.254% | 0.1%
26 | 11 | 0.125% | 0.302% | -0.171% | 0.1%
26 | 12 | 0.267% | 0.142% | 0.072% | 0.1%
26 | 13 | 0.357% | 0.09% | 0.25% | 0.1%
26 | 14 | 0.312% | -0.045% | 0.312% | 0.1%
26 | 15 | 0.321% | 0.009% | 0.33% | 0.1%
Deviations of %error on test id #27. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | -27.71% | 0% | 0% | 0.1%
27 | 1 | -17.22% | 10.49% | 0% | 0.1%
27 | 2 | -3.92% | 13.3% | -16.29% | 0.1%
27 | 3 | -0.451% | 3.469% | -7.198% | 0.1%
27 | 4 | 0.098% | 0.549% | -1.424% | 0.1%
27 | 5 | 0.187% | 0.089% | -0.055% | 0.1%
27 | 6 | 0.241% | 0.054% | 0.175% | 0.1%
27 | 7 | 0.321% | 0.08% | 0.25% | 0.1%
27 | 8 | 0.16% | -0.161% | 0.241% | 0.1%
Deviations of %error on test id #28. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | -17.37% | 0% | 0% | 0.1%
28 | 1 | -5.668% | 11.7% | 0% | 0.1%
28 | 2 | 0.375% | 6.043% | -7.554% | 0.1%
28 | 3 | 0.375% | 0% | -1.639% | 0.1%
28 | 4 | 0.375% | 0% | 0.375% | 0.1%
28 | 5 | 0.339% | -0.036% | 0.363% | 0.1%
Deviations of %error on test id #29. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | -18.68% | 0% | 0% | 0.1%
29 | 1 | -8.096% | 10.58% | 0% | 0.1%
29 | 2 | -1.781% | 6.315% | -9.518% | 0.1%
29 | 3 | 0.142% | 1.923% | -3.245% | 0.1%
29 | 4 | 0.258% | 0.116% | -0.46% | 0.1%
29 | 5 | 0.276% | 0.018% | 0.225% | 0.1%
29 | 6 | 0.08% | -0.196% | 0.205% | 0.1%
29 | 7 | 0.018% | -0.062% | 0.125% | 0.1%
29 | 8 | 0.062% | 0.044% | 0.053% | 0.1%
29 | 9 | 0.294% | 0.232% | 0.125% | 0.1%
29 | 10 | 0.384% | 0.09% | 0.247% | 0.1%
29 | 11 | 0.42% | 0.036% | 0.366% | 0.1%
29 | 12 | 0.428% | 0.008% | 0.411% | 0.1%
Deviations of %error on test id #30. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | -17.03% | 0% | 0% | 0.1%
30 | 1 | -4.328% | 12.7% | 0% | 0.1%
30 | 2 | -0.213% | 4.115% | -7.19% | 0.1%
30 | 3 | 0.125% | 0.338% | -1.472% | 0.1%
30 | 4 | 0.044% | -0.081% | -0.015% | 0.1%
30 | 5 | 0% | -0.044% | 0.056% | 0.1%
30 | 6 | 0.187% | 0.187% | 0.077% | 0.1%
30 | 7 | 0.214% | 0.027% | 0.134% | 0.1%
30 | 8 | 0.303% | 0.089% | 0.235% | 0.1%
30 | 9 | 0.393% | 0.09% | 0.303% | 0.1%
Deviations of %error on test id #31. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | -15.32% | 0% | 0% | 0.1%
31 | 1 | -5.573% | 9.751% | 0% | 0.1%
31 | 2 | 0.134% | 5.707% | -6.921% | 0.1%
31 | 3 | 0.107% | -0.027% | -1.777% | 0.1%
31 | 4 | 0.241% | 0.134% | 0.161% | 0.1%
31 | 5 | 0.33% | 0.089% | 0.226% | 0.1%
31 | 6 | 0.134% | -0.196% | 0.235% | 0.1%
31 | 7 | -0.062% | -0.196% | 0.134% | 0.1%
31 | 8 | 0.027% | 0.089% | 0.033% | 0.1%
31 | 9 | 0.232% | 0.205% | 0.066% | 0.1%
31 | 10 | 0.384% | 0.152% | 0.214% | 0.1%
31 | 11 | 0.411% | 0.027% | 0.342% | 0.1%
31 | 12 | 0.42% | 0.009% | 0.405% | 0.1%
31 | 13 | 0.384% | -0.036% | 0.405% | 0.1%
Deviations of %error on test id #32. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | -20.22% | 0% | 0% | 0.1%
32 | 1 | -7.109% | 13.11% | 0% | 0.1%
32 | 2 | -1.471% | 5.638% | -9.601% | 0.1%
32 | 3 | 0.027% | 1.498% | -2.851% | 0.1%
32 | 4 | -0.009% | -0.036% | -0.484% | 0.1%
32 | 5 | -0.098% | -0.089% | -0.027% | 0.1%
32 | 6 | 0.107% | 0.205% | 0% | 0.1%
32 | 7 | -0.169% | -0.276% | -0.053% | 0.1%
32 | 8 | -0.133% | 0.036% | -0.065% | 0.1%
32 | 9 | 0.134% | 0.267% | -0.056% | 0.1%
32 | 10 | 0.321% | 0.187% | 0.107% | 0.1%
32 | 11 | 0.375% | 0.054% | 0.277% | 0.1%
32 | 12 | 0.366% | -0.009% | 0.354% | 0.1%
32 | 13 | 0.384% | 0.018% | 0.375% | 0.1%
Deviations of %error on test id #33. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | -15.01% | 0% | 0% | 0.1%
33 | 1 | -5.897% | 9.113% | 0% | 0.1%
33 | 2 | -1.212% | 4.685% | -7.373% | 0.1%
33 | 3 | 0.196% | 1.408% | -2.304% | 0.1%
33 | 4 | 0.276% | 0.08% | -0.247% | 0.1%
33 | 5 | 0.214% | -0.062% | 0.229% | 0.1%
33 | 6 | 0.08% | -0.134% | 0.19% | 0.1%
33 | 7 | 0.036% | -0.044% | 0.11% | 0.1%
33 | 8 | 0.125% | 0.089% | 0.08% | 0.1%
33 | 9 | 0.384% | 0.259% | 0.182% | 0.1%
33 | 10 | 0.446% | 0.062% | 0.318% | 0.1%
33 | 11 | 0.473% | 0.027% | 0.434% | 0.1%
33 | 12 | 0.482% | 0.009% | 0.467% | 0.1%
Deviations of %error on test id #34. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | -18.66% | 0% | 0% | 0.1%
34 | 1 | -8.692% | 9.963% | 0% | 0.1%
34 | 2 | -1.995% | 6.697% | -9.781% | 0.1%
34 | 3 | 0.107% | 2.102% | -3.527% | 0.1%
34 | 4 | 0.223% | 0.116% | -0.555% | 0.1%
34 | 5 | 0.312% | 0.089% | 0.214% | 0.1%
34 | 6 | 0.08% | -0.232% | 0.205% | 0.1%
34 | 7 | -0.009% | -0.089% | 0.128% | 0.1%
34 | 8 | 0.062% | 0.071% | 0.044% | 0.1%
34 | 9 | 0.276% | 0.214% | 0.11% | 0.1%
34 | 10 | 0.393% | 0.117% | 0.244% | 0.1%
34 | 11 | 0.437% | 0.044% | 0.369% | 0.1%
34 | 12 | 0.473% | 0.036% | 0.434% | 0.1%
34 | 13 | 0.446% | -0.027% | 0.452% | 0.1%
Deviations of %error on test id #35. Total 18 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | -24.95% | 0% | 0% | 0.1%
35 | 1 | -13.29% | 11.66% | 0% | 0.1%
35 | 2 | -4.312% | 8.976% | -14.18% | 0.1%
35 | 3 | 0.169% | 4.481% | -5.81% | 0.1%
35 | 4 | 0.339% | 0.17% | -1.268% | 0.1%
35 | 5 | 0.42% | 0.081% | 0.309% | 0.1%
35 | 6 | 0.464% | 0.044% | 0.408% | 0.1%
35 | 7 | 0.232% | -0.232% | 0.372% | 0.1%
35 | 8 | 0.258% | 0.026% | 0.318% | 0.1%
35 | 9 | 0.348% | 0.09% | 0.279% | 0.1%
35 | 10 | -0.009% | -0.357% | 0.199% | 0.1%
35 | 11 | -0.186% | -0.177% | 0.051% | 0.1%
35 | 12 | 0.134% | 0.32% | -0.02% | 0.1%
35 | 13 | 0.312% | 0.178% | 0.087% | 0.1%
35 | 14 | 0.455% | 0.143% | 0.3% | 0.1%
35 | 15 | 0.455% | 0% | 0.407% | 0.1%
35 | 16 | 0.473% | 0.018% | 0.461% | 0.1%
35 | 17 | 0.446% | -0.027% | 0.458% | 0.1%
Deviations of %error on test id #36. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | -15.71% | 0% | 0% | 0.1%
36 | 1 | -4.296% | 11.41% | 0% | 0.1%
36 | 2 | 0.241% | 4.537% | -6.587% | 0.1%
36 | 3 | 0.134% | -0.107% | -1.307% | 0.1%
36 | 4 | 0.187% | 0.053% | 0.187% | 0.1%
36 | 5 | 0.009% | -0.178% | 0.11% | 0.1%
36 | 6 | 0.036% | 0.027% | 0.077% | 0.1%
36 | 7 | 0.062% | 0.026% | 0.036% | 0.1%
36 | 8 | 0.294% | 0.232% | 0.131% | 0.1%
36 | 9 | 0.42% | 0.126% | 0.259% | 0.1%
36 | 10 | 0.42% | 0% | 0.378% | 0.1%
36 | 11 | 0.437% | 0.017% | 0.426% | 0.1%
36 | 12 | 0.473% | 0.036% | 0.443% | 0.1%
Deviations of %error on test id #37. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | -18.54% | 0% | 0% | 0.1%
37 | 1 | -7.506% | 11.03% | 0% | 0.1%
37 | 2 | 0.116% | 7.622% | -8.642% | 0.1%
37 | 3 | 0.223% | 0.107% | -2.389% | 0.1%
37 | 4 | 0.366% | 0.143% | 0.235% | 0.1%
37 | 5 | 0.428% | 0.062% | 0.339% | 0.1%
37 | 6 | 0.303% | -0.125% | 0.366% | 0.1%
37 | 7 | 0.169% | -0.134% | 0.3% | 0.1%
37 | 8 | 0.33% | 0.161% | 0.267% | 0.1%
37 | 9 | 0.169% | -0.161% | 0.223% | 0.1%
37 | 10 | -0.169% | -0.338% | 0.11% | 0.1%
37 | 11 | 0.08% | 0.249% | 0.027% | 0.1%
37 | 12 | 0.321% | 0.241% | 0.077% | 0.1%
37 | 13 | 0.437% | 0.116% | 0.279% | 0.1%
37 | 14 | 0.473% | 0.036% | 0.41% | 0.1%
37 | 15 | 0.464% | -0.009% | 0.458% | 0.1%
37 | 16 | 0.473% | 0.009% | 0.47% | 0.1%
Deviations of %error on test id #38. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | -17.99% | 0% | 0% | 0.1%
38 | 1 | -5.287% | 12.7% | 0% | 0.1%
38 | 2 | -0.504% | 4.783% | -7.927% | 0.1%
38 | 3 | 0.232% | 0.736% | -1.853% | 0.1%
38 | 4 | 0.366% | 0.134% | 0.031% | 0.1%
38 | 5 | 0.232% | -0.134% | 0.277% | 0.1%
38 | 6 | 0.107% | -0.125% | 0.235% | 0.1%
38 | 7 | 0.027% | -0.08% | 0.122% | 0.1%
38 | 8 | 0.312% | 0.285% | 0.149% | 0.1%
38 | 9 | 0.411% | 0.099% | 0.25% | 0.1%
38 | 10 | 0.455% | 0.044% | 0.393% | 0.1%
38 | 11 | 0.509% | 0.054% | 0.458% | 0.1%
Deviations of %error on test id #39. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | -16.59% | 0% | 0% | 0.1%
39 | 1 | -7.002% | 9.584% | 0% | 0.1%
39 | 2 | -1.35% | 5.652% | -8.313% | 0.1%
39 | 3 | 0.339% | 1.689% | -2.671% | 0.1%
39 | 4 | 0.455% | 0.116% | -0.185% | 0.1%
39 | 5 | 0.384% | -0.071% | 0.393% | 0.1%
39 | 6 | 0.196% | -0.188% | 0.345% | 0.1%
39 | 7 | 0.276% | 0.08% | 0.285% | 0.1%
39 | 8 | 0.428% | 0.152% | 0.3% | 0.1%
39 | 9 | 0.482% | 0.054% | 0.395% | 0.1%
39 | 10 | 0.16% | -0.322% | 0.357% | 0.1%
39 | 11 | 0.009% | -0.151% | 0.217% | 0.1%
39 | 12 | 0.214% | 0.205% | 0.128% | 0.1%
39 | 13 | 0.402% | 0.188% | 0.208% | 0.1%
39 | 14 | 0.5% | 0.098% | 0.372% | 0.1%
39 | 15 | 0.509% | 0.009% | 0.47% | 0.1%
39 | 16 | 0.518% | 0.009% | 0.509% | 0.1%
Deviations of %error on test id #40. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | -15.13% | 0% | 0% | 0.1%
40 | 1 | -5.644% | 9.482% | 0% | 0.1%
40 | 2 | -0.925% | 4.719% | -7.232% | 0.1%
40 | 3 | 0.366% | 1.291% | -2.068% | 0.1%
40 | 4 | 0.455% | 0.089% | -0.035% | 0.1%
40 | 5 | 0.285% | -0.17% | 0.369% | 0.1%
40 | 6 | 0.071% | -0.214% | 0.27% | 0.1%
40 | 7 | 0.044% | -0.027% | 0.133% | 0.1%
40 | 8 | 0.276% | 0.232% | 0.13% | 0.1%
40 | 9 | 0.357% | 0.081% | 0.226% | 0.1%
40 | 10 | 0.42% | 0.063% | 0.351% | 0.1%
40 | 11 | 0.518% | 0.098% | 0.432% | 0.1%

## 25. PF = 1.0
[<<]
Accuracy Test Report of 'PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:11:19 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 20 | 20 |  | ±0.1% | ±0.5% | 0.03% | PASSED | 2023-07-27 16:50
Yes | 1 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.5% | 0.031% | PASSED | 2023-07-27 16:51
Yes | 3 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.5% | 0.062% | PASSED | 2023-07-27 16:51
Yes | 4 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0.077% | PASSED | 2023-07-27 16:51
Yes | 7 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.5% | 0.079% | PASSED | 2023-07-27 16:52
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 0.044% | 0% | 0% | 0.1%
0 | 1 | 0.044% | 0% | 0% | 0.1%
0 | 2 | 0.036% | -0.008% | 0.041% | 0.1%
0 | 3 | 0.009% | -0.027% | 0.03% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -0.039% | 0% | 0% | 0.1%
1 | 1 | 0.018% | 0.057% | 0% | 0.1%
1 | 2 | 0.043% | 0.025% | 0.007% | 0.1%
1 | 3 | 0.032% | -0.011% | 0.031% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.04% | 0% | 0% | 0.1%
3 | 1 | 0.071% | 0.031% | 0% | 0.1%
3 | 2 | 0.058% | -0.013% | 0.056% | 0.1%
3 | 3 | 0.058% | 0% | 0.062% | 0.1%
Deviations of %error on test id #4. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.027% | 0% | 0% | 0.1%
4 | 1 | 0.027% | 0% | 0% | 0.1%
4 | 2 | 0.089% | 0.062% | 0.048% | 0.1%
4 | 3 | 0.089% | 0% | 0.068% | 0.1%
4 | 4 | 0.053% | -0.036% | 0.077% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.093% | 0% | 0% | 0.1%
7 | 1 | 0.091% | -0.002% | 0% | 0.1%
7 | 2 | 0.071% | -0.02% | 0.085% | 0.1%
7 | 3 | 0.076% | 0.005% | 0.079% | 0.1%

## 26. PF = 0.5L
[<<]
Accuracy Test Report of 'PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:11:21 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | 0.327% | PASSED | 2023-07-27 16:52
Yes | 1 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | 0.319% | PASSED | 2023-07-27 16:53
Yes | 3 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.226% | PASSED | 2023-07-27 16:53
Yes | 4 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.217% | PASSED | 2023-07-27 16:54
Yes | 7 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.205% | PASSED | 2023-07-27 16:55
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -14.11% | 0% | 0% | 0.1%
0 | 1 | 0.357% | 14.47% | 0% | 0.1%
0 | 2 | 0.339% | -0.018% | -4.471% | 0.1%
0 | 3 | 0.357% | 0.018% | 0.351% | 0.1%
0 | 4 | -0.089% | -0.446% | 0.202% | 0.1%
0 | 5 | -0.354% | -0.265% | -0.029% | 0.1%
0 | 6 | -0.124% | 0.23% | -0.189% | 0.1%
0 | 7 | -0.213% | -0.089% | -0.23% | 0.1%
0 | 8 | -0.671% | -0.458% | -0.336% | 0.1%
0 | 9 | -0.213% | 0.458% | -0.366% | 0.1%
0 | 10 | 0.214% | 0.427% | -0.223% | 0.1%
0 | 11 | 0.321% | 0.107% | 0.107% | 0.1%
0 | 12 | 0.339% | 0.018% | 0.291% | 0.1%
0 | 13 | 0.321% | -0.018% | 0.327% | 0.1%
0 | 14 | 0.321% | 0% | 0.327% | 0.1%
Deviations of %error on test id #1. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -34.39% | 0% | 0% | 0.1%
1 | 1 | -16.84% | 17.54% | 0% | 0.1%
1 | 2 | -1.718% | 15.12% | -17.65% | 0.1%
1 | 3 | 0.309% | 2.027% | -6.084% | 0.1%
1 | 4 | 0.232% | -0.077% | -0.392% | 0.1%
1 | 5 | 0.101% | -0.131% | 0.214% | 0.1%
1 | 6 | -0.018% | -0.119% | 0.105% | 0.1%
1 | 7 | -0.26% | -0.242% | -0.059% | 0.1%
1 | 8 | -0.289% | -0.029% | -0.189% | 0.1%
1 | 9 | 0.119% | 0.408% | -0.143% | 0.1%
1 | 10 | 0.303% | 0.184% | 0.044% | 0.1%
1 | 11 | 0.333% | 0.03% | 0.252% | 0.1%
1 | 12 | 0.297% | -0.036% | 0.311% | 0.1%
1 | 13 | 0.327% | 0.03% | 0.319% | 0.1%
Deviations of %error on test id #3. Total 19 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -35.89% | 0% | 0% | 0.1%
3 | 1 | -27.62% | 8.271% | 0% | 0.1%
3 | 2 | -13.39% | 14.22% | -25.63% | 0.1%
3 | 3 | -3.45% | 9.944% | -14.82% | 0.1%
3 | 4 | 0.232% | 3.682% | -5.537% | 0.1%
3 | 5 | 0.258% | 0.026% | -0.987% | 0.1%
3 | 6 | 0.223% | -0.035% | 0.238% | 0.1%
3 | 7 | 0.027% | -0.196% | 0.169% | 0.1%
3 | 8 | -0.257% | -0.284% | -0.002% | 0.1%
3 | 9 | -0.169% | 0.088% | -0.133% | 0.1%
3 | 10 | -0.036% | 0.133% | -0.154% | 0.1%
3 | 11 | -0.151% | -0.115% | -0.119% | 0.1%
3 | 12 | -0.478% | -0.327% | -0.222% | 0.1%
3 | 13 | -0.248% | 0.23% | -0.292% | 0.1%
3 | 14 | -0.018% | 0.23% | -0.248% | 0.1%
3 | 15 | 0.134% | 0.152% | -0.044% | 0.1%
3 | 16 | 0.196% | 0.062% | 0.104% | 0.1%
3 | 17 | 0.241% | 0.045% | 0.19% | 0.1%
3 | 18 | 0.241% | 0% | 0.226% | 0.1%
Deviations of %error on test id #4. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -27.77% | 0% | 0% | 0.1%
4 | 1 | -6.336% | 21.44% | 0% | 0.1%
4 | 2 | 0.25% | 6.586% | -11.29% | 0.1%
4 | 3 | 0.062% | -0.188% | -2.008% | 0.1%
4 | 4 | -0.266% | -0.328% | 0.015% | 0.1%
4 | 5 | -0.222% | 0.044% | -0.142% | 0.1%
4 | 6 | -0.266% | -0.044% | -0.251% | 0.1%
4 | 7 | 0.134% | 0.4% | -0.118% | 0.1%
4 | 8 | 0.223% | 0.089% | 0.03% | 0.1%
4 | 9 | 0.187% | -0.036% | 0.181% | 0.1%
4 | 10 | 0.366% | 0.179% | 0.259% | 0.1%
4 | 11 | 0.446% | 0.08% | 0.333% | 0.1%
4 | 12 | 0.267% | -0.179% | 0.36% | 0.1%
4 | 13 | 0.205% | -0.062% | 0.306% | 0.1%
4 | 14 | 0.223% | 0.018% | 0.232% | 0.1%
4 | 15 | 0.223% | 0% | 0.217% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.112% | 0% | 0% | 0.1%
7 | 1 | 0.207% | 0.095% | 0% | 0.1%
7 | 2 | 0.201% | -0.006% | 0.173% | 0.1%
7 | 3 | 0.207% | 0.006% | 0.205% | 0.1%

## 27. PF = 0.8C
[<<]
Accuracy Test Report of 'PF = 0.8C'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / PF = 0.8C'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:11:26 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Current Y (I1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
5 | 0 | 0 | 5
3. Test Cases
Total 5 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Imax (40A) | Imax (40A) | Imax (40A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 10 | 20 |  | ±0.1% | ±0.6% | -0.059% | PASSED | 2023-07-27 16:56
Yes | 1 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 6 | 20 |  | ±0.1% | ±0.6% | 0.014% | PASSED | 2023-07-27 16:56
Yes | 3 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.107% | PASSED | 2023-07-27 16:56
Yes | 4 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | 0.5Ib (2.5A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.6% | 0.11% | PASSED | 2023-07-27 16:57
Yes | 7 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | 0.05Ib (0.25A) | Fb (50Hz) | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 0.8C, sin(ϕ), ϕ=306.87° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.025% | PASSED | 2023-07-27 16:58
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -44.6% | 0% | 0% | 0.1%
0 | 1 | 246.6% | 291.2% | 0% | 0.1%
0 | 2 | 979.7% | 733.1% | 393.9% | 0.1%
0 | 3 | -0.071% | -979.7% | 408.7% | 0.1%
0 | 4 | -0.106% | -0.035% | 326.5% | 0.1%
0 | 5 | -0.071% | 0.035% | -0.083% | 0.1%
0 | 6 | 0% | 0.071% | -0.059% | 0.1%
Deviations of %error on test id #1. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 191.7% | 0% | 0% | 0.1%
1 | 1 | 417.2% | 225.5% | 0% | 0.1%
1 | 2 | 58.18% | -359% | 222.3% | 0.1%
1 | 3 | 12.61% | -45.57% | 162.7% | 0.1%
1 | 4 | 0.071% | -12.54% | 23.62% | 0.1%
1 | 5 | 0.053% | -0.018% | 4.246% | 0.1%
1 | 6 | 0.03% | -0.023% | 0.051% | 0.1%
1 | 7 | -0.041% | -0.071% | 0.014% | 0.1%
Deviations of %error on test id #3. Total 10 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 328.7% | 0% | 0% | 0.1%
3 | 1 | 292.1% | -36.61% | 0% | 0.1%
3 | 2 | 25.18% | -266.9% | 215.3% | 0.1%
3 | 3 | 25.18% | 0% | 114.2% | 0.1%
3 | 4 | 4.263% | -20.92% | 18.21% | 0.1%
3 | 5 | -0.009% | -4.272% | 9.812% | 0.1%
3 | 6 | 0% | 0.009% | 1.418% | 0.1%
3 | 7 | 0.089% | 0.089% | 0.027% | 0.1%
3 | 8 | 0.151% | 0.062% | 0.08% | 0.1%
3 | 9 | 0.08% | -0.071% | 0.107% | 0.1%
Deviations of %error on test id #4. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 198% | 0% | 0% | 0.1%
4 | 1 | 170.6% | -27.39% | 0% | 0.1%
4 | 2 | 25.03% | -145.6% | 131.2% | 0.1%
4 | 3 | 0.036% | -24.99% | 65.23% | 0.1%
4 | 4 | 0% | -0.036% | 8.355% | 0.1%
4 | 5 | 0.116% | 0.116% | 0.051% | 0.1%
4 | 6 | 0.169% | 0.053% | 0.095% | 0.1%
4 | 7 | 0.214% | 0.045% | 0.166% | 0.1%
4 | 8 | 0.089% | -0.125% | 0.157% | 0.1%
4 | 9 | 0.062% | -0.027% | 0.122% | 0.1%
4 | 10 | 0.223% | 0.161% | 0.125% | 0.1%
4 | 11 | 0.196% | -0.027% | 0.16% | 0.1%
4 | 12 | 0.116% | -0.08% | 0.178% | 0.1%
4 | 13 | 0.018% | -0.098% | 0.11% | 0.1%
Deviations of %error on test id #7. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 15.01% | 0% | 0% | 0.1%
7 | 1 | 0.055% | -14.95% | 0% | 0.1%
7 | 2 | 0% | -0.055% | 5.021% | 0.1%
7 | 3 | 0.046% | 0.046% | 0.034% | 0.1%
7 | 4 | 0.028% | -0.018% | 0.025% | 0.1%

## 28. 5th Harmonics
[<<]
Accuracy Test Report of '5th Harmonics'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / 5th Harmonics'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:11:30 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Harmonics' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
1 | 0 | 0 | 1
3. Test Cases
Total 1 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 1 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | 0.5Imax (20A) | 0.5Imax (20A) | 0.5Imax (20A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° | {UcUbUa, 5th-10%, IcIbIa, 5th-40%} | 12 | 20 |  | ±0.1% | ±1% | 0.012% | PASSED | 2023-07-27 16:58
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #1. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -2.679% | 0% | 0% | 0.1%
1 | 1 | -3.428% | -0.749% | 0% | 0.1%
1 | 2 | -3.758% | -0.33% | -3.288% | 0.1%
1 | 3 | -1.506% | 2.252% | -2.897% | 0.1%
1 | 4 | 0.473% | 1.979% | -1.597% | 0.1%
1 | 5 | 0.888% | 0.415% | -0.048% | 0.1%
1 | 6 | 0.557% | -0.331% | 0.639% | 0.1%
1 | 7 | 0.273% | -0.284% | 0.573% | 0.1%
1 | 8 | 0.056% | -0.217% | 0.295% | 0.1%
1 | 9 | -0.009% | -0.065% | 0.107% | 0.1%
1 | 10 | 0.015% | 0.024% | 0.021% | 0.1%
1 | 11 | 0.03% | 0.015% | 0.012% | 0.1%

## 29. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:11:31 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RQ | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | 0.044% | PASSED | 2023-07-27 16:59
Yes | 1 | RQ | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.061% | PASSED | 2023-07-27 16:59
Yes | 2 | RQ | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.073% | PASSED | 2023-07-27 16:59
Yes | 3 | RQ | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.062% | PASSED | 2023-07-27 17:00
Yes | 4 | RQ | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.062% | PASSED | 2023-07-27 17:00
Yes | 5 | RQ | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±1.1% | 0.059% | PASSED | 2023-07-27 17:00
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | 1.301% | 0% | 0% | 0.1%
0 | 1 | 0.321% | -0.98% | 0% | 0.1%
0 | 2 | 0.049% | -0.272% | 0.557% | 0.1%
0 | 3 | 0.044% | -0.005% | 0.138% | 0.1%
0 | 4 | 0.04% | -0.004% | 0.044% | 0.1%
0 | 5 | 0.049% | 0.009% | 0.044% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0.067% | 0% | 0% | 0.1%
1 | 1 | 0.076% | 0.009% | 0% | 0.1%
1 | 2 | 0.031% | -0.045% | 0.058% | 0.1%
1 | 3 | 0.076% | 0.045% | 0.061% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | 0.08% | 0% | 0% | 0.1%
2 | 1 | 0.058% | -0.022% | 0% | 0.1%
2 | 2 | 0.098% | 0.04% | 0.079% | 0.1%
2 | 3 | 0.062% | -0.036% | 0.073% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.102% | 0% | 0% | 0.1%
3 | 1 | 0.067% | -0.035% | 0% | 0.1%
3 | 2 | 0.053% | -0.014% | 0.074% | 0.1%
3 | 3 | 0.067% | 0.014% | 0.062% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.076% | 0% | 0% | 0.1%
4 | 1 | 0.049% | -0.027% | 0% | 0.1%
4 | 2 | 0.053% | 0.004% | 0.059% | 0.1%
4 | 3 | 0.085% | 0.032% | 0.062% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | 0.076% | 0% | 0% | 0.1%
5 | 1 | 0.062% | -0.014% | 0% | 0.1%
5 | 2 | 0.085% | 0.023% | 0.074% | 0.1%
5 | 3 | 0.031% | -0.054% | 0.059% | 0.1%

## 30. Voltage Variatio...
[<<]
Accuracy Test Report of 'Voltage Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Voltage Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:11:34 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Voltage Y (U1)' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
6 | 0 | 0 | 6
3. Test Cases
Total 6 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RQ | 0.8Ub (176V) | 0.8Ub (176V) | 0.8Ub (176V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | 0.238% | PASSED | 2023-07-27 17:01
Yes | 1 | RQ | 0.9Ub (198V) | 0.9Ub (198V) | 0.9Ub (198V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.235% | PASSED | 2023-07-27 17:01
Yes | 2 | RQ | 0.95Ub (209V) | 0.95Ub (209V) | 0.95Ub (209V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.238% | PASSED | 2023-07-27 17:02
Yes | 3 | RQ | 1.05Ub (231V) | 1.05Ub (231V) | 1.05Ub (231V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.199% | PASSED | 2023-07-27 17:02
Yes | 4 | RQ | 1.1Ub (242V) | 1.1Ub (242V) | 1.1Ub (242V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1% | 0.149% | PASSED | 2023-07-27 17:03
Yes | 5 | RQ | 1.15Ub (253V) | 1.15Ub (253V) | 1.15Ub (253V) | Ib (5A) | Ib (5A) | Ib (5A) | Fb (50Hz) | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±1.8% | 0.196% | PASSED | 2023-07-27 17:03
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -26.91% | 0% | 0% | 0.1%
0 | 1 | -10.59% | 16.32% | 0% | 0.1%
0 | 2 | -0.407% | 10.18% | -12.63% | 0.1%
0 | 3 | 0.187% | 0.594% | -3.602% | 0.1%
0 | 4 | 0.169% | -0.018% | -0.017% | 0.1%
0 | 5 | -0.098% | -0.267% | 0.086% | 0.1%
0 | 6 | -0.266% | -0.168% | -0.065% | 0.1%
0 | 7 | -0.142% | 0.124% | -0.169% | 0.1%
0 | 8 | 0.08% | 0.222% | -0.109% | 0.1%
0 | 9 | 0.142% | 0.062% | 0.027% | 0.1%
0 | 10 | -0.177% | -0.319% | 0.015% | 0.1%
0 | 11 | -0.115% | 0.062% | -0.05% | 0.1%
0 | 12 | 0.134% | 0.249% | -0.053% | 0.1%
0 | 13 | 0.223% | 0.089% | 0.081% | 0.1%
0 | 14 | 0.241% | 0.018% | 0.199% | 0.1%
0 | 15 | 0.25% | 0.009% | 0.238% | 0.1%
Deviations of %error on test id #1. Total 18 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -29.96% | 0% | 0% | 0.1%
1 | 1 | -17.58% | 12.38% | 0% | 0.1%
1 | 2 | -7.025% | 10.56% | -18.19% | 0.1%
1 | 3 | 0.267% | 7.292% | -8.113% | 0.1%
1 | 4 | 0.241% | -0.026% | -2.172% | 0.1%
1 | 5 | 0.134% | -0.107% | 0.214% | 0.1%
1 | 6 | -0.186% | -0.32% | 0.063% | 0.1%
1 | 7 | -0.248% | -0.062% | -0.1% | 0.1%
1 | 8 | -0.195% | 0.053% | -0.21% | 0.1%
1 | 9 | 0.071% | 0.266% | -0.124% | 0.1%
1 | 10 | 0.16% | 0.089% | 0.012% | 0.1%
1 | 11 | -0.009% | -0.169% | 0.074% | 0.1%
1 | 12 | -0.213% | -0.204% | -0.021% | 0.1%
1 | 13 | -0.071% | 0.142% | -0.098% | 0.1%
1 | 14 | 0.16% | 0.231% | -0.041% | 0.1%
1 | 15 | 0.25% | 0.09% | 0.113% | 0.1%
1 | 16 | 0.232% | -0.018% | 0.214% | 0.1%
1 | 17 | 0.223% | -0.009% | 0.235% | 0.1%
Deviations of %error on test id #2. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -30.24% | 0% | 0% | 0.1%
2 | 1 | -13.85% | 16.39% | 0% | 0.1%
2 | 2 | -4.951% | 8.901% | -16.35% | 0.1%
2 | 3 | 0.241% | 5.192% | -6.187% | 0.1%
2 | 4 | 0.267% | 0.026% | -1.481% | 0.1%
2 | 5 | 0.258% | -0.009% | 0.255% | 0.1%
2 | 6 | 0.08% | -0.178% | 0.202% | 0.1%
2 | 7 | -0.027% | -0.107% | 0.104% | 0.1%
2 | 8 | -0.044% | -0.017% | 0.003% | 0.1%
2 | 9 | -0.106% | -0.062% | -0.059% | 0.1%
2 | 10 | -0.732% | -0.626% | -0.294% | 0.1%
2 | 11 | -0.504% | 0.228% | -0.447% | 0.1%
2 | 12 | -0.142% | 0.362% | -0.459% | 0.1%
2 | 13 | 0.151% | 0.293% | -0.165% | 0.1%
2 | 14 | 0.241% | 0.09% | 0.083% | 0.1%
2 | 15 | 0.223% | -0.018% | 0.205% | 0.1%
2 | 16 | 0.25% | 0.027% | 0.238% | 0.1%
Deviations of %error on test id #3. Total 18 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -30.05% | 0% | 0% | 0.1%
3 | 1 | -19.59% | 10.47% | 0% | 0.1%
3 | 2 | -4.838% | 14.75% | -18.16% | 0.1%
3 | 3 | -1.824% | 3.014% | -8.749% | 0.1%
3 | 4 | 0.241% | 2.065% | -2.14% | 0.1%
3 | 5 | 0.267% | 0.026% | -0.439% | 0.1%
3 | 6 | 0.142% | -0.125% | 0.217% | 0.1%
3 | 7 | -0.027% | -0.169% | 0.127% | 0.1%
3 | 8 | -0.018% | 0.009% | 0.032% | 0.1%
3 | 9 | -0.071% | -0.053% | -0.039% | 0.1%
3 | 10 | -0.592% | -0.521% | -0.227% | 0.1%
3 | 11 | -0.566% | 0.026% | -0.41% | 0.1%
3 | 12 | -0.381% | 0.185% | -0.513% | 0.1%
3 | 13 | -0.169% | 0.212% | -0.372% | 0.1%
3 | 14 | 0.134% | 0.303% | -0.139% | 0.1%
3 | 15 | 0.214% | 0.08% | 0.06% | 0.1%
3 | 16 | 0.196% | -0.018% | 0.181% | 0.1%
3 | 17 | 0.187% | -0.009% | 0.199% | 0.1%
Deviations of %error on test id #4. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -35.67% | 0% | 0% | 0.1%
4 | 1 | -31.07% | 4.596% | 0% | 0.1%
4 | 2 | -15.77% | 15.31% | -27.5% | 0.1%
4 | 3 | -9.696% | 6.072% | -18.85% | 0.1%
4 | 4 | 0.027% | 9.723% | -8.479% | 0.1%
4 | 5 | 0.214% | 0.187% | -3.152% | 0.1%
4 | 6 | 0.214% | 0% | 0.152% | 0.1%
4 | 7 | 0.16% | -0.054% | 0.196% | 0.1%
4 | 8 | -0.204% | -0.364% | 0.057% | 0.1%
4 | 9 | -0.098% | 0.106% | -0.047% | 0.1%
4 | 10 | 0.036% | 0.134% | -0.089% | 0.1%
4 | 11 | -0.151% | -0.187% | -0.071% | 0.1%
4 | 12 | -0.027% | 0.124% | -0.047% | 0.1%
4 | 13 | 0.08% | 0.107% | -0.033% | 0.1%
4 | 14 | -0.036% | -0.116% | 0.006% | 0.1%
4 | 15 | -0.292% | -0.256% | -0.083% | 0.1%
4 | 16 | -0.08% | 0.212% | -0.136% | 0.1%
4 | 17 | 0.036% | 0.116% | -0.112% | 0.1%
4 | 18 | 0.178% | 0.142% | 0.045% | 0.1%
4 | 19 | 0.232% | 0.054% | 0.149% | 0.1%
Deviations of %error on test id #5. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -35.08% | 0% | 0% | 0.1%
5 | 1 | -26.44% | 8.643% | 0% | 0.1%
5 | 2 | -18.98% | 7.459% | -26.83% | 0.1%
5 | 3 | -4.361% | 14.62% | -16.59% | 0.1%
5 | 4 | -2.301% | 2.06% | -8.546% | 0.1%
5 | 5 | 0.232% | 2.533% | -2.143% | 0.1%
5 | 6 | 0.187% | -0.045% | -0.627% | 0.1%
5 | 7 | 0.142% | -0.045% | 0.187% | 0.1%
5 | 8 | 0% | -0.142% | 0.11% | 0.1%
5 | 9 | -0.027% | -0.027% | 0.038% | 0.1%
5 | 10 | -0.009% | 0.018% | -0.012% | 0.1%
5 | 11 | -0.169% | -0.16% | -0.068% | 0.1%
5 | 12 | -0.636% | -0.467% | -0.271% | 0.1%
5 | 13 | -0.601% | 0.035% | -0.469% | 0.1%
5 | 14 | -0.372% | 0.229% | -0.536% | 0.1%
5 | 15 | -0.186% | 0.186% | -0.386% | 0.1%
5 | 16 | 0.125% | 0.311% | -0.144% | 0.1%
5 | 17 | 0.142% | 0.017% | 0.027% | 0.1%
5 | 18 | 0.223% | 0.081% | 0.163% | 0.1%
5 | 19 | 0.223% | 0% | 0.196% | 0.1%

## 31. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 1.0'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 1.0'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:11:42 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.025% | PASSED | 2023-07-27 17:03
Yes | 1 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | -0.009% | PASSED | 2023-07-27 17:04
Yes | 2 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.022% | PASSED | 2023-07-27 17:04
Yes | 3 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.023% | PASSED | 2023-07-27 17:04
Yes | 4 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.009% | PASSED | 2023-07-27 17:05
Yes | 5 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.027% | PASSED | 2023-07-27 17:05
Yes | 6 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.056% | PASSED | 2023-07-27 17:05
Yes | 7 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.062% | PASSED | 2023-07-27 17:06
Yes | 8 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.043% | PASSED | 2023-07-27 17:06
Yes | 9 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.021% | PASSED | 2023-07-27 17:06
Yes | 10 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.067% | PASSED | 2023-07-27 17:06
Yes | 11 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.083% | PASSED | 2023-07-27 17:07
Yes | 12 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.061% | PASSED | 2023-07-27 17:07
Yes | 13 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.062% | PASSED | 2023-07-27 17:07
Yes | 14 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.047% | PASSED | 2023-07-27 17:08
Yes | 15 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.049% | PASSED | 2023-07-27 17:08
Yes | 16 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.087% | PASSED | 2023-07-27 17:08
Yes | 17 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.11% | PASSED | 2023-07-27 17:08
Yes | 18 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.103% | PASSED | 2023-07-27 17:09
Yes | 19 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.092% | PASSED | 2023-07-27 17:09
Yes | 20 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.108% | PASSED | 2023-07-27 17:09
Yes | 21 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.082% | PASSED | 2023-07-27 17:10
Yes | 22 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.093% | PASSED | 2023-07-27 17:10
Yes | 23 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.116% | PASSED | 2023-07-27 17:10
Yes | 24 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.113% | PASSED | 2023-07-27 17:10
Yes | 25 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.107% | PASSED | 2023-07-27 17:11
Yes | 26 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.09% | PASSED | 2023-07-27 17:11
Yes | 27 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.09% | PASSED | 2023-07-27 17:11
Yes | 28 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.112% | PASSED | 2023-07-27 17:12
Yes | 29 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.104% | PASSED | 2023-07-27 17:12
Yes | 30 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.134% | PASSED | 2023-07-27 17:12
Yes | 31 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.153% | PASSED | 2023-07-27 17:13
Yes | 32 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.104% | PASSED | 2023-07-27 17:13
Yes | 33 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.113% | PASSED | 2023-07-27 17:13
Yes | 34 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.131% | PASSED | 2023-07-27 17:14
Yes | 35 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.121% | PASSED | 2023-07-27 17:14
Yes | 36 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.162% | PASSED | 2023-07-27 17:14
Yes | 37 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.154% | PASSED | 2023-07-27 17:15
Yes | 38 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.156% | PASSED | 2023-07-27 17:15
Yes | 39 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.141% | PASSED | 2023-07-27 17:16
Yes | 40 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 1.0, sin(ϕ), ϕ=90° | 120° | 240° |  | 2 | 20 |  | ±0.1% | ±0.7% | 0.126% | PASSED | 2023-07-27 17:16
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -0.027% | 0% | 0% | 0.1%
0 | 1 | -0.036% | -0.009% | 0% | 0.1%
0 | 2 | -0.004% | 0.032% | -0.022% | 0.1%
0 | 3 | -0.036% | -0.032% | -0.025% | 0.1%
Deviations of %error on test id #1. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | 0.04% | 0% | 0% | 0.1%
1 | 1 | -0.036% | -0.076% | 0% | 0.1%
1 | 2 | -0.009% | 0.027% | -0.002% | 0.1%
1 | 3 | 0.018% | 0.027% | -0.009% | 0.1%
Deviations of %error on test id #2. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | 0.027% | 0% | 0% | 0.1%
2 | 1 | 0.04% | 0.013% | 0% | 0.1%
2 | 2 | 0.004% | -0.036% | 0.024% | 0.1%
2 | 3 | 0.022% | 0.018% | 0.022% | 0.1%
Deviations of %error on test id #3. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | 0.009% | 0% | 0% | 0.1%
3 | 1 | 0.044% | 0.035% | 0% | 0.1%
3 | 2 | 0.013% | -0.031% | 0.022% | 0.1%
3 | 3 | 0.013% | 0% | 0.023% | 0.1%
Deviations of %error on test id #4. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | 0.058% | 0% | 0% | 0.1%
4 | 1 | 0.004% | -0.054% | 0% | 0.1%
4 | 2 | 0.013% | 0.009% | 0.025% | 0.1%
4 | 3 | 0.009% | -0.004% | 0.009% | 0.1%
Deviations of %error on test id #5. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | 0.022% | 0% | 0% | 0.1%
5 | 1 | 0.053% | 0.031% | 0% | 0.1%
5 | 2 | 0.009% | -0.044% | 0.028% | 0.1%
5 | 3 | 0.018% | 0.009% | 0.027% | 0.1%
Deviations of %error on test id #6. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | 0.093% | 0% | 0% | 0.1%
6 | 1 | 0.036% | -0.057% | 0% | 0.1%
6 | 2 | 0.053% | 0.017% | 0.061% | 0.1%
6 | 3 | 0.08% | 0.027% | 0.056% | 0.1%
Deviations of %error on test id #7. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | 0.058% | 0% | 0% | 0.1%
7 | 1 | 0.053% | -0.005% | 0% | 0.1%
7 | 2 | 0.049% | -0.004% | 0.053% | 0.1%
7 | 3 | 0.085% | 0.036% | 0.062% | 0.1%
Deviations of %error on test id #8. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | 0.04% | 0% | 0% | 0.1%
8 | 1 | 0.036% | -0.004% | 0% | 0.1%
8 | 2 | 0.049% | 0.013% | 0.042% | 0.1%
8 | 3 | 0.044% | -0.005% | 0.043% | 0.1%
Deviations of %error on test id #9. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | 0.018% | 0% | 0% | 0.1%
9 | 1 | 0.036% | 0.018% | 0% | 0.1%
9 | 2 | 0.013% | -0.023% | 0.022% | 0.1%
9 | 3 | 0.013% | 0% | 0.021% | 0.1%
Deviations of %error on test id #10. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | 0.049% | 0% | 0% | 0.1%
10 | 1 | 0.062% | 0.013% | 0% | 0.1%
10 | 2 | 0.067% | 0.005% | 0.059% | 0.1%
10 | 3 | 0.071% | 0.004% | 0.067% | 0.1%
Deviations of %error on test id #11. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | 0.044% | 0% | 0% | 0.1%
11 | 1 | 0.093% | 0.049% | 0% | 0.1%
11 | 2 | 0.093% | 0% | 0.077% | 0.1%
11 | 3 | 0.062% | -0.031% | 0.083% | 0.1%
Deviations of %error on test id #12. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | 0.12% | 0% | 0% | 0.1%
12 | 1 | 0.049% | -0.071% | 0% | 0.1%
12 | 2 | 0.076% | 0.027% | 0.082% | 0.1%
12 | 3 | 0.058% | -0.018% | 0.061% | 0.1%
Deviations of %error on test id #13. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | 0.053% | 0% | 0% | 0.1%
13 | 1 | 0.049% | -0.004% | 0% | 0.1%
13 | 2 | 0.076% | 0.027% | 0.059% | 0.1%
13 | 3 | 0.062% | -0.014% | 0.062% | 0.1%
Deviations of %error on test id #14. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | 0.049% | 0% | 0% | 0.1%
14 | 1 | 0.053% | 0.004% | 0% | 0.1%
14 | 2 | 0.049% | -0.004% | 0.05% | 0.1%
14 | 3 | 0.04% | -0.009% | 0.047% | 0.1%
Deviations of %error on test id #15. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | 0.062% | 0% | 0% | 0.1%
15 | 1 | 0.067% | 0.005% | 0% | 0.1%
15 | 2 | 0.027% | -0.04% | 0.052% | 0.1%
15 | 3 | 0.053% | 0.026% | 0.049% | 0.1%
Deviations of %error on test id #16. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | 0.089% | 0% | 0% | 0.1%
16 | 1 | 0.053% | -0.036% | 0% | 0.1%
16 | 2 | 0.12% | 0.067% | 0.087% | 0.1%
16 | 3 | 0.089% | -0.031% | 0.087% | 0.1%
Deviations of %error on test id #17. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | 0.089% | 0% | 0% | 0.1%
17 | 1 | 0.142% | 0.053% | 0% | 0.1%
17 | 2 | 0.098% | -0.044% | 0.11% | 0.1%
17 | 3 | 0.089% | -0.009% | 0.11% | 0.1%
Deviations of %error on test id #18. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | 0.098% | 0% | 0% | 0.1%
18 | 1 | 0.089% | -0.009% | 0% | 0.1%
18 | 2 | 0.085% | -0.004% | 0.091% | 0.1%
18 | 3 | 0.134% | 0.049% | 0.103% | 0.1%
Deviations of %error on test id #19. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | 0.058% | 0% | 0% | 0.1%
19 | 1 | 0.102% | 0.044% | 0% | 0.1%
19 | 2 | 0.089% | -0.013% | 0.083% | 0.1%
19 | 3 | 0.085% | -0.004% | 0.092% | 0.1%
Deviations of %error on test id #20. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | 0.071% | 0% | 0% | 0.1%
20 | 1 | 0.116% | 0.045% | 0% | 0.1%
20 | 2 | 0.089% | -0.027% | 0.092% | 0.1%
20 | 3 | 0.12% | 0.031% | 0.108% | 0.1%
Deviations of %error on test id #21. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | 0.085% | 0% | 0% | 0.1%
21 | 1 | 0.071% | -0.014% | 0% | 0.1%
21 | 2 | 0.085% | 0.014% | 0.08% | 0.1%
21 | 3 | 0.089% | 0.004% | 0.082% | 0.1%
Deviations of %error on test id #22. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | 0.062% | 0% | 0% | 0.1%
22 | 1 | 0.089% | 0.027% | 0% | 0.1%
22 | 2 | 0.089% | 0% | 0.08% | 0.1%
22 | 3 | 0.102% | 0.013% | 0.093% | 0.1%
Deviations of %error on test id #23. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | 0.129% | 0% | 0% | 0.1%
23 | 1 | 0.138% | 0.009% | 0% | 0.1%
23 | 2 | 0.085% | -0.053% | 0.117% | 0.1%
23 | 3 | 0.125% | 0.04% | 0.116% | 0.1%
Deviations of %error on test id #24. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | 0.116% | 0% | 0% | 0.1%
24 | 1 | 0.116% | 0% | 0% | 0.1%
24 | 2 | 0.134% | 0.018% | 0.122% | 0.1%
24 | 3 | 0.089% | -0.045% | 0.113% | 0.1%
Deviations of %error on test id #25. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | 0.098% | 0% | 0% | 0.1%
25 | 1 | 0.134% | 0.036% | 0% | 0.1%
25 | 2 | 0.089% | -0.045% | 0.107% | 0.1%
25 | 3 | 0.098% | 0.009% | 0.107% | 0.1%
Deviations of %error on test id #26. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | 0.102% | 0% | 0% | 0.1%
26 | 1 | 0.08% | -0.022% | 0% | 0.1%
26 | 2 | 0.093% | 0.013% | 0.092% | 0.1%
26 | 3 | 0.098% | 0.005% | 0.09% | 0.1%
Deviations of %error on test id #27. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | 0.102% | 0% | 0% | 0.1%
27 | 1 | 0.093% | -0.009% | 0% | 0.1%
27 | 2 | 0.093% | 0% | 0.096% | 0.1%
27 | 3 | 0.085% | -0.008% | 0.09% | 0.1%
Deviations of %error on test id #28. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | 0.12% | 0% | 0% | 0.1%
28 | 1 | 0.125% | 0.005% | 0% | 0.1%
28 | 2 | 0.085% | -0.04% | 0.11% | 0.1%
28 | 3 | 0.125% | 0.04% | 0.112% | 0.1%
Deviations of %error on test id #29. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | 0.129% | 0% | 0% | 0.1%
29 | 1 | 0.089% | -0.04% | 0% | 0.1%
29 | 2 | 0.093% | 0.004% | 0.104% | 0.1%
29 | 3 | 0.12% | 0.027% | 0.101% | 0.1%
29 | 4 | 0.098% | -0.022% | 0.104% | 0.1%
Deviations of %error on test id #30. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | 0.12% | 0% | 0% | 0.1%
30 | 1 | 0.125% | 0.005% | 0% | 0.1%
30 | 2 | 0.107% | -0.018% | 0.117% | 0.1%
30 | 3 | 0.169% | 0.062% | 0.134% | 0.1%
Deviations of %error on test id #31. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | 0.129% | 0% | 0% | 0.1%
31 | 1 | 0.156% | 0.027% | 0% | 0.1%
31 | 2 | 0.129% | -0.027% | 0.138% | 0.1%
31 | 3 | 0.174% | 0.045% | 0.153% | 0.1%
Deviations of %error on test id #32. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | 0.12% | 0% | 0% | 0.1%
32 | 1 | 0.089% | -0.031% | 0% | 0.1%
32 | 2 | 0.098% | 0.009% | 0.102% | 0.1%
32 | 3 | 0.125% | 0.027% | 0.104% | 0.1%
Deviations of %error on test id #33. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | 0.134% | 0% | 0% | 0.1%
33 | 1 | 0.089% | -0.045% | 0% | 0.1%
33 | 2 | 0.12% | 0.031% | 0.114% | 0.1%
33 | 3 | 0.089% | -0.031% | 0.099% | 0.1%
33 | 4 | 0.129% | 0.04% | 0.113% | 0.1%
Deviations of %error on test id #34. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | 0.085% | 0% | 0% | 0.1%
34 | 1 | 0.125% | 0.04% | 0% | 0.1%
34 | 2 | 0.129% | 0.004% | 0.113% | 0.1%
34 | 3 | 0.138% | 0.009% | 0.131% | 0.1%
Deviations of %error on test id #35. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | 0.08% | 0% | 0% | 0.1%
35 | 1 | 0.129% | 0.049% | 0% | 0.1%
35 | 2 | 0.093% | -0.036% | 0.101% | 0.1%
35 | 3 | 0.142% | 0.049% | 0.121% | 0.1%
Deviations of %error on test id #36. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | 0.138% | 0% | 0% | 0.1%
36 | 1 | 0.187% | 0.049% | 0% | 0.1%
36 | 2 | 0.147% | -0.04% | 0.157% | 0.1%
36 | 3 | 0.151% | 0.004% | 0.162% | 0.1%
Deviations of %error on test id #37. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | 0.125% | 0% | 0% | 0.1%
37 | 1 | 0.142% | 0.017% | 0% | 0.1%
37 | 2 | 0.174% | 0.032% | 0.147% | 0.1%
37 | 3 | 0.147% | -0.027% | 0.154% | 0.1%
Deviations of %error on test id #38. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | 0.183% | 0% | 0% | 0.1%
38 | 1 | 0.151% | -0.032% | 0% | 0.1%
38 | 2 | 0.178% | 0.027% | 0.171% | 0.1%
38 | 3 | 0.138% | -0.04% | 0.156% | 0.1%
Deviations of %error on test id #39. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | 0.12% | 0% | 0% | 0.1%
39 | 1 | 0.165% | 0.045% | 0% | 0.1%
39 | 2 | 0.129% | -0.036% | 0.138% | 0.1%
39 | 3 | 0.129% | 0% | 0.141% | 0.1%
Deviations of %error on test id #40. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | 0.12% | 0% | 0% | 0.1%
40 | 1 | 0.125% | 0.005% | 0% | 0.1%
40 | 2 | 0.129% | 0.004% | 0.125% | 0.1%
40 | 3 | 0.125% | -0.004% | 0.126% | 0.1%

## 32. Frequency Varia...
[<<]
Accuracy Test Report of 'Frequency Variation, PF = 0.5L'
Path in workspace, 'Smart Meter Test / 3P4W Meter / IEC 62052-23,24 Reactive Power Accuracy / Class 0.5S - Reverse / Frequency Variation, PF = 0.5L'
Workspace file path is 'D:\GitLab EM\smart_meter ticket 212\3p\rl78i1c (512)\02.Tests\Metrology\AccuracyTest\e2studio\20230728 SMAT 3PH4W NPL e2 studio SM Active Reactive Accuracy.zip'
Generated on 7/28/2023 3:11:59 PM by version Alpha5 - Build 2023 Jul. 27
1. Accuracy Chart
X axis is 'Freq.' column. Y axis is 'Error (%)' column.
2. Overall Test Result
PASSED
Passed | Failed | Inconclusive | Total
41 | 0 | 0 | 41
3. Test Cases
Total 41 test case(s) in below table
'Id' is identifier number of a test case, this number may not continous
'N' column specifies number of pulses which test equipment will capture to average %error before reply a %error reading to the test tool.
'M' column specifies maximum %error readings will be recorded & checked deviation for a test case
'D' column specifies number of first few %error readings from meter will be discarded before recorded into deviation table
Enable | Id | Dir. | Voltage R (U2) | Voltage Y (U1) | Voltage B (U3) | Current R (I2) | Current Y (I1) | Current B (I3) | Freq. | Phase R (U2I2) | Phase Y (U1I1) | Phase B (U3I3) | Phase YR (U1U2) | Phase YB (U1U3) | Harmonics | N | M | D | Deviation Threshold | Error Limit (%) | Error (%) | Test Result | Test Date
Yes | 0 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.071% | PASSED | 2023-07-27 17:17
Yes | 1 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 45.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.083% | PASSED | 2023-07-27 17:17
Yes | 2 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.071% | PASSED | 2023-07-27 17:18
Yes | 3 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 46.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.038% | PASSED | 2023-07-27 17:19
Yes | 4 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.089% | PASSED | 2023-07-27 17:20
Yes | 5 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 47.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.095% | PASSED | 2023-07-27 17:20
Yes | 6 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.154% | PASSED | 2023-07-27 17:21
Yes | 7 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 48.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.187% | PASSED | 2023-07-27 17:21
Yes | 8 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.178% | PASSED | 2023-07-27 17:22
Yes | 9 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 49.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.19% | PASSED | 2023-07-27 17:23
Yes | 10 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.154% | PASSED | 2023-07-27 17:23
Yes | 11 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 50.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.187% | PASSED | 2023-07-27 17:24
Yes | 12 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.187% | PASSED | 2023-07-27 17:25
Yes | 13 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 51.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.22% | PASSED | 2023-07-27 17:25
Yes | 14 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.244% | PASSED | 2023-07-27 17:26
Yes | 15 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 52.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.232% | PASSED | 2023-07-27 17:26
Yes | 16 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.238% | PASSED | 2023-07-27 17:27
Yes | 17 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 53.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.315% | PASSED | 2023-07-27 17:28
Yes | 18 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.312% | PASSED | 2023-07-27 17:28
Yes | 19 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 54.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.285% | PASSED | 2023-07-27 17:29
Yes | 20 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.306% | PASSED | 2023-07-27 17:30
Yes | 21 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 55.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.378% | PASSED | 2023-07-27 17:30
Yes | 22 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.3% | PASSED | 2023-07-27 17:31
Yes | 23 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 56.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.402% | PASSED | 2023-07-27 17:32
Yes | 24 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.378% | PASSED | 2023-07-27 17:32
Yes | 25 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 57.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.363% | PASSED | 2023-07-27 17:33
Yes | 26 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.617% | PASSED | 2023-07-27 17:34
Yes | 27 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 58.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.312% | PASSED | 2023-07-27 17:34
Yes | 28 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.408% | PASSED | 2023-07-27 17:35
Yes | 29 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 59.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.782% | PASSED | 2023-07-28 09:58
Yes | 30 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.384% | PASSED | 2023-07-28 11:19
Yes | 31 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 60.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.47% | PASSED | 2023-07-27 17:38
Yes | 32 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.345% | PASSED | 2023-07-27 17:38
Yes | 33 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 61.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.333% | PASSED | 2023-07-28 11:19
Yes | 34 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.437% | PASSED | 2023-07-28 10:24
Yes | 35 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 62.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | -0.077% | PASSED | 2023-07-28 11:20
Yes | 36 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.39% | PASSED | 2023-07-28 11:20
Yes | 37 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 63.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.602% | PASSED | 2023-07-28 10:32
Yes | 38 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.443% | PASSED | 2023-07-28 11:21
Yes | 39 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 64.5Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.479% | PASSED | 2023-07-28 11:22
Yes | 40 | RQ | Ub (220V) | Ub (220V) | Ub (220V) | Ib (5A) | Ib (5A) | Ib (5A) | 65Hz | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 0.5L, sin(ϕ), ϕ=30° | 120° | 240° |  | 1 | 20 |  | ±0.1% | ±0.8% | 0.713% | PASSED | 2023-07-28 10:03
4. Deviations of %error on test cases
Below tables show all %error responses from test equipment and analyze deviation of %error readings for all test cases
'Test Id' is test case identifier which some %error readings will belong to
'Average % Error' is only valid after 3 times of reading. First few may be 0%, not valid.
'Red cells in 'Deviation (%)' and 'Deviation Thres. (%)' columns indicates 'Deviation (%)' > 'Deviation Thres. (%)', when this contition becomes true, averaging of %error by 3 was restarted.
Deviations of %error on test id #0. Total 14 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
0 | 0 | -65.22% | 0% | 0% | 0.1%
0 | 1 | -8.707% | 56.51% | 0% | 0.1%
0 | 2 | 0.027% | 8.734% | -24.63% | 0.1%
0 | 3 | -1.082% | -1.109% | -3.254% | 0.1%
0 | 4 | -0.239% | 0.843% | -0.431% | 0.1%
0 | 5 | -0.872% | -0.633% | -0.731% | 0.1%
0 | 6 | 0.116% | 0.988% | -0.332% | 0.1%
0 | 7 | 0.16% | 0.044% | -0.199% | 0.1%
0 | 8 | 0.518% | 0.358% | 0.265% | 0.1%
0 | 9 | 0.312% | -0.206% | 0.33% | 0.1%
0 | 10 | 0.071% | -0.241% | 0.3% | 0.1%
0 | 11 | 0.071% | 0% | 0.151% | 0.1%
0 | 12 | 0.08% | 0.009% | 0.074% | 0.1%
0 | 13 | 0.062% | -0.018% | 0.071% | 0.1%
Deviations of %error on test id #1. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
1 | 0 | -64.08% | 0% | 0% | 0.1%
1 | 1 | -6.832% | 57.25% | 0% | 0.1%
1 | 2 | -0.098% | 6.734% | -23.67% | 0.1%
1 | 3 | -0.951% | -0.853% | -2.627% | 0.1%
1 | 4 | -0.389% | 0.562% | -0.479% | 0.1%
1 | 5 | 0.08% | 0.469% | -0.42% | 0.1%
1 | 6 | 0.089% | 0.009% | -0.073% | 0.1%
1 | 7 | 0.089% | 0% | 0.086% | 0.1%
1 | 8 | 0.071% | -0.018% | 0.083% | 0.1%
Deviations of %error on test id #2. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
2 | 0 | -65.83% | 0% | 0% | 0.1%
2 | 1 | -13.18% | 52.64% | 0% | 0.1%
2 | 2 | -0.169% | 13.01% | -26.39% | 0.1%
2 | 3 | -1.107% | -0.938% | -4.819% | 0.1%
2 | 4 | -1.54% | -0.433% | -0.939% | 0.1%
2 | 5 | -0.451% | 1.089% | -1.033% | 0.1%
2 | 6 | 0.116% | 0.567% | -0.625% | 0.1%
2 | 7 | 0.276% | 0.16% | -0.02% | 0.1%
2 | 8 | 0.788% | 0.512% | 0.393% | 0.1%
2 | 9 | 0.116% | -0.672% | 0.393% | 0.1%
2 | 10 | 0.08% | -0.036% | 0.328% | 0.1%
2 | 11 | 0.053% | -0.027% | 0.083% | 0.1%
2 | 12 | 0.08% | 0.027% | 0.071% | 0.1%
Deviations of %error on test id #3. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
3 | 0 | -38.87% | 0% | 0% | 0.1%
3 | 1 | 0.071% | 38.94% | 0% | 0.1%
3 | 2 | -1.064% | -1.135% | -13.29% | 0.1%
3 | 3 | -1.454% | -0.39% | -0.816% | 0.1%
3 | 4 | -0.89% | 0.564% | -1.136% | 0.1%
3 | 5 | 0.009% | 0.899% | -0.778% | 0.1%
3 | 6 | 0.044% | 0.035% | -0.279% | 0.1%
3 | 7 | 0.815% | 0.771% | 0.289% | 0.1%
3 | 8 | 0.151% | -0.664% | 0.337% | 0.1%
3 | 9 | 0.018% | -0.133% | 0.328% | 0.1%
3 | 10 | 0.018% | 0% | 0.062% | 0.1%
3 | 11 | 0.053% | 0.035% | 0.03% | 0.1%
3 | 12 | 0.044% | -0.009% | 0.038% | 0.1%
Deviations of %error on test id #4. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
4 | 0 | -54.86% | 0% | 0% | 0.1%
4 | 1 | -2.715% | 52.14% | 0% | 0.1%
4 | 2 | -0.627% | 2.088% | -19.4% | 0.1%
4 | 3 | -1.497% | -0.87% | -1.613% | 0.1%
4 | 4 | -1.09% | 0.407% | -1.071% | 0.1%
4 | 5 | 0.027% | 1.117% | -0.853% | 0.1%
4 | 6 | 0.107% | 0.08% | -0.319% | 0.1%
4 | 7 | 0.843% | 0.736% | 0.326% | 0.1%
4 | 8 | 0.223% | -0.62% | 0.391% | 0.1%
4 | 9 | 0.053% | -0.17% | 0.373% | 0.1%
4 | 10 | 0.071% | 0.018% | 0.116% | 0.1%
4 | 11 | 0.089% | 0.018% | 0.071% | 0.1%
4 | 12 | 0.107% | 0.018% | 0.089% | 0.1%
Deviations of %error on test id #5. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
5 | 0 | -55.84% | 0% | 0% | 0.1%
5 | 1 | -1.721% | 54.12% | 0% | 0.1%
5 | 2 | -0.504% | 1.217% | -19.35% | 0.1%
5 | 3 | -0.645% | -0.141% | -0.957% | 0.1%
5 | 4 | -1.134% | -0.489% | -0.761% | 0.1%
5 | 5 | -0.478% | 0.656% | -0.752% | 0.1%
5 | 6 | 0.169% | 0.647% | -0.481% | 0.1%
5 | 7 | 0.393% | 0.224% | 0.028% | 0.1%
5 | 8 | 0.788% | 0.395% | 0.45% | 0.1%
5 | 9 | 0.142% | -0.646% | 0.441% | 0.1%
5 | 10 | 0.071% | -0.071% | 0.334% | 0.1%
5 | 11 | 0.116% | 0.045% | 0.11% | 0.1%
5 | 12 | 0.098% | -0.018% | 0.095% | 0.1%
Deviations of %error on test id #6. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
6 | 0 | -68.94% | 0% | 0% | 0.1%
6 | 1 | -21.92% | 47.02% | 0% | 0.1%
6 | 2 | -0.372% | 21.55% | -30.41% | 0.1%
6 | 3 | -1.583% | -1.211% | -7.959% | 0.1%
6 | 4 | -0.389% | 1.194% | -0.781% | 0.1%
6 | 5 | 0.125% | 0.514% | -0.616% | 0.1%
6 | 6 | 0.16% | 0.035% | -0.035% | 0.1%
6 | 7 | 0.142% | -0.018% | 0.142% | 0.1%
6 | 8 | 0.16% | 0.018% | 0.154% | 0.1%
Deviations of %error on test id #7. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
7 | 0 | -17.37% | 0% | 0% | 0.1%
7 | 1 | -0.027% | 17.34% | 0% | 0.1%
7 | 2 | -0.986% | -0.959% | -6.128% | 0.1%
7 | 3 | -1.506% | -0.52% | -0.84% | 0.1%
7 | 4 | -0.151% | 1.355% | -0.881% | 0.1%
7 | 5 | 0.178% | 0.329% | -0.493% | 0.1%
7 | 6 | 0.689% | 0.511% | 0.239% | 0.1%
7 | 7 | 0.581% | -0.108% | 0.483% | 0.1%
7 | 8 | 0.196% | -0.385% | 0.489% | 0.1%
7 | 9 | 0.205% | 0.009% | 0.327% | 0.1%
7 | 10 | 0.178% | -0.027% | 0.193% | 0.1%
7 | 11 | 0.178% | 0% | 0.187% | 0.1%
Deviations of %error on test id #8. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
8 | 0 | -58.07% | 0% | 0% | 0.1%
8 | 1 | -4.214% | 53.85% | 0% | 0.1%
8 | 2 | -0.601% | 3.613% | -20.96% | 0.1%
8 | 3 | -0.557% | 0.044% | -1.791% | 0.1%
8 | 4 | 0.062% | 0.619% | -0.365% | 0.1%
8 | 5 | 0.178% | 0.116% | -0.106% | 0.1%
8 | 6 | 0.178% | 0% | 0.139% | 0.1%
8 | 7 | 0.16% | -0.018% | 0.172% | 0.1%
8 | 8 | 0.196% | 0.036% | 0.178% | 0.1%
Deviations of %error on test id #9. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
9 | 0 | -51.42% | 0% | 0% | 0.1%
9 | 1 | -1.54% | 49.88% | 0% | 0.1%
9 | 2 | -0.451% | 1.089% | -17.8% | 0.1%
9 | 3 | -1.644% | -1.193% | -1.212% | 0.1%
9 | 4 | -0.522% | 1.122% | -0.872% | 0.1%
9 | 5 | 0.196% | 0.718% | -0.657% | 0.1%
9 | 6 | 0.473% | 0.277% | 0.049% | 0.1%
9 | 7 | 0.77% | 0.297% | 0.48% | 0.1%
9 | 8 | 0.205% | -0.565% | 0.483% | 0.1%
9 | 9 | 0.187% | -0.018% | 0.387% | 0.1%
9 | 10 | 0.187% | 0% | 0.193% | 0.1%
9 | 11 | 0.196% | 0.009% | 0.19% | 0.1%
Deviations of %error on test id #10. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
10 | 0 | -60.04% | 0% | 0% | 0.1%
10 | 1 | 0.5% | 60.54% | 0% | 0.1%
10 | 2 | -0.328% | -0.828% | -19.96% | 0.1%
10 | 3 | -0.495% | -0.167% | -0.108% | 0.1%
10 | 4 | -1.194% | -0.699% | -0.672% | 0.1%
10 | 5 | -0.009% | 1.185% | -0.566% | 0.1%
10 | 6 | 0.312% | 0.321% | -0.297% | 0.1%
10 | 7 | 0.843% | 0.531% | 0.382% | 0.1%
10 | 8 | 0.464% | -0.379% | 0.54% | 0.1%
10 | 9 | 0.142% | -0.322% | 0.483% | 0.1%
10 | 10 | 0.151% | 0.009% | 0.252% | 0.1%
10 | 11 | 0.151% | 0% | 0.148% | 0.1%
10 | 12 | 0.16% | 0.009% | 0.154% | 0.1%
Deviations of %error on test id #11. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
11 | 0 | -33.61% | 0% | 0% | 0.1%
11 | 1 | 0.384% | 34% | 0% | 0.1%
11 | 2 | -0.881% | -1.265% | -11.37% | 0.1%
11 | 3 | -0.566% | 0.315% | -0.354% | 0.1%
11 | 4 | -0.968% | -0.402% | -0.805% | 0.1%
11 | 5 | 0.25% | 1.218% | -0.428% | 0.1%
11 | 6 | 0.375% | 0.125% | -0.114% | 0.1%
11 | 7 | 1.024% | 0.649% | 0.55% | 0.1%
11 | 8 | 0.294% | -0.73% | 0.564% | 0.1%
11 | 9 | 0.205% | -0.089% | 0.508% | 0.1%
11 | 10 | 0.187% | -0.018% | 0.229% | 0.1%
11 | 11 | 0.169% | -0.018% | 0.187% | 0.1%
Deviations of %error on test id #12. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
12 | 0 | -37.31% | 0% | 0% | 0.1%
12 | 1 | 0.062% | 37.37% | 0% | 0.1%
12 | 2 | -0.881% | -0.943% | -12.71% | 0.1%
12 | 3 | -1.626% | -0.745% | -0.815% | 0.1%
12 | 4 | 0.142% | 1.768% | -0.788% | 0.1%
12 | 5 | 0.241% | 0.099% | -0.414% | 0.1%
12 | 6 | 1.024% | 0.783% | 0.469% | 0.1%
12 | 7 | 0.339% | -0.685% | 0.535% | 0.1%
12 | 8 | 0.205% | -0.134% | 0.523% | 0.1%
12 | 9 | 0.169% | -0.036% | 0.238% | 0.1%
12 | 10 | 0.187% | 0.018% | 0.187% | 0.1%
12 | 11 | 0.205% | 0.018% | 0.187% | 0.1%
Deviations of %error on test id #13. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
13 | 0 | -41.08% | 0% | 0% | 0.1%
13 | 1 | -0.539% | 40.54% | 0% | 0.1%
13 | 2 | -0.846% | -0.307% | -14.16% | 0.1%
13 | 3 | -0.398% | 0.448% | -0.594% | 0.1%
13 | 4 | 0.196% | 0.594% | -0.349% | 0.1%
13 | 5 | 0.214% | 0.018% | 0.004% | 0.1%
13 | 6 | 0.241% | 0.027% | 0.217% | 0.1%
13 | 7 | 0.205% | -0.036% | 0.22% | 0.1%
Deviations of %error on test id #14. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
14 | 0 | -22.57% | 0% | 0% | 0.1%
14 | 1 | -0.363% | 22.2% | 0% | 0.1%
14 | 2 | -1.055% | -0.692% | -7.995% | 0.1%
14 | 3 | 0.062% | 1.117% | -0.452% | 0.1%
14 | 4 | 0.25% | 0.188% | -0.248% | 0.1%
14 | 5 | 0.25% | 0% | 0.187% | 0.1%
14 | 6 | 0.241% | -0.009% | 0.247% | 0.1%
14 | 7 | 0.241% | 0% | 0.244% | 0.1%
Deviations of %error on test id #15. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
15 | 0 | -1.359% | 0% | 0% | 0.1%
15 | 1 | -0.46% | 0.899% | 0% | 0.1%
15 | 2 | -1.747% | -1.287% | -1.189% | 0.1%
15 | 3 | -0.292% | 1.455% | -0.833% | 0.1%
15 | 4 | 0.205% | 0.497% | -0.611% | 0.1%
15 | 5 | 0.788% | 0.583% | 0.234% | 0.1%
15 | 6 | 0.509% | -0.279% | 0.501% | 0.1%
15 | 7 | 0.223% | -0.286% | 0.507% | 0.1%
15 | 8 | 0.223% | 0% | 0.318% | 0.1%
15 | 9 | 0.232% | 0.009% | 0.226% | 0.1%
15 | 10 | 0.241% | 0.009% | 0.232% | 0.1%
Deviations of %error on test id #16. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
16 | 0 | -25.53% | 0% | 0% | 0.1%
16 | 1 | 0.232% | 25.76% | 0% | 0.1%
16 | 2 | -0.724% | -0.956% | -8.673% | 0.1%
16 | 3 | -0.872% | -0.148% | -0.455% | 0.1%
16 | 4 | -0.151% | 0.721% | -0.582% | 0.1%
16 | 5 | 0.482% | 0.633% | -0.18% | 0.1%
16 | 6 | 0.879% | 0.397% | 0.403% | 0.1%
16 | 7 | 0.689% | -0.19% | 0.683% | 0.1%
16 | 8 | 0.267% | -0.422% | 0.612% | 0.1%
16 | 9 | 0.241% | -0.026% | 0.399% | 0.1%
16 | 10 | 0.232% | -0.009% | 0.247% | 0.1%
16 | 11 | 0.241% | 0.009% | 0.238% | 0.1%
Deviations of %error on test id #17. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
17 | 0 | -61.72% | 0% | 0% | 0.1%
17 | 1 | -1.583% | 60.14% | 0% | 0.1%
17 | 2 | -0.363% | 1.22% | -21.22% | 0.1%
17 | 3 | -1.858% | -1.495% | -1.268% | 0.1%
17 | 4 | -0.062% | 1.796% | -0.761% | 0.1%
17 | 5 | 0.321% | 0.383% | -0.533% | 0.1%
17 | 6 | 1.069% | 0.748% | 0.443% | 0.1%
17 | 7 | 0.527% | -0.542% | 0.639% | 0.1%
17 | 8 | 0.303% | -0.224% | 0.633% | 0.1%
17 | 9 | 0.312% | 0.009% | 0.381% | 0.1%
17 | 10 | 0.312% | 0% | 0.309% | 0.1%
17 | 11 | 0.321% | 0.009% | 0.315% | 0.1%
Deviations of %error on test id #18. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
18 | 0 | -51.29% | 0% | 0% | 0.1%
18 | 1 | 0.285% | 51.58% | 0% | 0.1%
18 | 2 | -0.284% | -0.569% | -17.1% | 0.1%
18 | 3 | -0.275% | 0.009% | -0.091% | 0.1%
18 | 4 | -0.732% | -0.457% | -0.43% | 0.1%
18 | 5 | 0.5% | 1.232% | -0.169% | 0.1%
18 | 6 | 0.689% | 0.189% | 0.152% | 0.1%
18 | 7 | 1.051% | 0.362% | 0.747% | 0.1%
18 | 8 | 0.357% | -0.694% | 0.699% | 0.1%
18 | 9 | 0.321% | -0.036% | 0.576% | 0.1%
18 | 10 | 0.294% | -0.027% | 0.324% | 0.1%
18 | 11 | 0.321% | 0.027% | 0.312% | 0.1%
Deviations of %error on test id #19. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
19 | 0 | -9.362% | 0% | 0% | 0.1%
19 | 1 | -0.275% | 9.087% | 0% | 0.1%
19 | 2 | -1.695% | -1.42% | -3.777% | 0.1%
19 | 3 | -0.142% | 1.553% | -0.704% | 0.1%
19 | 4 | 0.303% | 0.445% | -0.511% | 0.1%
19 | 5 | 0.96% | 0.657% | 0.374% | 0.1%
19 | 6 | 0.5% | -0.46% | 0.588% | 0.1%
19 | 7 | 0.267% | -0.233% | 0.576% | 0.1%
19 | 8 | 0.285% | 0.018% | 0.351% | 0.1%
19 | 9 | 0.285% | 0% | 0.279% | 0.1%
19 | 10 | 0.285% | 0% | 0.285% | 0.1%
Deviations of %error on test id #20. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
20 | 0 | -59.41% | 0% | 0% | 0.1%
20 | 1 | -0.248% | 59.17% | 0% | 0.1%
20 | 2 | -0.776% | -0.528% | -20.15% | 0.1%
20 | 3 | -0.513% | 0.263% | -0.512% | 0.1%
20 | 4 | 0.232% | 0.745% | -0.352% | 0.1%
20 | 5 | 0.303% | 0.071% | 0.007% | 0.1%
20 | 6 | 0.303% | 0% | 0.279% | 0.1%
20 | 7 | 0.312% | 0.009% | 0.306% | 0.1%
Deviations of %error on test id #21. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
21 | 0 | -7.056% | 0% | 0% | 0.1%
21 | 1 | -0.071% | 6.985% | 0% | 0.1%
21 | 2 | -0.018% | 0.053% | -2.382% | 0.1%
21 | 3 | -0.907% | -0.889% | -0.332% | 0.1%
21 | 4 | 0.428% | 1.335% | -0.166% | 0.1%
21 | 5 | 0.581% | 0.153% | 0.034% | 0.1%
21 | 6 | 1.16% | 0.579% | 0.723% | 0.1%
21 | 7 | 0.509% | -0.651% | 0.75% | 0.1%
21 | 8 | 0.357% | -0.152% | 0.675% | 0.1%
21 | 9 | 0.446% | 0.089% | 0.437% | 0.1%
21 | 10 | 0.357% | -0.089% | 0.387% | 0.1%
21 | 11 | 0.33% | -0.027% | 0.378% | 0.1%
Deviations of %error on test id #22. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
22 | 0 | -26.98% | 0% | 0% | 0.1%
22 | 1 | -0.266% | 26.71% | 0% | 0.1%
22 | 2 | -0.266% | 0% | -9.171% | 0.1%
22 | 3 | -1.177% | -0.911% | -0.57% | 0.1%
22 | 4 | -1.186% | -0.009% | -0.876% | 0.1%
22 | 5 | 0.321% | 1.507% | -0.681% | 0.1%
22 | 6 | 1.306% | 0.985% | 0.147% | 0.1%
22 | 7 | 1.215% | -0.091% | 0.947% | 0.1%
22 | 8 | 0.348% | -0.867% | 0.956% | 0.1%
22 | 9 | 0.294% | -0.054% | 0.619% | 0.1%
22 | 10 | 0.321% | 0.027% | 0.321% | 0.1%
22 | 11 | 0.285% | -0.036% | 0.3% | 0.1%
Deviations of %error on test id #23. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
23 | 0 | -56.53% | 0% | 0% | 0.1%
23 | 1 | -2.74% | 53.79% | 0% | 0.1%
23 | 2 | 0.258% | 2.998% | -19.67% | 0.1%
23 | 3 | -0.794% | -1.052% | -1.092% | 0.1%
23 | 4 | -0.381% | 0.413% | -0.306% | 0.1%
23 | 5 | 0.617% | 0.998% | -0.186% | 0.1%
23 | 6 | 1.033% | 0.416% | 0.423% | 0.1%
23 | 7 | 0.834% | -0.199% | 0.828% | 0.1%
23 | 8 | 0.366% | -0.468% | 0.744% | 0.1%
23 | 9 | 0.366% | 0% | 0.522% | 0.1%
23 | 10 | 0.464% | 0.098% | 0.399% | 0.1%
23 | 11 | 0.375% | -0.089% | 0.402% | 0.1%
Deviations of %error on test id #24. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
24 | 0 | -46.47% | 0% | 0% | 0.1%
24 | 1 | -0.337% | 46.13% | 0% | 0.1%
24 | 2 | 0.044% | 0.381% | -15.59% | 0.1%
24 | 3 | -0.495% | -0.539% | -0.263% | 0.1%
24 | 4 | -1.875% | -1.38% | -0.775% | 0.1%
24 | 5 | 0.366% | 2.241% | -0.668% | 0.1%
24 | 6 | 0.879% | 0.513% | -0.21% | 0.1%
24 | 7 | 1.736% | 0.857% | 0.994% | 0.1%
24 | 8 | 0.42% | -1.316% | 1.012% | 0.1%
24 | 9 | 0.384% | -0.036% | 0.847% | 0.1%
24 | 10 | 0.375% | -0.009% | 0.393% | 0.1%
24 | 11 | 0.375% | 0% | 0.378% | 0.1%
Deviations of %error on test id #25. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
25 | 0 | -9.66% | 0% | 0% | 0.1%
25 | 1 | -0.018% | 9.642% | 0% | 0.1%
25 | 2 | -0.601% | -0.583% | -3.426% | 0.1%
25 | 3 | -0.785% | -0.184% | -0.468% | 0.1%
25 | 4 | 0.572% | 1.357% | -0.271% | 0.1%
25 | 5 | 0.888% | 0.316% | 0.225% | 0.1%
25 | 6 | 1.051% | 0.163% | 0.837% | 0.1%
25 | 7 | 0.545% | -0.506% | 0.828% | 0.1%
25 | 8 | 0.608% | 0.063% | 0.735% | 0.1%
25 | 9 | 0.446% | -0.162% | 0.533% | 0.1%
25 | 10 | 0.393% | -0.053% | 0.482% | 0.1%
25 | 11 | 0.348% | -0.045% | 0.396% | 0.1%
25 | 12 | 0.348% | 0% | 0.363% | 0.1%
Deviations of %error on test id #26. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
26 | 0 | -5.422% | 0% | 0% | 0.1%
26 | 1 | -0.169% | 5.253% | 0% | 0.1%
26 | 2 | -1.523% | -1.354% | -2.371% | 0.1%
26 | 3 | 0.446% | 1.969% | -0.415% | 0.1%
26 | 4 | 0.788% | 0.342% | -0.096% | 0.1%
26 | 5 | 1.297% | 0.509% | 0.844% | 0.1%
26 | 6 | 0.644% | -0.653% | 0.91% | 0.1%
26 | 7 | 0.545% | -0.099% | 0.829% | 0.1%
26 | 8 | 0.545% | 0% | 0.578% | 0.1%
26 | 9 | 0.689% | 0.144% | 0.593% | 0.1%
26 | 10 | 0.608% | -0.081% | 0.614% | 0.1%
26 | 11 | 0.734% | 0.126% | 0.677% | 0.1%
26 | 12 | 0.572% | -0.162% | 0.638% | 0.1%
26 | 13 | 0.545% | -0.027% | 0.617% | 0.1%
26 | 14 | 0.608% | 0.063% | 0.575% | 0.1%
26 | 15 | 0.698% | 0.09% | 0.617% | 0.1%
Deviations of %error on test id #27. Total 8 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
27 | 0 | -50.6% | 0% | 0% | 0.1%
27 | 1 | -0.372% | 50.22% | 0% | 0.1%
27 | 2 | -0.195% | 0.177% | -17.05% | 0.1%
27 | 3 | 0.125% | 0.32% | -0.147% | 0.1%
27 | 4 | 0.42% | 0.295% | 0.117% | 0.1%
27 | 5 | 0.339% | -0.081% | 0.295% | 0.1%
27 | 6 | 0.285% | -0.054% | 0.348% | 0.1%
27 | 7 | 0.312% | 0.027% | 0.312% | 0.1%
Deviations of %error on test id #28. Total 12 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
28 | 0 | -49.96% | 0% | 0% | 0.1%
28 | 1 | 0.258% | 50.21% | 0% | 0.1%
28 | 2 | 0.267% | 0.009% | -16.48% | 0.1%
28 | 3 | -0.951% | -1.218% | -0.142% | 0.1%
28 | 4 | 0.375% | 1.326% | -0.103% | 0.1%
28 | 5 | 0.752% | 0.377% | 0.059% | 0.1%
28 | 6 | 1.297% | 0.545% | 0.808% | 0.1%
28 | 7 | 0.536% | -0.761% | 0.862% | 0.1%
28 | 8 | 0.393% | -0.143% | 0.742% | 0.1%
28 | 9 | 0.402% | 0.009% | 0.444% | 0.1%
28 | 10 | 0.411% | 0.009% | 0.402% | 0.1%
28 | 11 | 0.411% | 0% | 0.408% | 0.1%
Deviations of %error on test id #29. Total 15 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
29 | 0 | -12.03% | 0% | 0% | 0.1%
29 | 1 | -5.842% | 6.184% | 0% | 0.1%
29 | 2 | 0.617% | 6.459% | -5.75% | 0.1%
29 | 3 | 0.366% | -0.251% | -1.62% | 0.1%
29 | 4 | 0.509% | 0.143% | 0.497% | 0.1%
29 | 5 | 0.366% | -0.143% | 0.414% | 0.1%
29 | 6 | 0.258% | -0.108% | 0.378% | 0.1%
29 | 7 | 0.205% | -0.053% | 0.276% | 0.1%
29 | 8 | 0.402% | 0.197% | 0.288% | 0.1%
29 | 9 | 0.518% | 0.116% | 0.375% | 0.1%
29 | 10 | 0.644% | 0.126% | 0.521% | 0.1%
29 | 11 | 0.788% | 0.144% | 0.65% | 0.1%
29 | 12 | 0.806% | 0.018% | 0.746% | 0.1%
29 | 13 | 0.761% | -0.045% | 0.785% | 0.1%
29 | 14 | 0.779% | 0.018% | 0.782% | 0.1%
Deviations of %error on test id #30. Total 16 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
30 | 0 | -20.69% | 0% | 0% | 0.1%
30 | 1 | -8.373% | 12.31% | 0% | 0.1%
30 | 2 | -4.879% | 3.494% | -11.31% | 0.1%
30 | 3 | 0.089% | 4.968% | -4.388% | 0.1%
30 | 4 | 0.196% | 0.107% | -1.531% | 0.1%
30 | 5 | 0.205% | 0.009% | 0.163% | 0.1%
30 | 6 | 0% | -0.205% | 0.134% | 0.1%
30 | 7 | -0.106% | -0.106% | 0.033% | 0.1%
30 | 8 | -0.151% | -0.045% | -0.086% | 0.1%
30 | 9 | -0.009% | 0.142% | -0.089% | 0.1%
30 | 10 | 0.205% | 0.214% | 0.015% | 0.1%
30 | 11 | 0.285% | 0.08% | 0.16% | 0.1%
30 | 12 | 0.411% | 0.126% | 0.3% | 0.1%
30 | 13 | 0.375% | -0.036% | 0.357% | 0.1%
30 | 14 | 0.375% | 0% | 0.387% | 0.1%
30 | 15 | 0.402% | 0.027% | 0.384% | 0.1%
Deviations of %error on test id #31. Total 20 readings. Reached max number of deviation, final %error may not reliable
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
31 | 0 | -50.6% | 0% | 0% | 0.1%
31 | 1 | 0.509% | 51.11% | 0% | 0.1%
31 | 2 | 0.366% | -0.143% | -16.57% | 0.1%
31 | 3 | -0.257% | -0.623% | 0.206% | 0.1%
31 | 4 | 0.348% | 0.605% | 0.152% | 0.1%
31 | 5 | 1.078% | 0.73% | 0.39% | 0.1%
31 | 6 | 1.617% | 0.539% | 1.014% | 0.1%
31 | 7 | 0.951% | -0.666% | 1.215% | 0.1%
31 | 8 | 0.779% | -0.172% | 1.116% | 0.1%
31 | 9 | 0.725% | -0.054% | 0.818% | 0.1%
31 | 10 | 0.608% | -0.117% | 0.704% | 0.1%
31 | 11 | 0.635% | 0.027% | 0.656% | 0.1%
31 | 12 | 0.743% | 0.108% | 0.662% | 0.1%
31 | 13 | 0.491% | -0.252% | 0.623% | 0.1%
31 | 14 | 0.411% | -0.08% | 0.548% | 0.1%
31 | 15 | 0.59% | 0.179% | 0.497% | 0.1%
31 | 16 | 0.599% | 0.009% | 0.533% | 0.1%
31 | 17 | 0.527% | -0.072% | 0.572% | 0.1%
31 | 18 | 0.402% | -0.125% | 0.509% | 0.1%
31 | 19 | 0.482% | 0.08% | 0.47% | 0.1%
Deviations of %error on test id #32. Total 11 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
32 | 0 | -13.85% | 0% | 0% | 0.1%
32 | 1 | 0.42% | 14.27% | 0% | 0.1%
32 | 2 | -0.662% | -1.082% | -4.698% | 0.1%
32 | 3 | -0.68% | -0.018% | -0.307% | 0.1%
32 | 4 | 0.527% | 1.207% | -0.272% | 0.1%
32 | 5 | 0.951% | 0.424% | 0.266% | 0.1%
32 | 6 | 0.752% | -0.199% | 0.743% | 0.1%
32 | 7 | 0.33% | -0.422% | 0.678% | 0.1%
32 | 8 | 0.348% | 0.018% | 0.477% | 0.1%
32 | 9 | 0.339% | -0.009% | 0.339% | 0.1%
32 | 10 | 0.348% | 0.009% | 0.345% | 0.1%
Deviations of %error on test id #33. Total 13 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
33 | 0 | -4.499% | 0% | 0% | 0.1%
33 | 1 | -2.259% | 2.24% | 0% | 0.1%
33 | 2 | 0.018% | 2.277% | -2.247% | 0.1%
33 | 3 | 0% | -0.018% | -0.747% | 0.1%
33 | 4 | -0.106% | -0.106% | -0.029% | 0.1%
33 | 5 | 0.044% | 0.15% | -0.021% | 0.1%
33 | 6 | -0.089% | -0.133% | -0.05% | 0.1%
33 | 7 | -0.053% | 0.036% | -0.033% | 0.1%
33 | 8 | 0.08% | 0.133% | -0.021% | 0.1%
33 | 9 | 0.205% | 0.125% | 0.077% | 0.1%
33 | 10 | 0.285% | 0.08% | 0.19% | 0.1%
33 | 11 | 0.339% | 0.054% | 0.276% | 0.1%
33 | 12 | 0.375% | 0.036% | 0.333% | 0.1%
Deviations of %error on test id #34. Total 6 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
34 | 0 | -15.88% | 0% | 0% | 0.1%
34 | 1 | -8.581% | 7.3% | 0% | 0.1%
34 | 2 | 0.437% | 9.018% | -8.008% | 0.1%
34 | 3 | 0.42% | -0.017% | -2.575% | 0.1%
34 | 4 | 0.464% | 0.044% | 0.44% | 0.1%
34 | 5 | 0.428% | -0.036% | 0.437% | 0.1%
Deviations of %error on test id #35. Total 9 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
35 | 0 | -12.33% | 0% | 0% | 0.1%
35 | 1 | -6.546% | 5.782% | 0% | 0.1%
35 | 2 | -0.018% | 6.528% | -6.297% | 0.1%
35 | 3 | 0.107% | 0.125% | -2.152% | 0.1%
35 | 4 | 0.107% | 0% | 0.065% | 0.1%
35 | 5 | -0.142% | -0.249% | 0.024% | 0.1%
35 | 6 | -0.098% | 0.044% | -0.044% | 0.1%
35 | 7 | -0.115% | -0.017% | -0.118% | 0.1%
35 | 8 | -0.018% | 0.097% | -0.077% | 0.1%
Deviations of %error on test id #36. Total 5 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
36 | 0 | 25.06% | 0% | 0% | 0.1%
36 | 1 | 0.375% | -24.68% | 0% | 0.1%
36 | 2 | 0.42% | 0.045% | 8.617% | 0.1%
36 | 3 | 0.375% | -0.045% | 0.39% | 0.1%
36 | 4 | 0.375% | 0% | 0.39% | 0.1%
Deviations of %error on test id #37. Total 7 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
37 | 0 | -12.27% | 0% | 0% | 0.1%
37 | 1 | -7.255% | 5.018% | 0% | 0.1%
37 | 2 | 0.267% | 7.522% | -6.42% | 0.1%
37 | 3 | 0.626% | 0.359% | -2.121% | 0.1%
37 | 4 | 0.662% | 0.036% | 0.518% | 0.1%
37 | 5 | 0.563% | -0.099% | 0.617% | 0.1%
37 | 6 | 0.581% | 0.018% | 0.602% | 0.1%
Deviations of %error on test id #38. Total 18 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
38 | 0 | -20.9% | 0% | 0% | 0.1%
38 | 1 | -11.4% | 9.501% | 0% | 0.1%
38 | 2 | -6.093% | 5.303% | -12.8% | 0.1%
38 | 3 | 0.098% | 6.191% | -5.797% | 0.1%
38 | 4 | 0.214% | 0.116% | -1.927% | 0.1%
38 | 5 | 0.312% | 0.098% | 0.208% | 0.1%
38 | 6 | 0.384% | 0.072% | 0.303% | 0.1%
38 | 7 | 0.205% | -0.179% | 0.3% | 0.1%
38 | 8 | 0.089% | -0.116% | 0.226% | 0.1%
38 | 9 | 0.125% | 0.036% | 0.14% | 0.1%
38 | 10 | -0.31% | -0.435% | -0.032% | 0.1%
38 | 11 | -0.275% | 0.035% | -0.153% | 0.1%
38 | 12 | -0.106% | 0.169% | -0.23% | 0.1%
38 | 13 | 0.187% | 0.293% | -0.065% | 0.1%
38 | 14 | 0.348% | 0.161% | 0.143% | 0.1%
38 | 15 | 0.42% | 0.072% | 0.318% | 0.1%
38 | 16 | 0.455% | 0.035% | 0.408% | 0.1%
38 | 17 | 0.455% | 0% | 0.443% | 0.1%
Deviations of %error on test id #39. Total 17 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
39 | 0 | -6.801% | 0% | 0% | 0.1%
39 | 1 | 0.134% | 6.935% | 0% | 0.1%
39 | 2 | 0.258% | 0.124% | -2.136% | 0.1%
39 | 3 | 0.339% | 0.081% | 0.244% | 0.1%
39 | 4 | 0.473% | 0.134% | 0.357% | 0.1%
39 | 5 | 0.25% | -0.223% | 0.354% | 0.1%
39 | 6 | 0.142% | -0.108% | 0.288% | 0.1%
39 | 7 | 0.214% | 0.072% | 0.202% | 0.1%
39 | 8 | -0.222% | -0.436% | 0.045% | 0.1%
39 | 9 | -0.275% | -0.053% | -0.094% | 0.1%
39 | 10 | -0.089% | 0.186% | -0.195% | 0.1%
39 | 11 | 0.187% | 0.276% | -0.059% | 0.1%
39 | 12 | 0.339% | 0.152% | 0.146% | 0.1%
39 | 13 | 0.509% | 0.17% | 0.345% | 0.1%
39 | 14 | 0.464% | -0.045% | 0.437% | 0.1%
39 | 15 | 0.464% | 0% | 0.479% | 0.1%
39 | 16 | 0.509% | 0.045% | 0.479% | 0.1%
Deviations of %error on test id #40. Total 4 readings.
Test Id | Reading No. | Error (%) | Deviation (%) | Average Error (%) | Deviation Thres. (%)
40 | 0 | 0.572% | 0% | 0% | 0.1%
40 | 1 | 0.653% | 0.081% | 0% | 0.1%
40 | 2 | 0.734% | 0.081% | 0.653% | 0.1%
40 | 3 | 0.752% | 0.018% | 0.713% | 0.1%