## Summary
Workspace file path is 'D:\HaiVo\SmartMeter\20230406 SMAT 3PH4W Calibration.zip'
Generated on 7/20/2023 11:33:11 AM by version Alpha4 - Build 2023 Apr. 12
1. Configuration for 3Ph4W Meter
Equipment Setup
Option Name |  |  |  |  | Value
Communication Port |  |  |  |  | COM15
Communication Port Settings |  |  |  |  | 9600, 8, None, One
Timeout, Delay and Retries |  |  |  |  | 0, 50, 5000, 0, 50, 5000, 1
Equipment Version |  |  |  |  | GF302D
Control phase angle between YR (U1U2) and YB (U1U3) |  |  |  |  | Yes
Indicates RYB channels by colors |  |  |  |  | Yes
Meter Platform Name |  |  |  |  | Global Reference Meter
Serial Number |  |  |  |  | SED-IV0004
Meter Type |  |  |  |  | TYPE_3P4W
Meter Class |  |  |  |  | Class 0.5S
Meter Rating |  |  |  |  | Non-specified
Meter Maximum Current (Imax) |  |  |  |  | 40
Meter Basic Voltage (Un or Ub) |  |  |  |  | 220
Meter Basic Current (Ib or In) |  |  |  |  | 5
Meter Frequency (Fb) |  |  |  |  | 50
Meter Constant Active Power (P) |  |  |  |  | 3200
Meter Constant Reactive Power (Q) |  |  |  |  | 3200
Meter % Error Average Length |  |  |  |  | 3
Meter % Error Rounding Decimals |  |  |  |  | 3
Discard first few %error responses |  |  |  |  | Yes
Discard first few %error, number of responses will be discarded |  |  |  |  | 1
Chain voltage or current among test cases |  |  |  |  | No
Meter Configuration
Option Name |  |  |  |  | Value
Optical Port |  |  |  |  | COM7
Optical Port Settings |  |  |  |  | 9600, 8, None, One
External COM Port
External COM Port Settings |  |  |  |  | 19200, 8, None, One
Private Protocol Settings |  |  |  |  | abcd.efgh.xx@renesas.com, 0, 50, 100, 20000, 1
2. Auto Calibration
Details in sheet ''1. Auto Calibration #4'!B6'
Path in workspace, 'Workspace / 3Ph4W Meter / Renesas Private Protocol Group #1 / Auto Calibration #4'
3. Accuracy Test
Below table summarizes all test categories and test result
No | Path In Workspace | Test Category | Passed | Failed | Inconclusive | Total | Overall Result | Accuracy Chart | Details

## 1. Auto Calibration #4
[<<]
Auto Calibration Report of 'Auto Calibration #4'
Path in workspace, 'Workspace / 3Ph4W Meter / Renesas Private Protocol Group #1 / Auto Calibration #4'
Workspace file path is 'D:\HaiVo\SmartMeter\20230406 SMAT 3PH4W Calibration.zip'
Generated on 7/20/2023 11:33:10 AM by version Alpha4 - Build 2023 Apr. 12
1. Reference parameters for setup of Auto Calibration
Parameter Name |  | Value
Cycle for signal accumulation |  | 100
Cycle for phase angle accumulation |  | 100
Max Current |  | 60
Reference Voltage |  | 220
Reference Current |  | 5
Reference Frequency |  | 50
Reference Phase Angle |  | 0
2. Calibration Parameters Output
Total 4 line(s) calibrated and have parameters as shown in below table
Parameter Name | Phase R | Phase Y | Phase B | Neutral Channel
Selected | TRUE | TRUE | TRUE | TRUE
Sampling Frequency | 3907 | 3906 | 3908 | 0
Gain 0 | 2 | 2 | 2 | 2
Gain 1 | 0 | 0 | 0 | 0
Voltage Coeff | 37.2309 | 37.2042 | 37.2067 | 0
Current Coeff | 84121.48 | 83654.89 | 84209.59 | 83886.22
Power Coeff | 3131916 | 3112313 | 3133159 | 0
Angle Error 0 | -2.173 | -2.1316 | -2.0973 | 0
Angle Error 1 | 0 | 0 | 0 | 0
Angle 90 Offset | 0 | 0 | 0 | 0
Voltage Offset | 0 | 0 | 0 | 0
Current Offset | 0 | 0 | 0 | 0
3. Test Log
All Messages |  |  |  |  | GFUVE | Private Protocol
[2023-07-20 11:31:26] Start comms with test equipment in COM15. |  |  |  |  | [2023-07-20 11:31:28] 6805006D16 | [2023-07-20 11:31:47] 0201240017153C67C0E785A3ACD91B24C3C2AF784E2F625775BA3BE7AE0D31738F1718BA982603
[2023-07-20 11:31:27] COM15 was openned. |  |  |  |  | [2023-07-20 11:31:28] 689A000216 | [2023-07-20 11:31:47] 02030600000188A803
[2023-07-20 11:31:28] Set calib pararmeters to test equipment |  |  |  |  | [2023-07-20 11:31:29] 6801006916 | [2023-07-20 11:31:47] 02011301646442700000435C000040A0000000D87203
[2023-07-20 11:31:28] Exec UI output down |  |  |  |  | [2023-07-20 11:31:29] 689A000216 | [2023-07-20 11:32:00] 0203270100004574300040000000000000004214EC6A47A44CBE4A3F2831C00B12F500000000FF4CD803
[2023-07-20 11:31:28] 6805006D16 |  |  |  |  | [2023-07-20 11:31:29] 6810023435E316 | [2023-07-20 11:32:02] 0201240017153C67C0E785A3ACD91B24C3C2AF784E2F625775BA3BE7AE0D31738F1718BA982603
[2023-07-20 11:31:28] Waiting response in 5000ms |  |  |  |  | [2023-07-20 11:31:30] 689A000216 | [2023-07-20 11:32:02] 02030600000188A803
[2023-07-20 11:31:28] Response, NormalResponse, Data { None } |  |  |  |  | [2023-07-20 11:31:30] 6806006E16 | [2023-07-20 11:32:03] 02011301646442700000435C000040A0000001C9FB03
[2023-07-20 11:31:28] 689A000216 |  |  |  |  | [2023-07-20 11:31:30] 689A000216 | [2023-07-20 11:32:15] 0203270100004574200040000000000000004214D11A47A363724A3DF5E5C0086C0100000000FF1EB803
[2023-07-20 11:31:29] Delay 50ms. |  |  |  |  | [2023-07-20 11:31:30] 680702A3738716 | [2023-07-20 11:32:17] 0201240017153C67C0E785A3ACD91B24C3C2AF784E2F625775BA3BE7AE0D31738F1718BA982603
[2023-07-20 11:31:29] Exec Set meter type TYPE_3P4W |  |  |  |  | [2023-07-20 11:31:30] 689A000216 | [2023-07-20 11:32:17] 02030600000188A803
[2023-07-20 11:31:29] 6801006916 |  |  |  |  | [2023-07-20 11:31:31] 68070435538333B116 | [2023-07-20 11:32:18] 02011301646442700000435C000040A0000002FB6003
[2023-07-20 11:31:29] Waiting response in 5000ms |  |  |  |  | [2023-07-20 11:31:31] 689A000216 | [2023-07-20 11:32:30] 0203270100004574400040000000000000004214D3A147A478CC4A3F3B9AC006396A00000000FFFB3503
[2023-07-20 11:31:29] Response, NormalResponse, Data { None } |  |  |  |  | [2023-07-20 11:31:31] 680706435553333333F916 | [2023-07-20 11:32:32] 0201240017153C67C0E785A3ACD91B24C3C2AF784E2F625775BA3BE7AE0D31738F1718BA982603
[2023-07-20 11:31:29] 689A000216 |  |  |  |  | [2023-07-20 11:31:31] 689A000216 | [2023-07-20 11:32:32] 02030600000188A803
[2023-07-20 11:31:29] Delay 50ms. |  |  |  |  | [2023-07-20 11:31:31] 6807065355533333330916 | [2023-07-20 11:32:32] 02011301646442700000435C000040A0000003EAE903
[2023-07-20 11:31:29] Exec Set direction and meter type before UI output up, DP, TYPE_3P4W |  |  |  |  | [2023-07-20 11:31:31] 689A000216 | [2023-07-20 11:32:36] 0203270100000000000040000000000000000000000047A3D71C000000000000000000000000FFB12003
[2023-07-20 11:31:29] 6810023435E316 |  |  |  |  | [2023-07-20 11:31:32] 6807067355533333332916
[2023-07-20 11:31:29] Waiting response in 5000ms |  |  |  |  | [2023-07-20 11:31:32] 689A000216
[2023-07-20 11:31:30] Response, NormalResponse, Data { None } |  |  |  |  | [2023-07-20 11:31:32] 680706445338333333DD16
[2023-07-20 11:31:30] 689A000216 |  |  |  |  | [2023-07-20 11:31:32] 689A000216
[2023-07-20 11:31:30] Delay 50ms. |  |  |  |  | [2023-07-20 11:31:32] 680706545338333333ED16
[2023-07-20 11:31:30] Exec Clear all harmonics |  |  |  |  | [2023-07-20 11:31:32] 689A000216
[2023-07-20 11:31:30] 6806006E16 |  |  |  |  | [2023-07-20 11:31:32] 6807067453383333330D16
[2023-07-20 11:31:30] Waiting response in 5000ms |  |  |  |  | [2023-07-20 11:31:33] 689A000216
[2023-07-20 11:31:30] Response, NormalResponse, Data { None } |  |  |  |  | [2023-07-20 11:31:33] 6808007016
[2023-07-20 11:31:30] 689A000216 |  |  |  |  | [2023-07-20 11:31:33] 689A000216
[2023-07-20 11:31:30] Delay 50ms. |  |  |  |  | [2023-07-20 11:31:33] 6811093333333333333333334D16
[2023-07-20 11:31:30] Exec Up Voltage Phases R Y B |  |  |  |  | [2023-07-20 11:31:33] 689A000216
[2023-07-20 11:31:30] 680702A3738716 |  |  |  |  | [2023-07-20 11:31:33] 6803006B16
[2023-07-20 11:31:30] Waiting response in 5000ms |  |  |  |  | [2023-07-20 11:31:33] 689A000216
[2023-07-20 11:31:30] Response, NormalResponse, Data { None } |  |  |  |  | [2023-07-20 11:31:39] 680704465333337216
[2023-07-20 11:31:30] 689A000216 |  |  |  |  | [2023-07-20 11:31:39] 689A000216
[2023-07-20 11:31:30] Delay 50ms. |  |  |  |  | [2023-07-20 11:31:39] 680704565333338216
[2023-07-20 11:31:31] Exec Set frequency 50, actual set 50.00 |  |  |  |  | [2023-07-20 11:31:39] 689A000216
[2023-07-20 11:31:31] 68070435538333B116 |  |  |  |  | [2023-07-20 11:31:39] 68070476533333A216
[2023-07-20 11:31:31] Waiting response in 5000ms |  |  |  |  | [2023-07-20 11:31:39] 689A000216
[2023-07-20 11:31:31] Response, NormalResponse, Data { None } |  |  |  |  | [2023-07-20 11:31:40] 680704375453338416
[2023-07-20 11:31:31] 689A000216 |  |  |  |  | [2023-07-20 11:31:40] 689A000216
[2023-07-20 11:31:31] Delay 50ms. |  |  |  |  | [2023-07-20 11:31:40] 68070438557333A616
[2023-07-20 11:31:31] Exec Set voltage Y channel 220, actual set 0220.000000 |  |  |  |  | [2023-07-20 11:31:40] 689A000216
[2023-07-20 11:31:31] 680706435553333333F916 |  |  |  |  | [2023-07-20 11:32:39] 6805006D16
[2023-07-20 11:31:31] Waiting response in 5000ms |  |  |  |  | [2023-07-20 11:32:39] 689A000216
[2023-07-20 11:31:31] Response, NormalResponse, Data { None }
[2023-07-20 11:31:31] 689A000216
[2023-07-20 11:31:31] Delay 50ms.
[2023-07-20 11:31:31] Exec Set voltage R channel 220, actual set 0220.000000
[2023-07-20 11:31:31] 6807065355533333330916
[2023-07-20 11:31:31] Waiting response in 5000ms
[2023-07-20 11:31:31] Response, NormalResponse, Data { None }
[2023-07-20 11:31:31] 689A000216
[2023-07-20 11:31:31] Delay 50ms.
[2023-07-20 11:31:32] Exec Set voltage B channel 220, actual set 0220.000000
[2023-07-20 11:31:32] 6807067355533333332916
[2023-07-20 11:31:32] Waiting response in 5000ms
[2023-07-20 11:31:32] Response, NormalResponse, Data { None }
[2023-07-20 11:31:32] 689A000216
[2023-07-20 11:31:32] Delay 50ms.
[2023-07-20 11:31:32] Exec Set current Y channel 5, actual set 0005.000000
[2023-07-20 11:31:32] 680706445338333333DD16
[2023-07-20 11:31:32] Waiting response in 5000ms
[2023-07-20 11:31:32] Response, NormalResponse, Data { None }
[2023-07-20 11:31:32] 689A000216
[2023-07-20 11:31:32] Delay 50ms.
[2023-07-20 11:31:32] Exec Set current R channel 5, actual set 0005.000000
[2023-07-20 11:31:32] 680706545338333333ED16
[2023-07-20 11:31:32] Waiting response in 5000ms
[2023-07-20 11:31:32] Response, NormalResponse, Data { None }
[2023-07-20 11:31:32] 689A000216
[2023-07-20 11:31:32] Delay 50ms.
[2023-07-20 11:31:32] Exec Set current B channel 5, actual set 0005.000000
[2023-07-20 11:31:32] 6807067453383333330D16
[2023-07-20 11:31:32] Waiting response in 5000ms
[2023-07-20 11:31:33] Response, NormalResponse, Data { None }
[2023-07-20 11:31:33] 689A000216
[2023-07-20 11:31:33] Delay 50ms.
[2023-07-20 11:31:33] Exec Stop read result data
[2023-07-20 11:31:33] 6808007016
[2023-07-20 11:31:33] Waiting response in 5000ms
[2023-07-20 11:31:33] Response, NormalResponse, Data { None }
[2023-07-20 11:31:33] 689A000216
[2023-07-20 11:31:33] Delay 50ms.
[2023-07-20 11:31:33] Exec Stop verify error
[2023-07-20 11:31:33] 6811093333333333333333334D16
[2023-07-20 11:31:33] Waiting response in 5000ms
[2023-07-20 11:31:33] Response, NormalResponse, Data { None }
[2023-07-20 11:31:33] 689A000216
[2023-07-20 11:31:33] Delay 50ms.
[2023-07-20 11:31:33] Exec UI output up
[2023-07-20 11:31:33] 6803006B16
[2023-07-20 11:31:33] Waiting response in 5000ms
[2023-07-20 11:31:33] Response, NormalResponse, Data { None }
[2023-07-20 11:31:33] 689A000216
[2023-07-20 11:31:34] Delay 5050ms.
[2023-07-20 11:31:39] Exec Set phase Y channel 0, actual set 0000.00
[2023-07-20 11:31:39] 680704465333337216
[2023-07-20 11:31:39] Waiting response in 5000ms
[2023-07-20 11:31:39] Response, NormalResponse, Data { None }
[2023-07-20 11:31:39] 689A000216
[2023-07-20 11:31:39] Delay 50ms.
[2023-07-20 11:31:39] Exec Set phase R channel 0, actual set 0000.00
[2023-07-20 11:31:39] 680704565333338216
[2023-07-20 11:31:39] Waiting response in 5000ms
[2023-07-20 11:31:39] Response, NormalResponse, Data { None }
[2023-07-20 11:31:39] 689A000216
[2023-07-20 11:31:39] Delay 50ms.
[2023-07-20 11:31:39] Exec Set phase B channel 0, actual set 0000.00
[2023-07-20 11:31:39] 68070476533333A216
[2023-07-20 11:31:39] Waiting response in 5000ms
[2023-07-20 11:31:39] Response, NormalResponse, Data { None }
[2023-07-20 11:31:39] 689A000216
[2023-07-20 11:31:40] Delay 50ms.
[2023-07-20 11:31:40] Exec Set Uab 120, actual set 0120.00
[2023-07-20 11:31:40] 680704375453338416
[2023-07-20 11:31:40] Waiting response in 5000ms
[2023-07-20 11:31:40] Response, NormalResponse, Data { None }
[2023-07-20 11:31:40] 689A000216
[2023-07-20 11:31:40] Delay 50ms.
[2023-07-20 11:31:40] Exec Set Uac 240, actual set 0240.00
[2023-07-20 11:31:40] 68070438557333A616
[2023-07-20 11:31:40] Waiting response in 5000ms
[2023-07-20 11:31:40] Response, NormalResponse, Data { None }
[2023-07-20 11:31:40] 689A000216
[2023-07-20 11:31:40] Delay 50ms.
[2023-07-20 11:31:40] Finished all commands.
[2023-07-20 11:31:40] Waiting 5s for meter stable.
[2023-07-20 11:31:45] Start Renesas Comms (Private Protocol) with meter
[2023-07-20 11:31:46] Optical port COM7 was openned.
[2023-07-20 11:31:47] Start auto calibration.
[2023-07-20 11:31:47] Exec Authenticate with email "abcd.efgh.xx@renesas.com"
[2023-07-20 11:31:47] 0201240017153C67C0E785A3ACD91B24C3C2AF784E2F625775BA3BE7AE0D31738F1718BA982603
[2023-07-20 11:31:47] Waiting response in 20000ms
[2023-07-20 11:31:47] Response, Data { 000001, Authentication, OK }
[2023-07-20 11:31:47] 02030600000188A803
[2023-07-20 11:31:47] Delay 50ms.
[2023-07-20 11:31:47] Exec Start calibrate with (c = 100, cp = 100, imax = 60, v = 220, i = 5, line = Phase1
[2023-07-20 11:31:47] 02011301646442700000435C000040A0000000D87203
[2023-07-20 11:31:47] Waiting response in 20000ms
[2023-07-20 11:32:00] Response, Data { 0100004574300040000000000000004214EC6A47A44CBE4A3F2831C00B12F500000000FF, Calibration, OK }
[2023-07-20 11:32:00] 0203270100004574300040000000000000004214EC6A47A44CBE4A3F2831C00B12F500000000FF4CD803
[2023-07-20 11:32:00] Step = 255. Sampling frequency = 3907. Gain0 = 2. Gain1 = 0. Voltage coeff = 37.23087. Current coeff = 84121.48. Power coeff = 3131916. Angle error 0 = -2.173032. Angle error 1 = 0
[2023-07-20 11:32:00] Delay 2050ms.
[2023-07-20 11:32:00] Has calibration result for Phase1 channel.
[2023-07-20 11:32:02] Exec Authenticate with email "abcd.efgh.xx@renesas.com"
[2023-07-20 11:32:02] 0201240017153C67C0E785A3ACD91B24C3C2AF784E2F625775BA3BE7AE0D31738F1718BA982603
[2023-07-20 11:32:02] Waiting response in 20000ms
[2023-07-20 11:32:02] Response, Data { 000001, Authentication, OK }
[2023-07-20 11:32:02] 02030600000188A803
[2023-07-20 11:32:02] Delay 50ms.
[2023-07-20 11:32:03] Exec Start calibrate with (c = 100, cp = 100, imax = 60, v = 220, i = 5, line = Phase2
[2023-07-20 11:32:03] 02011301646442700000435C000040A0000001C9FB03
[2023-07-20 11:32:03] Waiting response in 20000ms
[2023-07-20 11:32:15] Response, Data { 0100004574200040000000000000004214D11A47A363724A3DF5E5C0086C0100000000FF, Calibration, OK }
[2023-07-20 11:32:15] 0203270100004574200040000000000000004214D11A47A363724A3DF5E5C0086C0100000000FF1EB803
[2023-07-20 11:32:15] Step = 255. Sampling frequency = 3906. Gain0 = 2. Gain1 = 0. Voltage coeff = 37.2042. Current coeff = 83654.89. Power coeff = 3112313. Angle error 0 = -2.131592. Angle error 1 = 0
[2023-07-20 11:32:15] Delay 2050ms.
[2023-07-20 11:32:15] Has calibration result for Phase2 channel.
[2023-07-20 11:32:17] Exec Authenticate with email "abcd.efgh.xx@renesas.com"
[2023-07-20 11:32:17] 0201240017153C67C0E785A3ACD91B24C3C2AF784E2F625775BA3BE7AE0D31738F1718BA982603
[2023-07-20 11:32:17] Waiting response in 20000ms
[2023-07-20 11:32:17] Response, Data { 000001, Authentication, OK }
[2023-07-20 11:32:17] 02030600000188A803
[2023-07-20 11:32:17] Delay 50ms.
[2023-07-20 11:32:18] Exec Start calibrate with (c = 100, cp = 100, imax = 60, v = 220, i = 5, line = Phase3
[2023-07-20 11:32:18] 02011301646442700000435C000040A0000002FB6003
[2023-07-20 11:32:18] Waiting response in 20000ms
[2023-07-20 11:32:30] Response, Data { 0100004574400040000000000000004214D3A147A478CC4A3F3B9AC006396A00000000FF, Calibration, OK }
[2023-07-20 11:32:30] 0203270100004574400040000000000000004214D3A147A478CC4A3F3B9AC006396A00000000FFFB3503
[2023-07-20 11:32:30] Step = 255. Sampling frequency = 3908. Gain0 = 2. Gain1 = 0. Voltage coeff = 37.20667. Current coeff = 84209.59. Power coeff = 3133159. Angle error 0 = -2.097254. Angle error 1 = 0
[2023-07-20 11:32:30] Has calibration result for Phase3 channel.
[2023-07-20 11:32:30] Delay 2050ms.
[2023-07-20 11:32:32] Exec Authenticate with email "abcd.efgh.xx@renesas.com"
[2023-07-20 11:32:32] 0201240017153C67C0E785A3ACD91B24C3C2AF784E2F625775BA3BE7AE0D31738F1718BA982603
[2023-07-20 11:32:32] Waiting response in 20000ms
[2023-07-20 11:32:32] Response, Data { 000001, Authentication, OK }
[2023-07-20 11:32:32] 02030600000188A803
[2023-07-20 11:32:32] Delay 50ms.
[2023-07-20 11:32:32] Exec Start calibrate with (c = 100, cp = 100, imax = 60, v = 220, i = 5, line = Neutral
[2023-07-20 11:32:32] 02011301646442700000435C000040A0000003EAE903
[2023-07-20 11:32:33] Waiting response in 20000ms
[2023-07-20 11:32:36] Response, Data { 0100000000000040000000000000000000000047A3D71C000000000000000000000000FF, Calibration, OK }
[2023-07-20 11:32:36] 0203270100000000000040000000000000000000000047A3D71C000000000000000000000000FFB12003
[2023-07-20 11:32:36] Step = 255. Sampling frequency = 0. Gain0 = 2. Gain1 = 0. Voltage coeff = 0. Current coeff = 83886.22. Power coeff = 0. Angle error 0 = 0. Angle error 1 = 0
[2023-07-20 11:32:36] Has calibration result for Neutral channel.
[2023-07-20 11:32:36] Delay 2050ms.
[2023-07-20 11:32:38] Finished all commands.
[2023-07-20 11:32:38] Optical port receiver stopped.
[2023-07-20 11:32:38] Renesas Comms stopped.
[2023-07-20 11:32:39] Optical port COM7 was closed.
[2023-07-20 11:32:39] Exec UI output down
[2023-07-20 11:32:39] 6805006D16
[2023-07-20 11:32:39] Waiting response in 5000ms
[2023-07-20 11:32:39] Response, NormalResponse, Data { None }
[2023-07-20 11:32:39] 689A000216
[2023-07-20 11:32:39] Delay 50ms.
[2023-07-20 11:32:39] Finished all commands.
[2023-07-20 11:32:40] Frame receiver on test equipment was stopped.
[2023-07-20 11:32:40] Comms on test equipment was stopped. COM15 was closed.